## RoomData
Standard Components - Items in room
15.12.2025
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
1BR-CCU | 1 Bed Room - CCU | 20 | 24 Hours | 1 patient; 1-2 visitors, 1-4 staff intermittently |  | Acoustic privacy is required | A 1 Bed Room - CCU will accommodate one patient and support the delivery of a range of clinical care, minor procedures and diagnostic tests, with clinical staff requiring access to both sides of the bed. Additional Considerations: The 1 Bed Room - CCU may vary in configuration to suit different Ensuite options including inboard, outboard and stacked (also known as 'back-to-back' or 'nested') configurations. In each configuration, sufficient clearance is required for equipment and staff to access both sides of the bed. Glazed doors and observation windows may be provided for visibility of high acuity patients and monitoring equipment. The provision of observation windows and the use of interstitial blinds or privacy curtains is to be supported by operational models for observation as well as support for patient privacy. The configuration of the Medical Services Panel (MSP), including medical gases, power and data, will be dependent on service requirements. Data provision is to take into consideration the increasing demand and sophistication of information systems and the technological development of monitoring and diagnostic equipment. Natural light and outlook is essential and the patient will need to have an external outlook from either the chair or the bed, preferably both.
ADLB | ADL Bathroom | 12 | up to 24 hours | 1 patient, 1 - 2 staff |  | Sound attenuation level medium, Acoustic privacy is required. | The ADL Bathroom provides a toilet, basin with vanity, shower and bath for the training and assessment of patients in Activities of Daily Living (ADL) tasks. The room is to provide domestic facilities with grab rails and other assistive devices to be included to reflect reasonable home modifications and equipment utilised by patients. The entrance of the room is to be flush with the adjoining corridor. The door is to be fitted with escape hardware for the door to open outwards if required. Additional Considerations: An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. The shower area may include a fixed raised hob to simulate a domestic scenario depending on service requirements.
ADLD | ADL Dining | 12 | up to 24 hours | up to 4 patients, 1-2 staff |  | sound attenuation level-high; acoustic privacy is required. | The ADL Dining provides a flexible space for the training and assessment of patients in Activities of Daily Living (ADL) tasks. Typically collocated with the ADL Kitchen, the Dining room supports small group discussions and demonstrations between both rooms and provides a space for debriefing after training and assessment of ADL tasks in the Kitchen. Additional Considerations: An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. If multiple ADL rooms are combined into an ADL suite nurse call buttons can be rationalised into a central location. External outlook and access to natural light is desirable.
ADLK | ADL Kitchen | 12 | up to 24 hours | up to 2 patients, 1-2 staff |  | Acoustics: sound attenuation level high. | The ADL Kitchen provides an accessible space for the training and assessment of patients in Activities of Daily Living (ADL) tasks. The space is to be functional for a range of patients including wheelchair users. The kitchen can include a combination of fittings and appliances such as a sink, cooktop, oven/grill, microwave and refrigerator to support common ADL training and assessment tasks as per service requirements. Room exhaust to be provided to range hood over cooktop. Additional Considerations: An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. If multiple ADL rooms are combined into an ADL suite nurse call buttons can be rationalised into a central location. The room may be enclosed with a lockable door depending on security requirements. External outlook and access to natural light is desirable.
ADLL | ADL Laundry | 8 | up to 24 Hours | 1 patient, 1 staff |  |  | The ADL Laundry provides domestic facilities for the training and assessment of patients in Activities of Daily Living (ADL) tasks. The room will include a domestic washing machine and dryer which, if front loading models are selected, may be placed on plinths to raise machines to an accessible height to support common ADL training and assessment tasks as per service requirements. A laundry tub and storage for ironing equipment are also provided. Additional Considerations: An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. If multiple ADL rooms are combined into an ADL suite nurse call buttons can be rationalised into a central location. The space may be provided as an open bay that is collocated with other ADL rooms. If open, the impact of noise on adjacent areas must be considered.
AHBF | After Hours Blood Fridge | 8 | Available 24 Hours | 1-2 staff, intermittently |  |  | The After Hours Blood Fridge is a discreet area for refrigerated storage of blood and blood products for access by authorised personnel after hours where a 24 hour pathology service is not available. This room is typically located on the periphery of a Pathology unit and may have direct, secure access to the Pathology laboratory with ready access to the transfusion service area. Additional Considerations: Refrigerator temperature monitoring is required and is to be connected to the building management system (BMS). The alarm is to provide alerts where 24-hour monitoring is undertaken, within the unit and/or remotely. Temperature monitoring is to meet ANZSBT guidelines. Additional fridges/freezers or lockable sections within the after hours fridge may be provided to support Pathology laboratory workflows. Provision of smart fridges for blood product tracking and other ICT requirements will be based on local policies and service requirements. Door width to allow for installation of fridge, including considerations for future proofing as equipment is replaced.
AHDR | After Hours Drug Store | 5 | 24 Hours | 1 - 2 staff intermittently |  |  | The After Hours Drug Store is a secure room for the storage of medications that may be required after hours. Stock is replenished by pharmacy personnel and, depending on the location of the room, there may be direct access from within the pharmacy. Secure access to the room for authorised personnel is required. A duress alarm may be required if room accessed from a public or patient care area. Refrigerated storage will be required for some medications. Refer to project specific requirements for automated medication management systems.
AIRLE-12 | Airlock - Entry, 12m2 | 12 | Up to 24 Hours |  |  |  | The Airlock is a transitional space associated with an entry area to provide wind/weather protection and to support efficient temperature regulation. Flooring in the Airlock typically supports the capture of dirt and moisture at the building entry to reduce cleaning requirements. Configuration will be dependent on site, internal layout and access requirements. The Airlock may be used by a wide range of people including fully ambulant people, persons using a wide range of walking aids and persons in wheelchairs. The Airlock may need to support the movement of beds/trolleys (e.g. at an ambulance or transit lounge entry) requiring a configuration that provides a relatively straight path of travel. Additional area to accommodate door movement and circulation will be required if entry from multiple directions provided. Additional Considerations: Entry doors may be glazed and automatic depending on access requirements. Access control may require the provision of swipe card readers, keypads, intercoms, push to open buttons (e.g. to exit after hours), door release buttons from nearby security or reception, etc. Provision of these items will depend on security requirements and operational policies (e.g. for after hours access). Lockdown and emergency egress requirements must also be considered. A hold open function to the internal door will typically be required for cleaning. Depending on layout and alignment of openings when the doors are open, an air curtain may be required to prevent the entry of draughts, insects, odours, etc. Wall materials shown are indicative and will be dependent on facade design and internal layout. Floor finishes are to be specified/detailed to reduce the risk of trips, slips and falls, (e.g. minimise changes in finished height at transitions). Drainage requirements to floor matting must be considered. Where allocated to an Ambulance entry floor finish is to provide a flat smooth surface suitable for transporting spinal and other patients. At public/patient entries, the provision of culturally appropriate artwork (e.g. with film on glazing) may be considered to create a welcoming atmosphere. Hand hygiene facilities, such as alcohol-based hand rub, should be provided nearby upon entry in line with local infection control policies. A small amount of concise signage may be required within the airlock however it is recommended that extensive wayfinding information is made readily accessible upon entry to keep airlock clear.
AIRLE-6 | Airlock - Entry, 6m2 | 6 | Up to 24 Hours |  |  |  | The Airlock is a transitional space associated with an entry area to provide wind/weather protection and to support efficient temperature regulation. Flooring in the Airlock typically supports the capture of dirt and moisture at the building entry to reduce cleaning requirements. Configuration will be dependent on site, internal layout and access requirements. The Airlock may be used by a wide range of people including fully ambulant people, persons using a wide range of walking aids and persons in wheelchairs. Additional area and a layout with a relatively straight path of travel is required if the movement of beds/trolleys is to be supported. Additional Considerations: Entry doors may be glazed and automatic depending on access requirements. Access control may require the provision of swipe card readers, keypads, intercoms, push to open buttons (e.g. to exit after hours), door release buttons from nearby security or reception, etc. Provision of these items will depend on security requirements and operational policies (e.g. for after hours access). Lockdown and emergency egress requirements must also be considered. A hold open function to the internal door will typically be required for cleaning. Depending on layout and alignment of openings when the doors are open, an air curtain may be required to prevent the entry of draughts, insects, odours, etc. Wall materials shown are indicative and will be dependent on facade design and internal layout. Floor finishes are to be specified/detailed to reduce the risk of trips, slips and falls, (e.g. minimise changes in finished height at transitions). Drainage requirements to floor matting must be considered. At public/patient entries, the provision of culturally appropriate artwork (e.g. with film on glazing) may be considered to create a welcoming atmosphere. Hand hygiene facilities, such as alcohol-based hand rub, should be provided nearby upon entry in line with local infection control policies. A small amount of concise signage may be required within the airlock however it is recommended that extensive wayfinding information is made readily accessible upon entry to keep airlock clear.
AUD-CR | Audiology Control Room | 9 | 8am to 6pm, daily | 1 staff |  | Acoustic privacy is required | The Audiology Control Room may be provided adjacent to the Audiology Testing Room to support the administering of tests that require discrete observation, result verification by a second clinician and/or remote control of equipment. The room will require a workspace and visibility of the Testing Room via an observation window. Additional Consideration: The floor of the Control Room should be flush with the floor of the Treatment Room and the adjacent corridor for access by patients in wheelchairs. Portable equipment may be shared between the Control Room and the Treatment Room.
AUD-TR | Audiology Testing Room | 14 | 8am to 6pm, daily | 1 patient, 1 visitor/carer, 1-2 staff |  | Acoustic isolation is required | The Audiology Testing Room provides an acoustically treated, soundproof space for clinical assessments to be carried out under controlled conditions. Testing Rooms may be modular proprietary items or constructed in situ and are generally designed as rooms within rooms in order to isolate the Testing Room from the rest of the building as much as possible. All designs should be reviewed and confirmed by an acoustic design expert. Additional Consideration: The Audiology Testing Room currently indicates a layout (with an adjacent control room) that is suitable for paediatric testing. Layouts will vary to suit service requirements and the provided layout will need to be altered to suit services for older children and adults. The audiometric booth is to be ventilated and connected to air-conditioning, power and data. Noise from mechanical services is to be minimised with air-conditioning to be designed for low noise operation to meet ambient noise requirements for testing. The floor of the Treatment Room should be flush with the floor of the Control Room, or the adjacent corridor where a Control Room is not provided, for access by patients in wheelchairs.
BATH | Bathroom | 15 | 24 Hours | 1 patient; 1-2 staff assisting |  | Sound attenuation level medium, Acoustic privacy is required. | The Bathroom provides facilities for assisted bathing, dressing and grooming of patients. The layout should allow for maneuvering of a patient on a lifting device, commode chair, wheelchair or shower trolley. Placement of the toilet should allow assistance from both sides. A height adjustable bath may be provided instead of a shower trolley depending on service requirements. Bariatric capability may be considered; refer to project specific requirements. Signage with braille tactiles to nominate 'Bathroom' to be provided.
BBEV | Bay - Beverage | 4 | Up to 24 Hours | 1-2 intermittently |  |  | The Beverage Bay is an area for preparation (including heating as required) of refreshments, snacks and small meals, washing some utensils, storing small mounts of food and drink and disposing of food waste (in accordance with food hygiene requirements). Beverage bays may allocated to patient care areas, staff areas and meeting/education facilities. Additional Considerations: Ready access to a handwashing basin is recommended. If an instantaneous boiling water tap/unit is provided, it should be hardwired and access restricted to the unit e.g. within cupboard. If system is located within a joinery unit, ensure appropriate ventilation is provided. A side approach to the sink for wheelchair access has been shown; front access may be required depending on local/jurisdictional policies and service requirements. The Beverage Bay should be adequately recessed to avoid encroachment into the corridor. If the space is to be enclosed and/or access is restricted, an increase in area is required. The door will require an observation panel. Size, type and quantity of waste bins is dependent on environmental management policies and service requirements. Where a facility has additional waste streams, beyond the general waste and comingled recycling indicatively shown in this Standard Component, additional area may be required to accommodate the receptacles. Depending on frequency of waste removal, bin sizes may be larger in high use areas or as guided by local waste management policies. Joinery to conceal bins may be provided depending on infection prevention and control policies. Fridge dimensions will vary depending on selection and service requirements; clear area provided for the fridge is to allow for adequate ventilation and full opening of door Provision and type of ice machine will be dependent on service requirements and local infection prevention and control policies.
BBLD | Bay - Blood Fridge | 2 | Available 24 Hours | accessed by staff intermittently |  |  | The Bay - Blood Fridge is an area provided within a secure, staff only area for the storage of blood products. This bay is typically provided in critical care and operating theatre units where the pathology department is not located nearby. A pass-through layout may be developed depending on service requirements and location (e.g. in relation to limiting access to the clean area of operating theatres). Additional Considerations: Refrigerator temperature monitoring is required and is to be connected to the building management system (BMS). The alarm is to provide alerts where 24-hour monitoring is undertaken, within the unit and/or remotely. Temperature monitoring is to meet ANZSBT guidelines. Requirements for accurate record keeping are to be supported through operational policies and service workflows. Provision of smart fridges to support blood product tracking and other ICT requirements will be based on local policies and service requirements.
BBW | Bay - Blanket / Fluid Warmer | 1 | up to 24 Hours | accessed intermittently |  |  | The Bay - Blanket / Fluid Warmer provides clear area and services for a single warming cabinet. The warming cabinet may be for fluids only, blankets only or a combination of fluids and blankets. Additional Considerations: Collocation of blanket warming cabinets with linen bays is common and supports efficient workflows. Locating the bay with ready access to patient areas or operating/procedure areas is recommended. The capacity of warming cabinets varies and must be considered in relation to the area/number of patients to be served. The throughput of patients must also be considered as a higher turnover will require higher capacity/additional cabinets. Warming cabinet sizes vary and the clear area of the bay is to suit the dimensions of the equipment selected as well as ensuring adequate space for ventilation as per manufactures specifications. The bay should deep enough to ensure that the warming cabinet itself does not impede adjacent corridor circulation. Clearance to be provided in adjacent circulation space to allow for opening of the door for intermittent use and restocking. Where fluids are kept, processes for labelling fluids need to be accommodated without impeding circulation. Various cabinet configurations, including benchtop, under bench, freestanding, mobile and built-in cabinets may be selected. See manufacturer's specifications for installation requirements. Temperatures for fluids need to be monitored. A data outlet, or alternative connection to suit project/equipment requirements, may be required for connection the Building Management System (BMS) depending on provision of fluid warming.
BES | Bay - Emergency Shower | 1 | Available 24 Hours | 1 |  |  | An unobstructed and accessible area with an emergency shower and eye wash equipment for staff who have been contaminated with toxic and/or corrosive substances.
BHW | Bay - Height/ Weight | 2 | Available 24 hours | 1 patient and 1 staff, intermittently |  | Sound attenuation level medium, Acoustic privacy is required. | An open bay for measuring and recording the height and weight of patients. Standing scales and a stadiometer are typically provided. Chair scales may be used here or may be stored here to be used elsewhere in the department as required. The type of scales and provision for bariatric patients will be dependent on service requirements. Additional Design Considerations: The bay should deep enough to ensure that equipment does not impede adjacent corridor circulation. Design and location of bay should allow for intermittent use without significant disruption to circulation in adjacent corridor. Grab rails for patient stability may be provided and may be either integral to the scales or wall mounted.
BHWS-A | Bay - Handwashing, Type A | 1 | 24 Hours | 1 |  |  | The Bay - Handwashing, Type A provides handwashing facilities for staff use in areas requiring clinical handwashing prior to undertaking selected procedures and where handwashing up to the elbow may be necessary. The basin is to be a large, wall-mounted, 'scrub' type, with hands-free taps that may be wall-mounted elbow operated, foot/knee operated or operated with electronic sensor taps. Warm and cold water are to be delivered by a common spout.
BHWS-B | Bay - Handwashing, Type B | 1 | 24 Hours | 1 |  |  | The Bay - Handwashing, Type B provides handwashing facilities for staff use before and after every patient contact and visitor use as required. The basin is to be a medium, wall-mounted type, with hands-free taps that may be either wall-mounted or basin-mounted elbow operated taps. Warm and cold water are to be delivered by a common spout.
BIRM-A | Birthing Room, Without Pool | 34 | 24 Hours | 1 mother; 1-2 support persons; up to 4 staff |  | Sound attenuation level high, Acoustic privacy is required. | The Birthing Room - Without Pool provides area and amenities for the entire birthing process including mother's preparation, active birthing exercises, relaxation and analgesia during labour, birthing, infant resuscitation, postnatal recovery and observation. The room includes an active birthing/bed zone, an area to accommodate a support person, a staff observation and write up area, an infant resuscitation area, an equipment store, and a hand hygiene area with type A handwashing basin. Additional Considerations: The room will have direct access to an Ensuite - Birthing. Air should not be recirculated to other areas from the Birthing Room. Negative pressure air conditioning is required for clearance of anaesthetic gases. Clinical items such as medical gases and services as well as neonatal resuscitation equipment may be concealed to support the provision of a calm environment. Even while concealed, items are to be within easy reach of the active birthing zone. The availability of proprietary items for in-room infant resuscitation may change and this can inform the use of built in or mobile equipment. Overall dimensions and configuration of storage for the resuscitaire equipment are to be evaluated based on the selection of equipment and it is recommended that a test-fit is conducted, ensuring appropriate clearances are provided for emergency procedures to be undertaken, in order to confirm room layout requirements. Equipment storage to be configured to suit service requirements including models for sharing equipment between multiple rooms. Storage requirements within the room will vary between facilities depending on models of centralising or decanting and restocking consumables. Configuration and extent of storage to be determined based on service requirements. Various supportive fixtures for active birthing are to be considered, for example, a wall mounted ladder, bench or mantlepiece for support and leaning, or ceiling mounted hooks for supportive fabrics to be hung. Consider mounting heights and providing support for various postures at multiple heights. A mother's entire weight may be supported on this item, therefore the strength of the item and fixings, including in-wall support, to be equivalent to heavy duty grab rails. A fixed workstation for staff write-up and observation has been shown. Sufficient space for dual monitors may be provided as required to suit operational, EMR and ICT requirements. A workstation on wheels is also acceptable. An examination light is required for suturing and small procedures. It is preferred that the light be as unobtrusive as possible and be able to be pushed away so that it is not dominating the area. Lighting options to be provided, including dimmable room lighting, reading lights and night lights to provide flexible lighting across the space. Coloured lighting and calming light artworks may be considered by projects. Systems, including portable items, to facilitate the selection and control of music by the mother are to be considered. Acoustics treatment may be required to support this without disruption to adjacent rooms. Provision of a telephone for patient use will be dependent on the operational model for transferring calls. Natural light and outlook essential. Window sill to be low enough to see view while laying in bed. Treatment, to windows to allow blinds to be kept open as desired but privacy to be maintained, may need to be considered depending on overlook from nearby buildings.
BIRM-B | Birthing Room, With Pool | 40.5 | 24 Hours | 1 mother; 1-2 support persons; up to 4 staff |  | Sound attenuation level high, Acoustic privacy is required. | The Birthing Room - With Pool provides area and amenities for the entire birthing process including mother's preparation, active birthing exercises, water immersion during labour, relaxation and analgesia during labour, birthing, infant resuscitation, postnatal recovery and observation. The room includes an active birthing/bed zone, birthing pool an area to accommodate a support person, a staff observation and write up area, an infant resuscitation area, an equipment store, and a hand hygiene area with type A handwashing basin. Birthing Pool: The use of birthing pools should be based on local policies and protocols for water immersion during labour and birth. Birthing pools can vary significantly in shape and the pool shown on the RLS is indicative only. Final selection may have implications for required area and clearances. The birthing pool must be deep and wide enough for effective water immersion for pain relief during labour and birth. Ergonomic assessment of birthing pool options should be undertaken prior to selection, including the height of the pool for mothers getting into and out of the pool, as well as access for staff assisting the mother from the side of the pool. The location of grabrails and supports, including those integral with the bath design are to be considered to support transfers without impeding access to the mother by staff. Pools may be installed with steps built around them to provide easy access into the pool from a seated position. Provision of stairs is to be assessed to ensure they do not limit access with floatation transfer devices in the event of an emergency. Handrails to be provided to stairs to reduce risk of slips in wet area around the pool. A ceiling mounted hoist has not been included as there are better and more timely ways to remove a mother from the birthing pool should her condition deteriorate. Consideration for ready access to floatation devices for transfer from the pool is required. The tapware and drain for the birthing pool are to be carefully considered. Taps need to allow the bath to be filled rapidly. Water also needs to be able to be drained quickly, preferably without needing to reach into the bath after birthing. Additional Considerations: The room will have direct access to an Ensuite - Birthing. Air should not be recirculated to other areas from the Birthing Room. Negative pressure air conditioning is required for clearance of anaesthetic gases. Clinical items such as medical gases and services as well as neonatal resuscitation equipment may be concealed to support the provision of a calm environment. Even while concealed, items are to be within easy reach of the active birthing zone. The availability of proprietary items for in-room infant resuscitation may change and this can inform the use of built in or mobile equipment. Overall dimensions and configuration of storage for the resuscitaire equipment are to be evaluated based on the selection of equipment and it is recommended that a test-fit is conducted, ensuring appropriate clearances are provided for emergency procedures to be undertaken, in order to confirm room layout requirements. Equipment storage to be configured to suit service requirements including models for sharing equipment between multiple rooms. Storage requirements within the room will vary between facilities depending on models of centralising or decanting and restocking consumables. Configuration and extent of storage to be determined based on service requirements. Various supportive fixtures for active birthing are to be considered, for example, a wall mounted ladder, bench or mantlepiece for support and leaning, or ceiling mounted hooks for supportive fabrics to be hung. Consider mounting heights and providing support for various postures at multiple heights. A mother's entire weight may be supported on this item, therefore the strength of the item and fixings, including in-wall support, to be equivalent to heavy duty grab rails. A fixed workstation for staff write-up and observation has been shown. Sufficient space for dual monitors may be provided as required to suit operational, EMR and ICT requirements. A workstation on wheels is also acceptable. An examination light is required for suturing and small procedures. It is preferred that the light be as unobtrusive as possible and be able to be pushed away so that it is not dominating the area. Lighting options to be provided, including dimmable room lighting, reading lights and night lights to provide flexible lighting across the space. Coloured lighting and calming light artworks may be considered by projects. Systems, including portable items, to facilitate the selection and control of music by the mother are to be considered. Acoustics treatment may be required to support this without disruption to adjacent rooms. Provision of a telephone for patient use will be dependent on the operational model for transferring calls. Natural light and outlook essential. Window sill to be low enough to see view while laying in bed. Treatment, to windows to allow blinds to be kept open as desired but privacy to be maintained, may need to be considered depending on overlook from nearby buildings.
BLIN | Bay - Linen | 2 | Up to 24 Hours | accessed intermittently |  |  | The Bay - Linen is a recessed bay to accommodate a clean linen supply or exchange trolley. Additional Considerations: The bay may be enclosed with hinged doors, sliding doors or a roller shutter where items need to be secured or enclosed. An open bay is acceptable if a covered trolley is supplied (as required by local infection control policies). Where lockable, hinged doors are provided, hold open hardware may be required to support easy movement of the trolley in and out of the bay. Door protection to suit movement of beds and equipment in adjacent area. Where a roller shutter is provided, the weight of the shutter, manual handling requirements and minimising the noise of the shutter when opening are to be considered. If bay is enclosed, a light switch or sensor may be required to provide adequate lighting when linen stock is accessed. Storage for additional items such as pillows, shower curtains, privacy curtains, etc. may be provided as either joinery, a modular shelving system, a wall mounted panel with wire baskets or a mobile trolley. The storage solution will be dependent on items to be stored and if items need to be secured or covered due to infection control requirements. Inclusion of storage should consider size and manoeuvrability of the linen trolley to be used to ensure adequate clearance. Linen trolleys can vary in height and often require a cover that is lifted from the top. In this instance, shelving above the trolley is not recommended. If the trolley is of a suitable height and a shelf is provided, it is to be mounted no higher than 1800AFFL and the reach over trolley is to be considered. Hooks may be provided for clean linen bags and may be differentiated for clinical, general and patient's personal linen. Linen bags on hooks should be able to be accessed without moving the trolley. Bay - Linen may sometimes be combined with Bay - Blanket / Fluid Warmer (see Standard Component BBW for additional space, FF&E and services requirements).
BMEQ | Bay - Mobile Equipment | 4 | Available 24 hours | accessed intermittently |  |  | An open bay for storage and recharging frequently used mobile equipment. Equipment may include wheelchairs, mobile scales, commode chairs, patient lifting devices or observation/diagnostic equipment. Additional Design Considerations: The bay should deep enough to ensure that equipment does not impede adjacent corridor circulation; a depth range from 800mm to 1500mm is recommended. Consider the manoeuverability of equipment to be stored as some items may have large turning circles and will need to be parked perpendicular, requiring a deeper bay. Power outlets may be located on the sides of deeper bays to be accessible when equipment is parked. If portable imaging equipment is to be stored additional area may be required. Data outlets may be required depending on equipment. Equipment may be recharged while stored in this space; provision of appropriately located power outlets is required. Items may be hung in this bay, e.g. a patient slide, and appropriate in-wall supports are required for mounting hooks.
BMFD-3 | Bay - Multifunction Device, 3m2 | 3 | Available 24 Hours | 1 staff intermittently |  |  | An open bay for the provision of printing, scanning and photocopying facilities and storage of stationery. Additional Considerations: Adequate ventilation to be provided over multifunction device. Adequate clearance is required around the multifunction device to allow access to easily open panels (e.g. for clearing paper jams, changing ink, restocking paper, etc.). Clearance requirements will be dependent on device specifications. Clearance to be provided in adjacent circulation space to allow for intermittent use and restocking.
BMFD-7 | Bay - Multifunction Device, 7m2 | 7 | Available 24 Hours | 1-2 staff intermittently |  |  | An open bay for the provision of printing, scanning and photocopying facilities and storage of stationery. Additional Considerations: Adequate ventilation to be provided over multifunction device. Adequate clearance is required around the multifunction device to allow access to easily open panels (e.g. for clearing paper jams, changing ink, restocking paper, etc.). Clearance requirements will be dependent on device specifications. Clearance to be provided in adjacent circulation space to allow for intermittent use and restocking.
BMT | Bay - Meal Trolley | 4 | Up to 24 Hours | accessed intermittently |  |  | The Bay - Meal Trolley is a recessed bay to accommodate meal trolleys during meal delivery and collection times. Additional Considerations: The size and type of meal trolleys to be accommodated will be dependent on service requirements and food service models (re-thermalisation, cook fresh, cook chill, mid-meals, etc.) and will need to be confirmed to inform spatial and services requirements. Consider manoeuvrability of the meal trolley to be used to ensure adequate clearance. Most meal trolleys have 2 x fixed and 2 x swivel wheels which limit movement to position in the bay. Bay may be designed without one end wall to improve trolley access and reduce WHS risks from push/pull movements with certain types of trolleys. The Bay should be deep enough to ensure trolleys do not impede corridor circulation; minimum recommended depth is 1200mm, however, the depth may need to increase to 1500mm depending on the trolley selected. Special power may be required for re-thermalisation trolleys.
BPATH | Bay - Pathology Point of Care Testing (POCT) | 3 | Available 24 Hours | 1-2 staff, intermittently |  |  | The Bay - Pathology Point of Care Testing (POCT) provides the area and facilities for the preparation and testing of specimens close to patient care areas allowing for rapid results to improve clinical decision making and patient care. Two benchtop POCT devices have been indicated, with additional devices requiring added space and services. The devices shown are placeholders as there are many POCT devices available and various combinations of devices may be provided depending on the service requirements of the unit in which this bay is located. Additional Considerations: Ready access to handwashing bay is recommended. Ready access may be required to a pneumatic tube bay that is in a staff supervised area. If providing wall mounted joinery over the bench consider clearance heights for benchtop machines, including for futureproofing. Weight requirements for machines should be considered to ensure adequate support is provided to the bench. POCT devices should not to be located in direct sunlight as they can overheat. Refrigerator temperature monitoring may be required and is to be connected to the building management system (BMS). The alarm is to provide alerts where 24-hour monitoring is undertaken, within the unit and/or remotely. Type and extent of storage to suit testing requirements, including ambient temperature storage for cartridges. Ensure adequate lighting over workspace. Power requirements (type/quantity) to be considered to suit equipment and service requirements.
BPROP | Bay - Property, Staff | 2 | Available 24 Hours | Up to 4 staff, intermittently |  | Sound attenuation level medium, Acoustic privacy is required. | A bay containing lockers for the secure storage of property including clothing, bags and personal effects. The number of lockers provided will be dependent on staff complement per shift. Additional Design Considerations: Approximately 1m2 will be required for each additional 3 x tiers of lockers. Each tier may provide up to 4 lockers. Lockers may be provided as freestanding items or built-in joinery. A proportion (minimum 2%) of lockers may be required to be accessible with locking devices located between 900mm to 1100mm AFFL.
BPTS | Bay - Pneumatic Tube Station | 1 | Available 24 Hours | accessed by staff intermittently |  |  | The Bay - Pneumatic Tube Station (PTS) provides an area for dispatching and receiving of items via the pneumatic tube system. This Standard Component represents a department based bay and additional area, FFE and services will be required for a central receiving and processing area in a pathology or pharmacy department. Additional Considerations: To be collocated or located with clear oversite from a staff station or within a staff only area. Arrival notification alarms must be programmable for different times and scenarios depending on service requirements. Security requirements are to be considered. If the station is to receive deliveries from Pharmacy or is in an area accessed by the public then secure access to a lockable receiving basket may be required depending on location of the bay. Station may be a pass through or end station and receive/dispatch through the floor or ceiling depending on location and system configuration. Radio frequency identification (RFID) technology may be considered by projects to reduce errors, improve traceability and to enable dynamic/intuitive routing to reduce "traffic jams". This technology involves embedded RFID tags in cannisters that are read by receivers at tube stations on dispatch/arrival and at interchange points in the system. Where multiple services send/receive products via the pneumatic tube system, e.g. pathology and pharmacy, colour coding the cannisters for the different streams is recommended. Colour coding to differentiate between cannisters from high priority areas, e.g. emergency and intensive care vs general wards may also be considered.
BRES | Bay - Resuscitation Trolley | 1.5 | Available 24 hours | accessed intermittently |  |  | A open bay to provide supervised holding of the resuscitation trolley and associated equipment. The bay should be located close to a Staff Station. Rapid access to the trolley and from this bay to patient areas is essential. Additional Design Considerations: Requirements for specialised resuscitation trolleys should be considered (e.g. specific trolleys are required for paediatric and neonatal departments and non-magnetic trolleys are required for some medical imaging modalities). If a department requires more than one type of resuscitation trolley to be accommodated in their bay (e.g. both adult and paediatric trolleys) additional area will be required. Additional equipment, such as ECG machines, may be accommodated in this bay in some departments. The bay shall be a minimum depth of 800mm to ensure that equipment does not impede adjacent corridor circulation. A patient slide may need to be readily available if the resuscitation trolley is required and may be provided within this bay.
BVM | Bay - Vending Machine | 2 | Available 24 Hours | 1 - 2 intermittently |  |  | The Bay - Vending Machine provides space and services for vending machines in public and staff areas. The area may be in a recessed bay along a corridor or in an open area and should be located to ensure adjacent circulation paths are not impeded with intermittent use. Each additional vending machine will require approximately 1.5m2 of additional area as well as associated services. Additional Considerations: Sufficient clearance is required around the machines for manoeuvring in the case of removal for maintenance, repair or replacement. Bulkheads and joinery surrounds may be provided to prevent users from tipping or shaking machines but should be designed to allow for access for repairs and maintenance (e.g. removable). Alternative types of vending machines such as for toiletries, phone charging accessories, etc. that may be useful to patients, visitors, and staff, e.g. for unexpected overnight stays, and staff at end of trip areas, may also be considered. Hand hygiene facilities, such as alcohol-based hand rub, should be provided nearby in line with local infection control policies.
BWC | Bay - Wheelchair Park | 4 | Available 24 hours |  |  |  | An open bay for storing wheelchairs. Typically located in close proximity to the main entrance or department entry. Additional Design Considerations: The bay should be deep enough to ensure that wheelchairs do not impede adjacent circulation space; a minimum depth of 1300mm is recommended. The ability to secure wheelchairs (e.g. with locks or by enclosing the bay) to be considered as required.
BWTR | Bay - Water Fountain | 1 | Available 24 Hours | 1 intermittently |  |  | An area, typically associated with a main entry lobby or waiting area, for a drinking water dispenser. The water dispenser may be in a recessed bay or in an open area and should be located to ensure adjacent circulation paths are not impeded with intermittent use. Additional Considerations: A water dispenser with a spout to fill disposable cups or refill drink bottles is recommended. Hand hygiene facilities, such as alcohol-based hand rub, should be provided nearby in line with local infection control policies.
CHPT | Change Cubicle - Patient | 2 | up to 24 Hours | 1 patient |  |  | The Change Cubicle - Patient provides area and amenities for ambulant patients to change into a hospital gown or appropriate clothes prior to physical therapy, examination, diagnostic imaging, treatment, procedures or surgery, and to re-dress afterwards. The cubicle should provide facilities for temporarily hanging clothes and a bench seat or chair for the patient to sit on whilst changing. Additional Considerations: Provision of a door, cubicle partition system, or privacy curtain to enclose the change cubicle is to be determined to suit patient cohort and service requirements. Where possible, the cubicle should be located so the entry opening is discreet and not facing into high traffic areas. Where door is provided it must swing outward. For door to swing inward, additional area for circulation and emergency release hardware is required. In both cases it is required that the door can be opened from the outside in an emergency. The selection of door hardware, namely handles and locks/privacy latches, must ensure usability by people with limited dexterity and motor skills. Lighting control may be provided external to the change cubicle to serve a number of change cubicles in an area, particularly where cubicle systems or privacy curtains are provided. Enclosed change cubicles may have individual light switches or motion controlled lighting. Where two-way access is provided, e.g. from a corridor and directly from a treatment or imaging room, additional area may be required to provide appropriate circulation at the second entry. Where adjacent or with direct access to an imaging room or treatment room where imaging devices are used, the change room must be adequately shielded. Radiation shielding is to be as advised and confirmed by a Radiation Consultant. Where paediatric services are provided, additional area within change cubicles will be required to support parents/carers assisting children to change. The provision of multiple benches with different seating heights (450mm-520mm for adults and 350mm for children) may also be considered to support paediatric patients. Operational models for handling patient belongings are to be considered. Close proximity to patient lockers or other secure property storage may be required where belongings are not to be kept with the patient. A chair may be provided in lieu of a bench seat depending on patient cohort and service requirements. The depth of the seat for both chair and bench seat options is to be adequate to support movement while changing, and the weight rating of all seating, including provision for bariatric capacity, is to be considered. The colour of seating is to contrast with the flooring and walls to support visibility and reduce the risk of falls. An operational model for responding to calls from staff/patients will be required to support the provision of the patient to staff call button, staff assist call button and emergency call button, and inclusion is to be determined based on service requirements. Additional staff assist and emergency call buttons my be located outside of the cubicle and shared between a number of change rooms depending on service requirements.
CHPT-AC | Change Cubicle - Patient, Accessible | 4.5 | up to 24 Hours | 1 patient, 1 staff / carer assisting as required |  |  | The Change Cubicle - Patient, Accessible provides area and amenities for patients using wheelchairs, walking aids and/or requiring assistance to change into a hospital gown or appropriate clothes prior to physical therapy, examination, diagnostic imaging, treatment, procedures or surgery, and to re-dress afterwards. The cubicle should provide facilities for temporarily hanging clothes and a bench seat or chair for the patient to sit on whilst changing. Additional Considerations: Provision of a door, cubicle partition system, or privacy curtain to enclose the change cubicle is to be determined to suit patient cohort and service requirements. Where possible, the cubicle should be located so the entry opening is discreet and not facing into high traffic areas. Where door is provided it must swing outward. For door to swing inward, additional area for circulation and emergency release hardware is required. In both cases it is required that the door can be opened from the outside in an emergency. The selection of door hardware, namely handles and locks/privacy latches, must ensure usability by people with limited dexterity and motor skills. Lighting control may be provided external to the change cubicle to serve a number of change cubicles in an area, particularly where cubicle systems or privacy curtains are provided. Enclosed change cubicles may have individual light switches or motion controlled lighting. Where two-way access is provided, e.g. from a corridor and directly from a treatment or imaging room, additional area may be required to provide appropriate circulation at the second entry. Where adjacent or with direct access to an imaging room or treatment room where imaging devices are used, the change room must be adequately shielded. Radiation shielding is to be as advised and confirmed by a Radiation Consultant. Where paediatric services are provided, additional area within change cubicles will be required to support parents/carers assisting children to change. The provision of multiple benches with different seating heights (450mm-520mm for adults and 350mm for children) may also be considered to support paediatric patients. Operational models for handling patient belongings are to be considered. Close proximity to patient lockers or other secure property storage may be required where belongings are not to be kept with the patient. The depth of the seat is to be adequate to support movement while changing, and the weight rating of all seating, including provision for bariatric capacity, is to be considered. The colour of seating is to contrast with the flooring and walls to support visibility and reduce the risk of falls. An operational model for responding to calls from staff/patients will be required to support the provision of the patient to staff call button, staff assist call button and emergency call button, and inclusion is to be determined based on service requirements. Additional staff assist and emergency call buttons my be located outside of the cubicle and shared between a number of change rooms depending on service requirements.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | 12 | 24 Hours | Up to 3 |  | Sound attenuation level - medium; acoustic privacy required. | The Patient Change areas are provided for patients to change into gowns, to store their street clothing and to perform personal ablutions. Change Rooms should be divided into Male and Female areas on a proportional basis and will include Change Cubicle, Toilet, handbasin and lockers for secure storage of personal property.
CHST-10 | Change - Staff, 10m2 | 10 | Available 24 Hours | 1-3 staff, intermittently |  | Sound attenuation level high, Acoustic privacy is required. | A room for staff to change into appropriate work clothing or gowns, to store their street clothing and to perform personal ablutions. Change rooms are typically allocated as separate male and female areas on a proportional basis. The room will be comprised of showers, toilets, basins and lockers, with numbers to suit service and staff requirements. Secure storage for personal property may be provided dependent on operational policies and service requirements. Additional Design Considerations: Accessible change facilities for staff to be provided nearby. A proportion (minimum 2%) of lockers may be required to be accessible with locking devices located between 900mm to 1100mm AFFL. Secure, staff-only access to change rooms is required. Visual privacy from the entry door to the changing areas and cubicles is to be considered during planning. If provided for staff to change for sterile procedures, access to clean linen will be required
CHST-35 | Change - Staff, 35m2 | 35 | Available 24 Hours | Up to 16 staff, intermittently |  | Sound attenuation level high, Acoustic privacy is required. | A room for staff to change into appropriate work clothing or gowns, to store their street clothing and to perform personal ablutions. Change rooms are typically allocated as separate male and female areas on a proportional basis. The room will be comprised of showers, toilets, basins and lockers, with numbers to suit service and staff requirements. Secure storage for personal property may be provided dependent on operational policies and service requirements. Additional Design Considerations: Accessible change facilities for staff to be provided nearby. A proportion (minimum 2%) of lockers may be required to be accessible with locking devices located between 900mm to 1100mm AFFL. Secure, staff-only access to change rooms is required. Visual privacy from the entry door to the changing areas and cubicles is to be considered during planning.
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | 60 | Available 24 Hours | 1 patient; up to 6 staff |  | Acoustic privacy is required | A highly specialised procedure room where interventional Cardiac Catheter procedures are undertaken in an aseptic environment and may include administration of sedation or general anaesthesia. The Catheter Laboratory - Electrophysiology Studies may also be used for electrophysiology (EP) studies and treatment and is typically only provided in hospitals of tertiary level or higher, where required. A Control room will be located adjacent and may serve more than one procedure room. A clinical scrub area should be located immediately adjacent to the staff entry. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. A strut channel system able to carry moving loads may be required. Cable ducts from Control Room and Equipment room will be required. 'X-Ray in Use' lights are required outside each entry door. Power to imaging equipment should be on emergency supply. The Catheter Laboratory for EP studies will require radiofrequency shielding.
CLAB-I | Catheter Laboratory - Interventional | 55.5 | Available 24 Hours | 1 patient; up to 6 staff |  | Acoustic privacy is required | A procedure room where interventional Cardiac Catheter procedures are undertaken in a controlled, clean environment and may include administration of sedation or, much less commonly, general anaesthesia. A Control room will be located adjacent and may serve more than one procedure room. A clinical scrub area should be located immediately adjacent to the staff entry. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. A strut channel system able to carry moving loads may be required. Cable ducts from Control Room and Equipment room will be required. 'X-Ray in Use' lights are required outside each entry door. Power to imaging equipment should be on emergency supply.
CLCR-EP | Catheter Laboratory EP Control Room | 17 | Available 24 Hours | Up to 4 staff |  | Acoustic privacy is required | Room adjacent to Catheter Lab EP Procedure Room/s with direct visibility to the patient for remote control of equipment, and review and reporting of images. The Control Room may serve two procedure rooms. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Cable ducts from Catheter Lab Procedure Room and Computer/Equipment room will be required.
CLCR-I | Catheter Laboratory Interventional Control Room | 14 | Available 24 Hours | Up to 3 staff |  | Acoustic privacy is required | Room adjacent to Catheter Lab Interventional Procedure Room/s with direct visibility to the patient for remote control of equipment, and review and reporting of images. The Control Room may serve two procedure rooms. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Cable ducts from Catheter Lab Procedure Room and Equipment room will be required.
CLN-10 | Clean Store, 10m2 | 10 | 24 Hours | 1 - 2 staff intermittently |  |  | The Clean Store, 10m² provides space for the storage of clean consumables and equipment for patient care.
CLN-MED-20 | Clean Store / Medication Room, 20m2 | 20 | 24 Hours | 1 - 2 staff intermittently |  |  | The Clean Store / Medication Room, 20m² provides a space for the storage of clean consumables and the secure storage of medications, including intravenous fluids. The room may provide storage for drugs of dependence in accordance with relevant legislation. Controlled temperature and humidity is required for medication storage and internal temperatures should not rise above 25°C. Secured staff only access to the room is required. The provision and placement of CCTV cameras is dependent on jurisdictional policies. Refer to project specific requirements for the provision of automated medication management systems.
CLN-MED-S | Clean Store / Medication Room, Sub | 8 | 24 Hours | 1 - 2 staff intermittently |  |  | The Clean Store / Medication Room, Sub provides a small space for the storage of clean consumables and the secure storage of medications, including intravenous fluids. The room may provide storage for drugs of dependence in accordance with relevant legislation. Controlled temperature and humidity is required for medication storage and internal temperatures should not rise above 25°C. Secured staff only access to the room is required. The provision and placement of CCTV cameras is dependent on jurisdictional policies. Refer to project specific requirements for the provision of automated medication management systems.
CLRM | Cleaner's Room | 5 | 24 Hours | 1-2 staff intermittently |  |  | The Cleaner's Room is a secure room for the storage and decanting of cleaning materials, washing and storage of cleaning equipment and cleaner's trolley, and for liquid waste disposal. This room may be shared between multiple units depending on service requirements. Additional Considerations: Size requirements for a Cleaner's Room will be dependent on the department to be serviced and operational models for staffing and cleaning regimens that may dictate the type of equipment and number of trolleys to be accommodated. Where more than one trolley will be stored an additional 1.5m2 is required per trolley. The requirements to accommodate vacuum cleaners, floor scrubbers or other floor cleaning equipment within the Cleaner&#8217;s Room is to be determined at a project/department level in conjunction with assessment of finishes and cleaning requirements. Where operational models do not provide a centralised area for the storage of large cleaning equipment (e.g. floor scrubbers), additional area may be required in units or in shared areas between units depending on the additional equipment to be stored. Considerations regarding detergent/chemical management system to be confirmed at a project level, including water connection and associated services requirements (including back flow prevention). Extent and type of storage for chemicals and consumables is to suit service requirements. Consideration of the weight of chemical bottles is required when determining appropriate storage solutions (i.e. if storing chemicals in bottles over 1 litre then low-height shelving is required to avoid lifting of heavy items from high shelves). Hold open door hardware may be provided for hinged doors to allow for easy movement of trolleys in and out of the room. Controlled, staff-only access may be required if this room is provided in a public or patient care area.
CLUP-7 | Clean-Up Room, 7m2 | 7 | 24 Hours | 1 - 2 |  |  | The Clean-Up Room is used for holding of used trolleys and articles from an Operating Room, Procedure Room or Catheter Laboratory. Items may be sorted, rinsed or disposed of and despatched to waste holding/sterilising areas as appropriate.
COMM | Communications Room | 15 | 24 Hours | 1 - 2 intermittently |  |  | A secure room for engineering services to accommodate racks, backbone cabling and interface control panels for information and communication systems. Systems may include master antenna television (MATV), internet protocol television (IPTV), nurse call, security, telecommunications, building management system (BMS) and Fire/EWIS systems. Additional Considerations: The area shown is indicative; the size of the Communications Room will vary according to engineering consultant advice based on systems to be accommodated. Head end units for security, nurse call, BMS, TMV, IPTV/MATV to be accommodated as required. Equipment located in this room will require uninterruptible power fed from the essential supply (generator). Dedicated air conditioning that is monitored by the Building Management System (BMS) is required Provide adequate lighting to ensure visibility to both front and rear of racks. The configuration of power and data (i.e., wall mounted or ceiling mounted services) to supply the racks is to be determine base on the room layout and equipment selected. The Communication Room must be located such that it is not susceptible to possible water damage or excessive electromagnetic interference. Water or gas pipes must not pass through a Communication Room. Where a Distributed Antenna System (DAS) for in-building mobile reception is provided, the telco carrier base station equipment is to be located in the Communications Room. Refer to telco for spatial requirements to determine sizing of the Communications Room. A DAS shall conform to the most recent edition of the Mobile Carriers Forum Design Specification for Distributed Antenna Systems.
CONS | Consult Room | 12 | 24 Hours | 1 patient: 1-2 staff, 1-2 visitors |  | Sound attenuation level high, Acoustic privacy is essential. | The Consult Room will provide for private consultation and physical examination of patients with or without support persons present. A second door may be required for staff safety and security; the location of the second door will be dependent on the availability an appropriate area to egress. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements.
CONS-UN | Consult Room - Universal Access | 17 | 24 Hours | 1 patient: 1-2 staff, 1-2 visitors |  | Sound attenuation level high, Acoustic privacy is essential. | The Consult Room - Universal Access will provide for private consultation and physical examination of a range of patients with or without support persons present. The spatial allowance will accommodate patients who (a) are using a wheelchair, including a bariatric sized wheelchair; or (b) require a hoist to transfer. The position of the basin, dispensers and other fittings also allows for use by staff in wheelchairs. A second door may be required for staff safety and security; the location of the second door will be dependent on the availability an appropriate area to egress. It is anticipated that one of these rooms may be provided within a 'pod' of consult rooms. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements.
CORR | Corridor - Patient Treatment Areas | 0 | 24 hours |  |  |  | The Corridor - Patient Treatment Areas indicates requirements for circulation spaces between patient treatment spaces, patient accommodation and clinical support spaces within departments. Corridor widths will vary between departments and will depend on the requirement for frequent bed transportation; widths must allow for the safe movement and passing of trolleys, beds, wheelchairs and other mobile equipment as required. Provision should be made for situations where oversized additional equipment, e.g. bed extensions or bariatric beds, are in use. Guidance for the calculation of areas (m2) for corridors is provided for in Health Planning Unit Schedules of Accommodation as circulation percentages. Refer to AusHFG Part C: Design for Access, Mobility, Safety and Security for further information. Additional Design Considerations: Wall protection is recommended to protect wall linings and finishes from damage from impact. The height of the wall protection will be dependent on equipment frequently moved through the space; a height range of 900mm minimum up to 1500mm is recommended depending on service requirements. Handrails are to be provided along not less than one side of all corridors in patient care areas to provide assistance and support for patients and visitors. Handrails are to be continuous along the corridors wherever practical. Cleaner's electrical outlets to be provided at regular intervals as per mandatory requirements. Corridor lighting, where patients are transported on beds/trolleys, should be positioned to prevent unnecessary glare to patient eyes. Light fittings shall be recessed, flush fitting and sealed to prevent dust ingress. Manual controls to power-operated doors shall be located on the continuous accessible path of travel no closer than 500 mm from an internal corner and between 1000 mm to 2000 mm from the hinged door leaf in any position as per the requirements of AS1428.1.
CTPR | CT Planning Room | 45 | 8am to 6pm, Monday – Friday (extended hours, weekends and/or emergency access may be required) | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | The CT Planning Room provides space and facilities for Computed Tomography (CT) scanning procedures as part of the planning of radiation therapy treatment for patients. A collocated Control Room with observation window is required for the remote operation of equipment. An Equipment Room may be required and will need to be in close proximity due to acceptable cable lengths. Cable ducts between the Scanning Room, Control Room and Equipment Room will be required. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access. The CT Planning Room must be adequately shielded. Radiation shielding is to be as advised and confirmed by a Radiation Consultant to suit the equipment procured. 'X-ray in Use' lights are required outside each entry door to be illuminated when scanning begins. Interlock system to doors recommended to prevent machine from running if door open and to stop the machine upon entry if in use. Additional Considerations: Additional medical gases may be provided (including additional oxygen, additional suction, nitrous oxide and scavenge) depending on clinical requirements, such as the provision of anaesthetic services. Dimmable lights are to be provided and lights are to be placed to avoid glare to patients. Back lit ceiling artwork may be installed for patient comfort and distraction. Where possible, it is preferred that ductwork in the ceiling is designed to the perimeter of room to allow for future flexibility to provide ceiling mounted equipment. The need for and extent of emergency and UPS power must be reviewed to suit site specific requirements and confirmed based on risk assessment considering the impact of a power outage on patient care/safety. Unistrut ceiling able to carry moving loads may be required. A mix of respiratory gating, surface guidance, motion management and/or patient position technologies may be provided depending on clinical/service requirements and will be coordinated to support the technologies provided in the Linear Accelerator Treatment Room. The provision of these technologies is to be confirmed by the project and may have additional design requirements including additional power/data, water cooling, heat load, possible increased set down, increased number of trenches/conduits, space requirements for computer cabinets in control room and additional penetrations through the suspended ceiling. Requirements to be confirmed with vendors.
DEN-MLB | Dental Laboratory, Minor | 12 | 8am to 6pm, daily | 1-2 staff |  | Sound attenuation high: Acoustic privacy required | The Minor Dental Laboratory provides facilities and equipment for dental clinicians/technicians to make adjustments to dentures or other dental prosthetics, and to store prosthetic supplies. Additional Design Considerations: Dust extraction is required at the source and may be integral to equipment in some instances. If Bunsen burner to be used in this room, gas outlet or portable gas is required and overhead joinery may need to be reviewed to provide appropriate clearance. A workstation with computer and barcode scanner may be provided for digital tracking and records.
DENSR-1 | Dental Surgery, Type 1 | 14.5 | Available 24 Hours | 1 patient, 1-2 staff, 1-2 visitors |  | Sound attenuation level high, Acoustic privacy is required. | The Dental Surgery provides for dental consultation, examination and treatment for patients. This Standard Component provides a configuration for switching between a a right-handed and left-handed set-up and should not be mirrored. Dental surgery layouts will vary and will be dependent on existing/proposed workflows and the equipment procured. Additional Design Considerations: Services will be required for the dental chair unit; these will be dependent on the chair model and supplier and include dental air, dental suction, filtered water, power and drainage. Facilities for viewing physical x-ray films brought in by patients may be provided in a central location as required. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. Radiation shielding to be advised by Radiation Consultant. An external outlook is desirable.
DENSR-2 | Dental Surgery, Type 2 | 14.5 | Available 24 Hours | 1 patient, 1-2 staff, 1-2 visitors |  | Sound attenuation level high, Acoustic privacy is required. | The Dental Surgery provides for dental consultation, examination and treatment for patients. This Standard Component shows a right-handed configuration and should not be mirrored. The layout may be adjusted to provide an ambidextrous set up. Dental surgery layouts will vary and will be dependent on existing/proposed workflows and the equipment procured. Additional Design Considerations: Services will be required for the dental chair unit; these will be dependent on the chair model and supplier and include dental air, dental suction, filtered water, power and drainage. Facilities for viewing physical x-ray films brought in by patients may be provided in a central location as required. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. Radiation shielding to be advised by Radiation Consultant. An external outlook is desirable.
DINR | Dining - Patient | 40 | Up to 24 Hours | 16 patients seated, 3-4 staff intermittently |  | Sound attenuation level high; acoustic privacy is required. | The Dining Room is an area provided for patients to eat their meals away from the bedroom environment in an inpatient setting. The Dining Room may be used for other patient recreational activities when not in use for meals. In this instance, additional transfer of noise to adjacent patient care areas to be considered. Additional Considerations: Food services model, e.g. meal trolley (tray or re-thermalisation), servery, beverage bay, etc. will impact room design. A kitchen, servery or beverage bay may be collocated or combined with the Dining Room. Size of space will depend on number of beds and models of care; designed area based on 3m2 per person. Access will be required for patients with limited mobility including wheelchairs, walking frames and other walking aids. Design must also consider the needs of patients with dementia and frail/elderly. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. Patient to nurse call button and handset is optional and may be provided in units where high patient observation is required. Duress provision will be dependent on service requirements. Furniture shown in this Standard Component is indicative and items selected will vary depending on patient cohort. Considerations regarding furniture include assessing the weight rating of all seating, including for bariatric capacity as required, providing a mix of chairs with and without arms to support patient mobility, and providing built-in furniture or securing furniture where there may be safety/security concerns. Lighting to be consistent across the space. Size, type and quantity of waste bins is dependent on environmental management policies and service requirements. Where a facility has additional waste streams, beyond the general waste and comingled recycling indicatively shown in this Standard Component, additional area may be required to accommodate the receptacles. Depending on frequency of waste removal, bin sizes may be larger in high use areas or as guided by local waste management policies. Joinery to conceal bins may be provided depending on infection prevention and control policies. Where possible a achieving a domestic appearance with furniture, fittings and finishes is recommended. The display of artwork and cultural assets will be guided by the facility. A rail display system may be considered to minimise risks and allow for flexibility in display. Ease of cleaning for infection prevention and control also needs to be considered when selecting artwork. Natural light and external outlook is essential.
DISP-10 | Disposal Room, 10m2 | 10 | Up to 24 Hours | 1-2 staff, accessed intermittently |  |  | The Disposal Room provides for the temporary holding of waste, separated into various waste streams (general waste, clinical waste, cardboard recycling, co-mingled recycling, cytotoxic waste, confidential waste, sharps, soiled linen, etc.) prior to removal to a designated waste zone in the loading dock. Waste will be collected at regular times each day, or when requested, to prevent build-up of waste in the unit. Additional Considerations: Size requirements for a Disposal Room will be dependent on a department's estimated waste output, the frequency of waste collection and local operational policies for waste management that may dictate the number of waste streams and minimum bin sizes. Adequate space must be provided to allow for segregation of waste streams. The use of 660 L bins should be carefully considered due to the WHS risks associated with manual handling, however it is acknowledged in some areas with significant waste output they may be required. Where the weight of full waste receptacles is a concern, the use of tugs/ACVs is often required. Waste transfer may be manual, by motorised movers/tugs or via automated guided vehicle (AGV) systems. Where tugs and AGVs are being considered, it is important to confirm clearances, door widths and turning circles within the room, in the adjacent corridor and along path of travel to designated waste zone. Hold open door hardware may be required to allow for easy movement of bins and trolleys in and out of the room. Shelf may be provided for storage of small sharps bins prior to collection. If a handwashing basin is not provided within the Disposal Room, one should be located in close proximity. The location of hand hygiene facilities may be dependent on local infection control policies. Staff only access may be required if this room is provided in a public or patient care area.
DISP-15 | Disposal Room, 15m2 | 15 | Up to 24 Hours | 1-2 staff, accessed intermittently |  |  | The Disposal Room provides for the temporary holding of waste, separated into various waste streams (general waste, clinical waste, cardboard recycling, co-mingled recycling, cytotoxic waste, confidential waste, sharps, soiled linen, etc.) prior to removal to a designated waste zone in the loading dock. Waste will be collected at regular times each day, or when requested, to prevent build-up of waste in the unit. Additional Considerations: Size requirements for a Disposal Room will be dependent on a department's estimated waste output, the frequency of waste collection and local operational policies for waste management that may dictate the number of waste streams and minimum bin sizes. Adequate space must be provided to allow for segregation of waste streams. The use of 660 L bins should be carefully considered due to the WHS risks associated with manual handling, however it is acknowledged in some areas with significant waste output they may be required. Where the weight of full waste receptacles is a concern, the use of tugs/ACVs is often required. Waste transfer may be manual, by motorised movers/tugs or via automated guided vehicle (AGV) systems. Where tugs and AGVs are being considered, it is important to confirm clearances, door widths and turning circles within the room, in the adjacent corridor and along path of travel to designated waste zone. Dual access from within a unit and from an external services corridor may be provided depending on overall planning arrangement. Hold open door hardware may be required to allow for easy movement of bins and trolleys in and out of the room. Shelf may be provided for storage of small sharps bins prior to collection. If a handwashing basin is not provided within the Disposal Room, one should be located in close proximity. The location of hand hygiene facilities may be dependent on local infection control policies. Staff only access may be required if this room is provided in a public or patient care area.
DTUR-10 | Dirty Utility, 10m2 | 10 | 24 Hours | 1 - 2 intermittently |  | Sound attenuation level high, Acoustic privacy is required. | The Dirty Utility, 10m² provides a space for the decontamination and storage of bedpans, urinals and bowls, testing and disposal of patient specimens, disposal of clinical and other wastes, cleaning and holding of used equipment for collection, and holding of soiled linen for collection. Note: due to the availability of washers that do not require pre-handling of items a spray hose and flushing rim sink are no longer included in the Dirty Utility, 10m² fit out.
DTUR-12 | Dirty Utility, 12m2 | 12 | 24 Hours | 1 - 2 intermittently |  | Sound attenuation level high, Acoustic privacy is required. | The Dirty Utility, 12m² provides a space for the decontamination and storage of bedpans, urinals and bowls, testing and disposal of patient specimens, disposal of clinical and other wastes, cleaning and holding of used equipment for collection, and holding of soiled linen for collection. Note: due to the availability of washers that do not require pre-handling of items a spray hose and flushing rim sink are no longer included in the Dirty Utility, 12m² fit out.
DTUR-14 | Dirty Utility, 14m2 | 14 | 24 Hours | 1 - 2 intermittently |  | Sound attenuation level high, Acoustic privacy is required. | The Dirty Utility, 14m² provides a space for the decontamination and storage of bedpans, urinals and bowls, testing and disposal of patient specimens, disposal of clinical and other wastes, cleaning and holding of used equipment for collection, and holding of soiled linen for collection. The Dirty Utility, 14m2 may provide access doors from both sides of the room. Note: due to the availability of washers that do not require pre-handling of items a spray hose and flushing rim sink are no longer included in the Dirty Utility, 14m² fit out.
DTUR-S | Dirty Utility, Sub | 8 | 24 Hours | 1 - 2 intermittently |  | Sound attenuation level high, Acoustic privacy is required. | The Dirty Utility, Sub provides a small space for the decontamination and storage of bedpans, urinals and bowls, testing and disposal of patient specimens, disposal of clinical and other wastes, cleaning and holding of used equipment for collection, and holding of soiled linen for collection. Note: due to the availability of washers that do not require pre-handling of items a spray hose and flushing rim sink are no longer included in the Dirty Utility, Sub fit out.
ECHO-TOE | Echocardiography - Transoesophageal | 26 | Available 24 Hours | 1 patient; 1 - 2 staff |  | Acoustic privacy is required | Room for diagnostic echocardiography procedures including transoesophageal echocardiography. Echocardiography procedures may be performed under local anaesthesia and sedation. The Echocardiography room should be located close to patient toilet and change rooms.
ECL-10 | Equipment Clean-Up, 10m2 | 10 | 24 Hours | 1 - 2 |  |  | The Equipment Clean-Up Room is used for cleaning and maintaining of equipment used in the Unit. Storage may be required for holding cleaned equipment and spares. Equipment may also be recharged in this room. Natural light is desirable.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | 12 | 24 Hours | 1 - 2 staff intermittently |  |  | The Equipment Clean Up Room is used for cleaning of equipment used in patient care or loaned to outpatients following return. Storage will be required for holding cleaning equipment.
ECL-14 | Equipment Clean-Up, 14m2 | 14 | 24 Hours | 1 - 2 |  |  | The Equipment Clean-Up Room is used for cleaning and maintaining of equipment used in the Unit. Storage may be required for holding cleaned equipment and spares. Equipment may also be recharged in this room. Natural light is desirable.
ECL-8 | Equipment Clean-Up, 8m2 | 8 | 24 Hours | 1 - 2 |  |  | The Equipment Clean-Up Room is used for cleaning and maintaining of equipment used in the Unit. Storage may be required for holding cleaned equipment and spares. Equipment may also be recharged in this room. Natural light is desirable.
ENS-BR | Ensuite - Birthing | 7 | 24 Hours | 1 mother; 1-2 staff assisting |  | Sound attenuation level high, Acoustic privacy is required. | An Ensuite - Birthing Room is for a mother to wash, shower and use the toilet, either independently or with assistance. During labour, the shower may be used for extended periods to manage pain and two shower heads are provided to support this. The area and layout must accommodate an assisting midwife and a support person. Additional Considerations: The Ensuite - Birthing will have direct access from the Birthing Room. The shower floor must not have a raised hob and the entrance to the ensuite is to be flush with the adjoining flooring. A ladder or other wall mounted supportive fixtures may be included to provide support at various heights and facilitate different postures during active birthing. A birthing woman's entire weight may be supported on this item, therefore the strength of the item and fixings, including in-wall support, to be equivalent to heavy duty grab rails. Item to be suitable for wet area.
GYAH-GP | Gymnasium, Group Therapy | 105 | 7am to 6pm, daily | up to 15 patients, 2-3 staff |  | Acoustic privacy is required | The Gymnasium - Group Therapy provides space and facilities for patient evaluation, rehabilitation exercise activities, therapy and ambulation training in a group setting. An open area is typically provided for aerobic/movement exercises and may also provide space for group debriefing, discussion and education after sessions. An external outlook and ready access to outdoor areas is desirable. The range of specialised exercise equipment, open exercise areas and number of clients / patients to be accommodated within a gymnasium will vary from unit to unit and will require confirmation to inform overall space requirements. Additional Design Requirements: RCD protection is required to body protected areas in accordance with AS/NZS3003. Quantity and location dependent on selected equipment and circuiting arrangements.
GYAH-ID | Gymnasium, Individual Treatment | 76 | 7am to 6pm, daily | up to 8 patients, 3 staff |  | Acoustic privacy is required | The Gymnasium - Individual Treatment provides space and facilities for one on one patient evaluation, therapy and ambulation training. A number of curtained bays with plinths are typically provided along with an open area accommodating a range of equipment such as stairs, parallel bars and exercise bikes. An external outlook and ready access to outdoor areas is desirable. The range of specialised exercise equipment, open exercise areas and number of clients / patients to be accommodated within a gymnasium will vary from unit to unit and will require confirmation to inform overall space requirements. Additional Design Requirements: RCD protection is required to body protected areas in accordance with AS/NZS3003. Quantity and location dependent on selected equipment and circuiting arrangements.
HYDP | Hydrotherapy Pool | 260 | 7am to 8pm, daily | Up to 12 patients, 2-3 staff |  |  | The Hydrotherapy Pool is for patient therapy, rehabilitation and exercise under supervision. The pool size indicated is 15m x 6m (90m2 - not including pool surrounds). Pool and surrounds to comply with AS 3979 - Hydrotherapy Pools. Additional Design Considerations: Ready access is required to adequate change facilities for clients/patients and staff, including accessible showers and toilets. Emergency call points may be wall mounted and/or ceiling-suspended depending on the staffing model and local operational policies. Appropriate storage for aids and equipment will be required separately to the storage of pool maintenance equipment and chemicals. Depth 0.9m to 1.5 m with gradual slope. Depth at deep end may be reduced if paediatric use is likely to be frequent. Ramp access and stair access to be provided with handrails. Pool handrails are required and are not to be recessed. All stainless steel railing to be chlorine resistant. Temperature of water to be 33-34 degrees. Temperature of air to be no more than 6 degrees below water temperature. Wet deck - single channel. Water to be earthed. Internal pool surface and pool surrounds shown as tiled. Tiles or other suitable surface to be non-slip; ensure a high level of slip resistance to tiles on the ramp (vertical surfaces of tiling do not need to be slip resistant). Pool blanket is optional for insulation for energy conservation. Hydraulic hoist may be required; weight rating to be confirmed at project level to suit service requirements. Underwater pressure jets for therapeutic purposes are optional. Departmental planning to provide external windows is recommended. Disinfection system to be advised by expert consultant. Pool may be made available to the community after-hours and on the weekends. Careful consideration will need to be given to location, controlled access and security.
INTV | Interview Room | 12 | 24 Hours | Up to 8 persons |  | Sound attenuation level high, Acoustic privacy is required. | The Interview Room provides space for confidential discussion and/or counselling between staff, patients and/or family members where required. The Interview Room may also be used as a multipurpose room for small staff groups or discussions. An external outlook is desirable. Assumes staff will bring their own devices as required. A second door may be required for staff safety and security; the location of the second door will be dependent on the availability an appropriate area to egress. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements
LINAC | Linear Accelerator Treatment Room | 165 | 8am to 6pm, Monday – Friday (extended hours, weekends and/or emergency access may be required) | 1 patient, 1-4 staff |  | Acoustic privacy is required | The Linear Accelerator Treatment Room provides space and facilities for the treatment of patients requiring external beam radiation therapy (EBRT). A collocated Control Room is required for the remote operation of equipment. Cable pathways for services between the Control Room, Modulator Room and Treatment Room must be established and coordinated prior to construction and developed based on vendor and physics requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access. The Treatment Room must be adequately shielded to ensure that radiation doses to surrounding areas do not exceed radiation dose limits and satisfy design/dose constraints, including the provision of a maze entry or shielded door. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. The modulator may need to be accommodated within the Treatment Room. If so, this will be concealed by a wall, with a door for access. Depending on the quality of the power supply at a site, a power conditioner may be required and is typically located with the modulator. Additional noise from this equipment should also be accounted for with acoustic treatment. Visibility of the Treatment Room from exit and via observation cameras is required to ensure all persons are clear of the room before treatment commences. A last person out button is to be provided, as well as a sensor gate and/or physical gate at the entry of the maze. A &#8216;Beam-On/Off&#8217; light system is required at the entry to be illuminated when treatment begins/ends. Additional Considerations: Additional medical gases may be provided (including additional oxygen, additional suction, nitrous oxide and scavenge) depending on clinical requirements such as the provision of anaesthetic services. Acoustic requirements will vary depending on the equipment procured and its proximity to the patient treatment area. For patient and staff comfort, acoustics should be considered to reduce the impact of noise from the equipment. Dimmable lights are to be provided and are to be placed to avoid glare to patients. Back lit ceiling artwork may be installed for patient comfort and speakers and connections should be provided to allow music to be played and controlled individually for each bunker. Air conditioning is to be provided to control the room temperature and humidity to meet the equipment suppliers&#8217; requirements; high equipment loads shall be accounted for. Air conditioning equipment shall be located outside of this room with ductwork to be installed to follow the entry maze with any penetrations considered to ensure radiation protection. Where possible, it is preferred that ductwork in the ceiling is designed to the perimeter of room for future flexibility to provide ceiling mounted equipment. Cooling shall be provided to the linear accelerator equipment in the form of chilled water to meet the supplier&#8217;s requirements, independent chiller equipment is preferred. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. Small UPS units are normally required as a minimum to support control and operating equipment. In some cases, these are supplied with the medical equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. A mix of respiratory gating, surface guidance, motion management and/or patient position technologies may be provided depending on clinical/service requirements. The provision of these technologies is to be confirmed by the project and may have additional requirements including additional power/data, water cooling, heat load, possible increased set down, increased number of trenches/conduits, space requirements for computer cabinets in control room and additional penetrations through the suspended ceiling. Requirements to be confirmed with vendors. Allow sufficient distance between the isocentre of the machine and primary shielding wall for the flexibility to provide specialised treatments, e.g. Total Body Irradiation (TBI), depending on service requirements.
LINAC-CR | Linear Accelerator Control Room | 21 | 8am to 6pm, Monday – Friday (extended hours, weekends and/or emergency access may be required) | 2-5 staff |  | Acoustic privacy is required | The Linear Accelerator Control room is located adjacent to the Treatment Room for remote control of equipment with audio and visual contact with the patient via cameras, speakers and microphones. Each Treatment Room requires a Control Room. Cable pathways for services between the Control room to the Treatment Room must be established and coordinated prior to construction. Radiation shielding is to be as advised by a Radiation Consultant. Additional Considerations: Air conditioning to provide comfort conditions; high equipment loads shall be accounted for. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. Small UPS units are normally required as a minimum to support control and operating equipment. In some cases, these are supplied with the medical equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. Cable management and the potential for computers to be consolidated into a rack (either within the room or in a dedicated, collocated air conditioned room/cupboard) to be determined at project level. Provision of a height adjustable workstation may be considered depending on local policies and staff accessibility requirements.
LNGE-10 | Lounge - Patient / Visitor, 10m2 | 10 | Up to 24 Hours | Up to 6 patients/visitors; 1-2 staff intermittently |  | Sound attenuation level medium; acoustic privacy is required. | The Patient Lounge provides an environment away from clinical areas and/or patient bedrooms for patients and visitors to meet and socialise. It is an area where family groups can visit together with patients. A range of occupants will need to be accommodated including adults, children, able-bodied people as well as people with disabilities. Additional Considerations: The size of lounge areas can vary depending on service requirements; add 1.2m2 per additional person and allow 1.5m2 for additional wheelchair spaces as required. Access for a patient lifter or patient bed may be required depending on patient cohort and service requirements. The Patient Lounge is often located in close proximity to or combined with a Bay - Beverage. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. Patient to nurse call button and handset is optional and may be provided in units where high patient observation is required. Duress provision will be dependent on service requirements. Provision of a hearing augmentation system may be required depending on television and amplification system provided. Furniture shown in this Standard Component is indicative and items selected will vary depending on patient cohort. Considerations regarding furniture include assessing the weight rating of all seating, including for bariatric capacity as required, providing a mix of chairs with and without arms to support patient mobility, and providing built-in furniture or securing furniture where there may be safety/security concerns. Room may be designed to be open to a corridor. Additional transfer of noise to adjacent patient care areas to be considered if room is not enclosed. The display of artwork and cultural assets will be guided by the facility. A rail display system may be considered to minimise risks and allow for flexibility in display. Ease of cleaning for infection prevention and control also needs to be considered when selecting artwork. Natural light and external outlook is essential.
LNGE-20 | Lounge - Patient / Visitor, 20m2 | 20 | Up to 24 Hours | Up to 10 patients/visitors; 1-2 staff intermittently |  | Sound attenuation level medium; acoustic privacy is required. | The Patient Lounge provides an environment away from clinical areas and/or patient bedrooms for patients and visitors to meet and socialise. It is an area where family groups can visit together with patients. A range of occupants will need to be accommodated including adults, children, able-bodied people as well as people with disabilities. Additional Considerations: The size of lounge areas can vary depending on service requirements; add 1.2m2 per additional person and allow 1.5m2 for additional wheelchair spaces as required. Access for a patient lifter or patient bed may be required depending on patient cohort and service requirements. The Patient Lounge is often located in close proximity to or combined with a Bay - Beverage. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. Patient to nurse call button and handset is optional and may be provided in units where high patient observation is required. Duress provision will be dependent on service requirements. Provision of a hearing augmentation system may be required depending on television and amplification system provided. Furniture shown in this Standard Component is indicative and items selected will vary depending on patient cohort. Considerations regarding furniture include assessing the weight rating of all seating, including for bariatric capacity as required, providing a mix of chairs with and without arms to support patient mobility, and providing built-in furniture or securing furniture where there may be safety/security concerns. Room may be designed to be open to a corridor. Additional transfer of noise to adjacent patient care areas to be considered if room is not enclosed. The display of artwork and cultural assets will be guided by the facility. A rail display system may be considered to minimise risks and allow for flexibility in display. Ease of cleaning for infection prevention and control also needs to be considered when selecting artwork. Natural light and external outlook is essential.
MED-14 | Medication Room, 14m2 | 14 | 24 Hours | 1 - 2 staff intermittently |  |  | The Medication Room, 14m² is used for the secure storage of medications, including intravenous fluids, and a small amount of medication-related clean consumables. The room may provide storage for drugs of dependence in accordance with relevant legislation. Controlled temperature and humidity is required for medication storage and internal temperatures should not rise above 25°C. Secured staff only access to the room is required. The provision and placement of CCTV cameras is dependent on jurisdictional policies. Refer to project specific requirements for the provision of automated medication management systems.
MEET-15 | Meeting Room, 15m2 | 15 | Available 24 Hours | Up to 8 seated |  | Sound attenuation level high, Acoustic privacy is required. | A multipurpose room for staff meetings, education and training activities. Requirement for Hearing Augmentation to be confirmed. Videoconferencing systems and AV equipment are not required in all 15m2 meeting rooms, particularly where most use is for ad hoc meetings with BYO devices. Provision of equipment and services will need to be confirmed based on service requirements. Additional Considerations: Video conferencing systems and AV equipment will be dependent on service requirements and local ICT policies. The display system, sound system, microphone, room and video conferencing control and supporting AV equipment must be carefully considered to provide a cohesive solution for users. Where a camera is provided for video conferencing the furniture layout should be considered to allow for visibility of users within the field of view. To support video conferencing, decor should be light/neutral and free from busy patterns within the camera view for optimal image quality. Room booking panels may be provided to suit service/facility approach to room scheduling. Occupancy will vary, with the total number of people that may be accommodated in the space dependent on furniture layout. Furniture and layout will vary to suit service requirements. Furniture considerations include provision of stackable chairs and folding tables to support reconfiguration of the layout for flexible use of the space. Floor boxes for services may be provided to support flexible use of the space. Location to be confirmed to suit service requirements. Acoustics to be carefully considered to ensure appropriate sound levels, particularly with video conferencing equipment provided, and to reduce noise transfer to/from adjacent areas for privacy and to reduce disruption. Background noise (e.g. from air conditioning) should be minimised. Hard surfaces may be covered to reduce reverberation as required. Where a meeting room is intended to be used with patients and/or family, additional features for staff safety (e.g. duress buttons, additional egress doors, etc.) may be considered to suit service requirements. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. When used with patients and/or family in mental health settings, an alternative egress door should be provided. Where possible, equipment should be recessed into walls and finishes, fittings, fixtures and furniture should be suitable for mental health patients. All glazing (including external windows) is to be Grade A safety glass to comply with AS/NZS 2208 and AS 1288. Ready access to a telephone in an adjacent space may be required to answer pagers. An external outlook is desirable.
MEET-20 | Meeting Room, 20m2 | 20 | Available 24 Hours | Up to 12 seated |  | Sound attenuation level high, Acoustic privacy is required. | A multipurpose room for staff meetings, education and training activities. Requirement for Hearing Augmentation to be confirmed. Additional Considerations: Video conferencing systems and AV equipment will be dependent on service requirements and local ICT policies. The display system, sound system, microphone, room and video conferencing control and supporting AV equipment must be carefully considered to provide a cohesive solution for users. Where a camera is provided for video conferencing the furniture layout should be considered to allow for visibility of users within the field of view. To support video conferencing, decor should be light/neutral and free from busy patterns within the camera view for optimal image quality. Room booking panels may be provided to suit service/facility approach to room scheduling. Occupancy will vary, with the total number of people that may be accommodated in the space dependent on furniture layout. Furniture and layout will vary to suit service requirements. Furniture considerations include provision of stackable chairs and folding tables to support reconfiguration of the layout for flexible use of the space. Floor boxes for services may be provided to support flexible use of the space. Location to be confirmed to suit service requirements. Acoustics to be carefully considered to ensure appropriate sound levels, particularly with video conferencing equipment provided, and to reduce noise transfer to/from adjacent areas for privacy and to reduce disruption. Background noise (e.g. from air conditioning) should be minimised. Hard surfaces may be covered to reduce reverberation as required. Where a meeting room is intended to be used with patients and/or family, additional features for staff safety (e.g. duress buttons, additional egress doors, etc.) may be considered to suit service requirements. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. When used with patients and/or family in mental health settings, an alternative egress door should be provided. Where possible, equipment should be recessed into walls and finishes, fittings, fixtures and furniture should be suitable for mental health patients. All glazing (including external windows) is to be Grade A safety glass to comply with AS/NZS 2208 and AS 1288. Ready access to a telephone in an adjacent space may be required to answer pagers. An external outlook is desirable.
MEET-30 | Meeting Room, 30m2 | 30 | Available 24 Hours | Up to 18 seated |  | Sound attenuation level high, Acoustic privacy is required. | A multipurpose room for staff meetings, education and training activities. Requirement for Hearing Augmentation to be confirmed. Additional Considerations: Video conferencing systems and AV equipment will be dependent on service requirements and local ICT policies. The display system, sound system, microphone, room and video conferencing control and supporting AV equipment must be carefully considered to provide a cohesive solution for users. Where a camera is provided for video conferencing the furniture layout should be considered to allow for visibility of users within the field of view. To support video conferencing, decor should be light/neutral and free from busy patterns within the camera view for optimal image quality. Room booking panels may be provided to suit service/facility approach to room scheduling. Occupancy will vary, with the total number of people that may be accommodated in the space dependent on furniture layout. Furniture and layout will vary to suit service requirements. Furniture considerations include provision of stackable chairs and folding tables to support reconfiguration of the layout for flexible use of the space. Floor boxes for services may be provided to support flexible use of the space. Location to be confirmed to suit service requirements. Acoustics to be carefully considered to ensure appropriate sound levels, particularly with video conferencing equipment provided, and to reduce noise transfer to/from adjacent areas for privacy and to reduce disruption. Background noise (e.g. from air conditioning) should be minimised. Hard surfaces may be covered to reduce reverberation as required. Where a meeting room is intended to be used with patients and/or family, additional features for staff safety (e.g. duress buttons, additional egress doors, etc.) may be considered to suit service requirements. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. When used with patients and/or family in mental health settings, an alternative egress door should be provided. Where possible, equipment should be recessed into walls and finishes, fittings, fixtures and furniture should be suitable for mental health patients. All glazing (including external windows) is to be Grade A safety glass to comply with AS/NZS 2208 and AS 1288. Ready access to a telephone in an adjacent space may be required to answer pagers. An external outlook is desirable.
MEET-55 | Meeting Room, 55m2 | 55 | Available up to 24 Hours | Up to 45 seated |  | Sound attenuation level high, Acoustic privacy is required. | A large multipurpose room for staff meetings, education and training activities. Requirement for Hearing Augmentation to be confirmed. Additional Considerations: An operable wall may be provided to open this room to an adjoining meeting room, increasing capacity for lectures, training activities or events. Connections and coordination between AV and video conferencing equipment in adjoining rooms will need to be considered to support both separate and combined activities. Video conferencing systems and AV equipment will be dependent on service requirements and local ICT policies. The lectern design, display system, sound system, microphone, room and video conferencing control and supporting AV equipment must be carefully considered to provide a cohesive solution for users. To support video conferencing, decor should be light/neutral and free from busy patterns within the camera view for optimal image quality. Room booking panels may be provided to suit service/facility approach to room scheduling. Occupancy will vary, with the total number of people that may be accommodated in the space dependent on furniture layout. Furniture and layout will vary to suit service requirements. Furniture considerations include provision of chairs with integral note tables, stackable chairs and folding tables to support reconfiguration of the layout for flexible use of the space. A storage zone or enclosed storeroom may be provided to hold additional furniture to support alternate layouts. If multiple rooms are provided, furniture storage areas may be shared. Floor boxes for services may be provided to support flexible use of the space. Location to be confirmed to suit service requirements. Acoustics to be carefully considered to ensure appropriate sound levels, particularly with video conferencing equipment provided, and to reduce noise transfer to/from adjacent areas for privacy and to reduce disruption. Background noise (e.g. from air conditioning) should be minimised. Hard surfaces may be covered to reduce reverberation as required. Where a meeting room is intended to be used with patients and/or family, additional features for staff safety (e.g. duress buttons, additional egress doors, etc.) may be considered to suit service requirements. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. When used with patients and/or family in mental health settings, an alternative egress door should be provided. Where possible, equipment should be recessed into walls and finishes, fittings, fixtures and furniture should be suitable for mental health patients. All glazing (including external windows) is to be Grade A safety glass to comply with AS/NZS 2208 and AS 1288. Ready access to a telephone in an adjacent space may be required to answer pagers.
MEET-9 | Meeting Room, 9m2 | 9 | Available 24 Hours | Up to 4 seated |  | Sound attenuation level high, Acoustic privacy is required. | A small multipurpose room for staff interviews, meetings, education and training activities. This room may also function as a 'breakout' room in an open workstation environment. Requirement for Hearing Augmentation to be confirmed. Videoconferencing systems and AV equipment are not required in all 9m2 meeting rooms, particularly where most use is for ad hoc meetings with BYO devices. Provision of equipment and services will need to be confirmed based on service requirements. Additional Considerations: Video conferencing systems and AV equipment will be dependent on service requirements and local ICT policies. The display system, sound system, microphone, room and video conferencing control and supporting AV equipment must be carefully considered to provide a cohesive solution for users. Where a camera is provided for video conferencing the furniture layout should be considered to allow for visibility of users within the field of view. To support video conferencing, decor should be light/neutral and free from busy patterns within the camera view for optimal image quality. Room booking panels may be provided to suit service/facility approach to room scheduling. Acoustics to be carefully considered to ensure appropriate sound levels, particularly with video conferencing equipment provided, and to reduce noise transfer to/from adjacent areas for privacy and to reduce disruption. Background noise (e.g. from air conditioning) should be minimised. Hard surfaces may be covered to reduce reverberation as required. For small meeting spaces suitable for use with patients please refer to the Interview Room Standard Component. Ready access to a telephone in an adjacent space may be required to answer pagers. An external outlook is desirable.
MOR-AU | Mortuary - Autopsy Room | 30 | Available 24 Hours | Up to 4 |  |  | The Mortuary - Autopsy Room is for undertaking post mortem examinations. An observation area is to be provided to at least one Autopsy Room within a department. Layout of tables and orientation of observation bay may be adjusted to suit the type of autopsy being performed (e.g. coronial or non-coronial). The Mortuary - Autopsy Room should be located with ready access to Mortuary - Body Holding. Additional Design Considerations: doors to be sealed for odours, post mortem table to have water flushing system. Appropriate Smoke/Heat detection to be considered.
MOR-BH | Mortuary - Body Holding | 22 | 24 Hours | 1 - 2 staff intermittently |  |  | The Mortuary - Body Holding provides refrigerated storage for the temporary holding of bodies with an adjacent area to transfer a body from a trolley into storage. The room can be configured with refrigerated cabinets or as a cool room depending on service requirements. The Mortuary - Body Holding should be located with ready access to the Mortuary - Viewing Room.
MOR-BR | Mortuary - Body Reception | 7 | up to 24 Hours | 1 - 2 staff intermittently |  |  | The Mortuary - Body Reception is an area for receiving bodies from the hospital and checking-in to the holding room/cool room as well as the checking-out of bodies for release to funeral directors, relatives or for transfer to the coroner. This room should be located on a corridor that is in close proximity to both the staff-only mortuary entry and the exit to the funeral director's covered parking area. Entry to this area by authorised personnel only. Handwashing facilities should be located in close proximity.
MOR-VR | Mortuary - Viewing Room | 8 | up to 24 Hours | 4-5 visitors; 1 staff |  | Sound attenuation level high, Acoustic privacy is essential | The Mortuary - Viewing Room is provided for relatives and others to view a body with the support of a social worker. The Mortuary - Viewing Room will be located adjacent to the Mortuary - Waiting. A body may be viewed through a window from within the Mortuary - Waiting, with access provided between the two rooms provided to allow for viewing a body directly.
MOR-W | Mortuary - Waiting | 9 | up to 24 Hours | 4-5 visitors; 1 staff |  | Sound attenuation level high, Acoustic privacy is essential | The Mortuary - Waiting provides an area for relatives and others to await the viewing of a body and to have private discussions with social workers. The area should be comfortable and non-clinical in appearance with furnishings and artwork that are culturally appropriate. Natural light is desirable. The Mortuary - Waiting will be located adjacent to the Mortuary - Viewing Room. A body may be viewed through the window from within the Mortuary - Waiting, with access provided between the two rooms to view a body directly.
OFF-1P-12 | Office - 1 Person, 12m2 | 12 | 8am to 6pm, daily | 1 staff, up to 3 visitors |  | Sound attenuation level medium, Acoustic privacy is required. | An office to accommodate 1 staff member carrying out computer based and/or administrative tasks in privacy. The room size allows for a small meeting table within the room for meetings or interviews with visitors (including staff, patients or family) as required. Additional Considerations: Consider placement of power and data to support flexible arrangement of the layout to suit user requirements. Workstation length to service requirements and local WHS/office accommodation policies. An external outlook is desirable.
OFF-1P-9 | Office - 1 Person, 9m2 | 9 | 8am to 6pm, daily | 1 staff, up to 2 visitors |  | Sound attenuation level medium, Acoustic privacy is required. | An office to accommodate 1 staff member carrying out computer based and/or administrative tasks in privacy. Small meetings, ad hoc discussions or interviews with visitors (including staff, patients or family) may be held in this space, depending on service requirements. Additional Considerations: Consider placement of power and data to support flexible arrangement of the layout to suit user requirements. Workstation length to service requirements and local WHS/office accommodation policies. An external outlook is desirable.
OFF-2P | Office - 2 Person | 12 | 8am to 6pm, daily | 2 staff, up to 2 visitors |  | Sound attenuation level medium, Acoustic privacy is required. | An office to accommodate 2 staff members carrying out computer based and/or administrative tasks in privacy. Small meetings, ad hoc discussions or interviews with visitors (including staff, patients or family) may be held in this space, depending on service requirements. Additional Considerations: Workstation length to service requirements and local WHS/office accommodation policies. An external outlook is desirable.
OFF-CLN | Office - Clinical Workroom | 15 | Available 24 hours | 4 workstations |  | Sound attenuation level high, Acoustic privacy is required. | An Office - Clinical Workroom, typically adjoining or in close proximity to a Staff Station, is for staff to write up notes, review patient records and digital images, hold confidential discussions, store records and conduct handovers. One computer has been shown as a downtime computer with UPS power; this requirement will need to be confirmed by the department. The downtime computer, if required by a department, may alternatively be provided in a Staff Station. Provision of a downtime computer will need to be rationalised across the unit to avoid duplication. Additional Design Considerations: Location of power and data outlets (above or below the bench) is to meet local operational/WHS requirements. Cable management solutions to be considered to meet infection control and cleaning requirements. A desktop printer may be provided. Printing/scanning/photocopying facilities should be consolidated across the unit to suit service requirements. Provision of height adjustable workstations may be considered depending on local policies and staff accessibility requirements.
OFF-WI-5 | Office - Write-up, 5m2 | 5 | 24 Hours | 2 staff members |  |  | The Write-up Bay provides a write-up bench for use by members of the patient care team. In this space the User may review and write-up patient records, enter patient data on computer and make telephone calls. If located within the Operating Unit it may also include an optional area for porters to wait until paged or called by staff. The Write-up space must not impede the corridor.
OFF-WS | Office - Workstation | 4.5 | 8am to 6pm, daily | 1 staff/visitor |  |  | A workstation within an open plan arrangement. Multiple staff may utilise this space at different times to carry out computer based and/or administrative tasks. Additional Considerations: Depending on overall layout, services may be wall mounted or workstations with soft-wired cable trays may be connected via an umbilical to the floor or ceiling. Workstation length to service requirements and local WHS/office accommodation policies. A modesty panel or extended acoustic panel may be required depending on layout. Shared document storage may be allocated separately to the workstation. Small meeting tables for informal discussions may be provided between a number of workstations in larger open plan areas. Where devices are to be brought with the user a dock and monitor(s) may be provided in lieu of a computer. Adequate coverage for Wireless Access Point (WAP) to be provided. An external outlook is desirable.
OVBR | Overnight Stay - Bedroom | 10 | Available 24 Hours | 1 staff/visitor |  | Sound attenuation level high, Acoustic privacy is required. | A 1 bed room with an adjoining ensuite for staff, parents or carers needing to remain on close call overnight. The room shall be lockable. Additional Design Considerations: An external outlook is required. Overnight stay bedroom and ensuite layouts may be adjusted to provide accessible facilities i.e. as required if a standalone suite is provided or in a portion of rooms if multiple overnight stay suites are provided.
OVES | Overnight Stay - Ensuite | 5 | Available 24 Hours | 1 staff/visitor |  | Acoustic privacy is required | An ensuite with shower, toilet and basin adjoining the overnight stay bedroom for staff, carers or parents needing to remain on close call overnight. Additional Design Considerations: Overnight stay bedroom and ensuite layouts may be adjusted to provide accessible facilities e.g. as required if a standalone suite is provided or in a portion of rooms if multiple overnight stay suites are provided.
PAR | Parenting Room | 9 | Available 24 Hours | Up to 2 |  | Sound attenuation level high, Acoustic privacy is required. | The Parenting Room is provided as a private space for parents to feed and change babies and will provide access for prams. The Parenting Room will be located to have ready access to a public amenities and waiting areas. Additional Design Considerations: Provision of a fridge may be considered if servicing a staff area; this will be dependent on operational policies. Custom joinery for baby changing facilities may be considered, dependent on operational policies and project requirements.
PBSC | Patient Bay - Specimen Collection | 9 | 8am to 6pm, daily | 1 patient, 1 staff |  |  | A Patient Bay used for the collection of specimens to be transferred to a laboratory for testing. Additional Considerations: Ready access to a handwashing basin is required. Ready access may be required to a pneumatic tube bay that is in a staff supervised area. Recliner must be able to lay flat to support the patient in the event of an adverse reaction (e.g. fainting). A desk may be provided in lieu of Workstation on Wheels depending on operational model. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. A risk assessment should be undertaken to inform staff safety requirements with adjoining cubicles which may require adjustments to the design and/or operational strategies, such as the use of mobile duress systems. Refer to the Pathology HPU for paediatric requirements.
PBTR-MD | Patient Bay - Medical Day Treatment | 9 | 8am to 6pm, daily | 1 patient: 1-2 staff, 1 visitor |  |  | The Patient Bay - Medical Day Treatment may be used for assessment and treatment of non-admitted patients receiving scheduled treatment in either a bed or recliner (e.g. for infusions such as antibiotics and chemotherapy). A staff handwashing basin will be provided for each 4 bays and the Patient Bay will have, wherever possible, clear line of sight to a staff station and close access to a patient toilet. Access to natural light and views is required. Medical Gases may be provided individually to each bay or shared between two bays depending on service requirements.
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | 9 | 7am to 9pm, daily | 1 patient, 1 visitor, 1-2 staff |  |  | The Patient Bay - Renal Dialysis is for patients receiving haemodialysis therapy. Treatments for each patient may take from 3 to 6 hours, occurring up to 3 days per week. A staff handwashing basin will be provided for each 4 bays and the Patient Bay will have, wherever possible, clear line of sight to a staff station. Access to natural light and views is required. Furniture and furnishings should provide a high level of patient comfort and amenity.
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | 9 | 7am to 9pm, daily | 1 patient, 1 visitor, 1-2 staff |  |  | The Patient Bay - Renal Dialysis is for patients receiving haemodialysis therapy. Treatments for each patient may take from 3 to 6 hours, occurring up to 3 days per week. A staff handwashing basin will be provided for each 4 bays and the Patient Bay will have, wherever possible, clear line of sight to a staff station. Access to natural light and views is required. Furniture and furnishings should provide a high level of patient comfort and amenity.
PET-CT | PET-CT Imaging Room | 50 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | The PET-CT Imaging Room provides the area and facilities for Positron Emission Tomography-Computed Tomography (PET/CT) scanning procedures with a Control Room and Equipment Room (as required) located nearby. It is not essential that the control is directly adjacent to the Imaging Room where patients are monitored via observation cameras, however ready access is required to attend to patient needs. If the Control Room is adjacent an observation window may be provided and shielding requirements will need to be confirmed. Cable pathways for services between the Control Room, Equipment Room and Imaging Room must be established and coordinated prior to construction and developed based on vendor requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access and support structure as required. The PET-CT Imaging Room must be adequately shielded to ensure that radiation doses to surrounding areas do not exceed radiation dose limits and satisfy design/dose constraints. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. Additional Considerations: Configuration of medical gases is to be confirmed to meet clinical requirements such as the provision of anaesthetic services. Dimmable lights are recommended and are to be placed to avoid glare to patients. Back lit ceiling artwork may be installed for patient comfort. 'X-ray in Use' lights are required outside each entry door. Laser marker system for procedure planning may be required subject to model of care and equipment selected. Structural assessment is required to ensure the floor will be rigid with no vibration and will accommodate the weight of the PET-CT unit. Room temperature and humidity control is required. Temperature requirements for equipment are to be confirmed by the vendor, including requirement for circulated cooling water. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. UPS power is required to the PET-CT scanning and control equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. An optional laser system for patient positioning and procedure planning has been indicated and may be provided subject to service requirements. A mix of respiratory gating, surface guidance, motion management and/or patient position technologies may be considered depending on clinical/service requirements (e.g. for radiation therapy planning). The provision of these technologies is to be confirmed by the project and may have additional requirements including additional power/data, water cooling, heat load, possible increased set down, increased number of trenches/conduits, space requirements for computer cabinets in control room and additional penetrations through the suspended ceiling. Requirements to be confirmed with vendors. Consideration needs to be given to the clearances of equipment and replacement part delivery routes.
PHA-CO | Pharmacy - Counter | 9 | 8am to 8pm, daily | 1-2 staff |  | Sound attenuation level - Medium; acoustic privacy required. | The Pharmacy &#8211; Counter provides a reception point for the Pharmacy Unit for patients and staff to drop off prescriptions to be filled and collect dispensed prescriptions. The Counter is to have good visibility of the waiting area and visitors as they approach. Additional Design Requirements: Information on prescribed medications will be provided to patients at the counter and a degree of privacy is required for these discussions. Interview rooms may be located nearby to provide more private information or counselling. Cash handling and EFTPOS facilities may be required in this space depending on operational models and the proximity of the Unit to the central hospital cashier. Security glazing will be required. The design must support clear communication between pharmacy staff and patients with communication slots. A hearing augmentation system may be considered. Small items will also need to be passed through and forms may need to be signed &#8211; the design of the glazing needs to support this function. The ability to lock down the Unit, e.g. in evenings when the pharmacy is closed, is required. This may be achieved by lockable doors from public corridors to the waiting area, however lockable shutters to the counter may be considered.
PHA-DB | Pharmacy - Distribution Workstation | 6 | 8am to 8pm, daily | 1 staff |  |  | The Distribution Workstation provides workspace to sort and process dispensed medications to be distributed from the Pharmacy to clinical units (e.g. imprest medications) and may be used to prepare specialised packs (e.g. resuscitation trolley supplies). Additional Considerations: Ready access to bulk store and refrigerated storage is required from this area. Direct access to a dispatch/collection area is required for holding items prior to delivery. Printer and telephone facilities to be shared between multiple workstations. A Multifunction device is to be in close proximity to the distribution area. Ensure adequate lighting over workspace.
PHA-DS | Pharmacy - Dispensing Workstation | 4 | 8am to 8pm, daily | 1 staff |  |  | The Dispensing Workstation provides workspace to prepare, assemble and label patient medications for dispensing to inpatients and outpatients. The workstations can be configured to provide dispensing/assembly stations and a checking station as required. Dispensing workstations should be arranged to support effective workflows, reduce travel time and avoid error. These stations should be arranged such that extended reaches are avoided. Where dispensing workstations are arranged adjacent to each other, corridors in between them should be wide enough to provide unobstructed movement for staff and trolleys. Additional Considerations: Ready access to the bulk store and refrigerated storage is required. Printer and telephone facilities to be shared between multiple workstations. A Multifunction device is to be in close proximity to the dispensing area. Ensure adequate lighting over workspace - task lighting may be required.
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | 12 | 8am to 8pm, daily | Up to 2 Staff |  |  | The Preparation Room provides a dedicated area for the extemporaneous preparation and compounding of non-aseptically prepared products and cleaning of equipment as required. The Preparation Room should be located in close proximity to the dispensing area.
PHA-RE | Pharmacy - Returns Workstation | 6 | 8am to 8pm, daily | 1 staff |  |  | The Returns Workstation provides workspace to sort and process medication returns from within the hospital that are to be returned to stock or sent for disposal. Storage is required to accommodate the returns before they are processed. Additional Considerations: Ready access to refrigerated storage and disposal room is required from this area. Printer and telephone facilities to be shared between multiple workstations. A Multifunction device is to be in close proximity to the returns area. Ensure adequate lighting over workspace.
PLAY | Play Area - Paediatric | 10 | Available 24 Hours | Up to 4 |  | Acoustics: sound attenuation level - high. | An area, typically associated with a waiting area, where children may play while waiting with parents. The Play Area should allow access for prams and be located in close proximity to public amenities, including baby change facilities. The Play Area should be designed to minimise noise transfer to other areas. Additional Considerations: Provision of age appropriate furnishing, artwork and play equipment (if provided) to be based on service and population. Furniture that is easily wiped clean is recommended. Furniture suitable for adults is typically provided within the space to support supervision of children but may be provided in the adjacent waiting area if layout allows for adequate visibility. Provision of interactive board/projector technologies and associated services may be considered depending on operational policies and ICT support.
PLNT-WT | Plant - Water Treatment | 18 | 24 Hours | 1 - 2 staff |  | High level sound isolation | The Plant - Water Treatment is a lockable room for water treatment systems, including particle filters, water softeners, carbon filters and reverse osmosis (RO) water plant/systems, for the provision of safe water for patients receiving haemodialysis therapy. The Water Treatment Plant Room should be located in close proximity to the Renal Dialysis Unit to permit short tubing runs to each Treatment Bay, and permit staff to monitor and service the water treatment systems. Refer to manufacturer's specifications for details of RO water treatment plant equipment. In particular, this equipment will incorporate a heat disinfection function, and water saving features. The final components of the system will be determined by the quality of feed water and the ability of the overall system to produce and maintain appropriate water quality. Water treatment requirements to be confirmed by a dialysis water treatment specialist. Special Design Requirements: High level sound isolation is required to ensure noise generated from this room does not invade treatment spaces. Structural engineer's assessment must be sought for floor load bearing capacity with respect to water treatment and pre-treatment plant equipment. Ventilation, exhaust and/or air-conditioning must be designed to accommodate the heat loads of the specified equipment. Light is to be minimised; there are to be no windows provided. Service access will be required around the perimeter of plant equipment, clearances will be determined by manufacturer's specifications. A second door may be provided for separate deliveries and maintenance access if located in the patient treatment area. Pipework and components installed after the water inlet in this room shall not contain brass or copper. Options for cooling of high temperature (circa 90°C) waste water discharged from RO plant is to be considered. Where chilled water is not available, a chiller may be required outside the plant room. The maximum temperature of the feed water to be assessed to allow determination of any requirement for a heat exchanger or chiller design as determined by expert. The intent of this is ensure the temperature at the dialysis point for patient care is between 30-37°C. Redundancy within the pre-treatment and treatment devices/filters should be considered based on the hospitals service criticality level, the feed water quality or risk around quality, and the hospital location where supply of replacement equipment is difficult.
PLST | Plaster Room | 18 | 24 Hours | 1 patient; 2 -4 staff |  | Sound attenuation level - high; acoustic privacy is required | The Plaster Room allows for the application of plaster or other splints and for the closed reduction of displaced fractures or dislocations under sedative or regional anaesthesia. Typically provided as a dedicated room for high volume services, otherwise, a generic Treatment Room may be used with the inclusion of a plaster trap and other specific plastering equipment. The cast saw shall incorporate an integrated dust extraction system or the room shall have an independent dust extraction system suitably located to capture the dust at source. A room exhaust shall also be included to prevent dust from entering adjacent spaces. The plaster trap shall be easily accessible for cleaning and maintenance.
PROC | Procedure Room | 20 | 24 Hours | 1 patient; 2-4 staff |  | Sound attenuation level - High. Acoustic privacy is required. | The Procedure Room provides a controlled environment and facilities for carrying out procedures that do not require the full facilities of the Operating Suite. This may include suturing of wounds, complex dressings, lumbar puncture, catheterisations, administration of local anaesthesia and use of medical gases in connection with the procedures being performed. General anaesthesia will not be administered in this room. Additional Design Considerations: Lead shielding may be required depending on use of mobile imaging equipment. Treatment of bariatric patients including mobile, bariatric lifter access. The ceiling height has been nominally shown at 3000mm, however 2700mm is also acceptable and the height is to be determined based on ceiling mounted equipment and operational requirements.
PT-HOLD-A | Patient Bay - Holding, Chair | 6.5 | up to 24 Hours | 1 patient, 1 visitor, 1-2 staff intermittently |  |  | A Patient Bay - Holding provides facilities for patient holding and preparation prior to examination, diagnostic imaging, treatment, procedures or surgery. A handwashing basin will be provided close by to be shared between multiple bays. Staff observation is required and, where possible, a clear line of sight from a staff station is to be provided. Close access to a patient toilet is required. Additional Considerations: The configuration of the Medical Services Panel (MSP), including medical gases, power and data, will be dependent on service requirements. Data provision is to take into consideration the increasing demand and sophistication of information systems and the technological development of monitoring equipment. The requirement for fixed wall mounted patient monitoring will be dependent on service requirements. A patient monitor may be located on the left or right side of bedhead depending on monitor functionality and staff workflows. Appropriate wall supports are required for mounting the bracket where a monitor is provided. Where patient monitoring is provided, criticality of monitoring is to be assessed. Provision of emergency/standby power or UPS is to be confirmed to suit site and service requirements. Confirmation to be based on risk assessment considering the impact of a power outage on patient care/safety. Facilities to accommodate bariatric patients within a select number of bays, including accommodating bariatric recliners, is to be considered to ensure sufficient amenity is provided. Provision will be dependent on service and patient cohort requirements. Depending on service requirements, enclosed holding bays may be provided. Where a Patient Bay is to be enclosed, additional area is required to accommodate a dedicated handwashing basin and circulation at the door. Observation windows may be required and the extent of glazing to be provided is to suit requirements for staff supervision. PPE may be consolidated between multiple bays. Provision of PPE and placement throughout the overall department will be dependent on local infection control policies. A mobile sharps bin will be shared between multiple bays as required. Where use of sharps is frequent, a wall mounted bin sharps bin may be considered. Electronic Medical Record (EMR) management may be performed on a wirelessly connected workstation on wheels (WOW) that is shared between a number of bays. Ensure circulation space is sufficient for bedside and bed end data entry. Parking for WOWs and facilities for battery charging are required to support this model of care. Alternative configurations, such as wall mounted monitors in bays, may be considered depending on local ICT policies. Patient entertainment systems may be considered depending on anticipated length of stay within the holding bay and patient cohort requirements (e.g. paediatrics). Solutions may include individual systems provided to each bay (with integration to patient handset for audio and content control), a central display screen visible from multiple bays, or portable solutions, such as tablets, brought to patients as needed. Operational models for handling patient belongings are to be considered. Close proximity to patient lockers or other secure property storage may be required where belongings are not to be kept with the patient. Depending on anticipated length of stay and pre-procedure requirements, local control of lighting within the bay may be provided to allow patients to dim/adjust lighting levels to increase comfort. Access to natural light is required and an external outlook is recommended.
PT-HOLD-B | Patient Bay - Holding, Bed | 9 | up to 24 Hours | 1 patient, 1 visitor, 1-2 staff intermittently |  |  | A Patient Bay - Holding provides facilities for patient holding and preparation prior to examination, diagnostic imaging, treatment, procedures or surgery. A handwashing basin will be provided close by to be shared between multiple bays. Staff observation is required and, where possible, a clear line of sight from a staff station is to be provided. Close access to a patient toilet is required. Additional Considerations: The configuration of the Medical Services Panel (MSP), including medical gases, power and data, will be dependent on service requirements. Data provision is to take into consideration the increasing demand and sophistication of information systems and the technological development of monitoring equipment. The requirement for fixed wall mounted patient monitoring will be dependent on service requirements. A patient monitor may be located on the left or right side of bedhead depending on monitor functionality and staff workflows. Appropriate wall supports are required for mounting the bracket where a monitor is provided. Where patient monitoring is provided, criticality of monitoring is to be assessed. Provision of emergency/standby power or UPS is to be confirmed to suit site and service requirements. Confirmation to be based on risk assessment considering the impact of a power outage on patient care/safety. Facilities to accommodate bariatric patients within a select number of bays, including accommodating bariatric beds, is to be considered to ensure sufficient amenity is provided. Provision will be dependent on service and patient cohort requirements. Patient transfer aids such as mobile hoists, patient slides and hover mats are to be considered for assisted patient transfers to and from beds in holding bays. A ceiling mounted hoist may also be considered for patient transfers. Selection of transfer aids will be dependent on service and patient cohort requirements as well as WHS risk assessment. Appropriate ceiling support is required for the provision of a ceiling mounted hoist. Depending on service requirements, enclosed holding bays may be provided. Where a Patient Bay is to be enclosed, additional area is required to accommodate a dedicated handwashing basin and circulation at the door. Observation windows may be required and the extent of glazing to be provided is to suit requirements for staff supervision. PPE may be consolidated between multiple bays. Provision of PPE and placement throughout the overall department will be dependent on local infection control policies. A mobile sharps bin will be shared between multiple bays as required. Where use of sharps is frequent, a wall mounted bin sharps bin may be considered. Electronic Medical Record (EMR) management may be performed on a wirelessly connected workstation on wheels (WOW) that is shared between a number of bays. Ensure circulation space is sufficient for bedside and bed end data entry. Parking for WOWs and facilities for battery charging are required to support this model of care. Alternative configurations, such as wall mounted monitors in bays, may be considered depending on local ICT policies. Patient entertainment systems may be considered depending on anticipated length of stay within the holding bay and patient cohort requirements (e.g. paediatrics). Solutions may include individual systems provided to each bay (with integration to patient handset for audio and content control), a central display screen visible from multiple bays, or portable solutions, such as tablets, brought to patients as needed. Operational models for handling patient belongings are to be considered. Close proximity to patient lockers or other secure property storage may be required where belongings are not to be kept with the patient. Depending on anticipated length of stay and pre-procedure requirements, local control of lighting within the bay may be provided to allow patients to dim/adjust lighting level to increase comfort. Access to natural light is required and an external outlook is recommended.
PT-RS1 | Patient Bay - Recovery, Stage 1 | 9 | up to 24 Hours | 1 patient; 1-2 staff |  |  | A Patient Bay - Recovery, Stage 1 provides facilities for the recovery of patients from anaesthetic post procedure, and is located within the Operating Unit or Day Procedures Unit. A handwashing basin will be provided close by to be shared between up to 4 Patient Bays. Staff observation is required and a clear line of sight from a staff station should be provided. Additional Considerations: The configuration of the Medical Services Panel (MSP), including medical gases, power and data, will be dependent on service requirements. Data provision is to take into consideration the increasing demand and sophistication of information systems and the technological development of monitoring and diagnostic equipment. A patient monitor may be located on the left or right side of bedhead depending on monitor functionality and staff workflows. Appropriate wall supports are required for mounting the bracket where a monitor is provided. Provision of emergency/standby power or UPS for patient monitoring is to be confirmed to suit site and service requirements. Confirmation to be based on risk assessment considering the impact of a power outage on patient care/safety. Facilities to accommodate bariatric patients within a select number of bays, including accommodating bariatric beds, is to be considered to ensure sufficient amenity is provided. Provision will be dependent on service and patient cohort requirements. Patient transfer aids such as mobile hoists, patient slides and hover mats are to be considered for assisted patient transfers to and from beds in holding bays. A ceiling mounted hoist may also be considered for patient transfers. Selection of transfer aids will be dependent on service and patient cohort requirements as well as WHS risk assessment. Appropriate ceiling support is required for the provision of a ceiling mounted hoist. Depending on service requirements, enclosed recovery bays may be provided. Where a Patient Bay is to be enclosed, additional area is required to accommodate a dedicated handwashing basin and circulation at the door. Observation windows may be required and the extent of glazing to be provided is to suit requirements for staff supervision. Additional area may be required in a percentage of Recovery Stage 1 bays to suit various services requirements, e.g. to support maternity patients or paediatric patients, for accommodating bariatric patients, for care of complex patients, or for extended recovery models. The percentage of bays that may be provided with increased area will be dependent on the anticipated case mix. Electronic Medical Record (EMR) management may be performed on a wirelessly connected workstation on wheels (WOW) that is shared between a number of bays. Ensure circulation space is sufficient for bedside and bed end data entry. Parking for WOWs and facilities for battery charging required to support this model of data entry. Alternative configurations, such as wall mounted monitors in bays, may be considered depending on local ICT policies. Access to natural light is required. Where windows are provided, integral blinds are required for infection prevention and control.
PTRY | Pantry | 9 | Up to 24 Hours | 1-2 staff intermittently |  |  | The Pantry is for staff preparing and/or heating refreshments and snacks for patients, washing utensils and crockery or temporarily storing dirty items for collection, storing food and drink and disposing of food waste. This room may be shared between multiple units depending on service requirements. Additional Considerations: If food re-thermalisation trolleys are to be located in this room during meal times an additional 2m2 per trolley will be required as well as appropriate power outlets. Refer to Bay - Meal Trolley (BMT) for requirements. Local food preparation regulations and infection control policies may require a Type B hand wash basin, separate to the sink, to be included in the room or may allow for it to be located nearby. Fridge dimensions will vary depending on selection and service requirements; clear area provided for the fridge is to allow for adequate ventilation and full opening of door. Where double door fridges are provided, additional area will be required. If an instantaneous boiling water tap/unit is provided, access should be restricted to the water unit e.g. within a cupboard. If system is located within a joinery unit, ensure appropriate ventilation is provided. Provision and type of ice machine will be dependent on service requirements and local infection prevention and control policies. Where ice machine is provided, ensure adequate support to bench for weight of appliance. Where tea service is provided the height of the hot water dispenser is to be considered to ensure large kettles can be filled easily. Size, type and quantity of waste bins is dependent on environmental management policies and service requirements. Where a facility has additional waste streams, beyond the general waste and comingled recycling streams indicatively shown in this Standard Component, additional area may be required to accommodate the receptacles. Depending on frequency of waste removal and anticipated traffic, bin sizes may be larger in high use areas or as guided by local waste management policies. Joinery to conceal bins may be provided depending on infection prevention and control policies. Storage of nutritional supplements may require a lockable cupboard and will be dependent on service requirements. Temperature of fridge may need to be monitored. A data outlet, or alternative connection to suit project/equipment requirements, may be required for connection the Building Management System (BMS) depending on service requirements. Controlled, staff-only access may be required if this room is provided in a public or patient care area.
RECP-10 | Reception, 10m2 | 10 | 8am to 6pm, daily, dependent on function of unit | 1 staff |  | Sound attenuation level - Medium; acoustic privacy required. | A Reception area where visitors to a Unit can be received and directed to their destination or to a Waiting Area. The Reception may also act as an administrative base as well as controlling access to the Unit. Additional Considerations: Dependent on location, security glazing or shutters may be required. A recessed area for front access by wheelchair users is typically required. Depending on check in requirements at Reception, a set down area for visual access only may be sufficient. Adjustable height workstations may be provided in line with WHS policies and access for disabled staff working within the Reception area is to be considered. CCTV monitor, intercom and door release buttons linked to entry points may be required for access control depending on location. A desktop printer and labels printer may be provided as required. Printing, scanning and photocopying facilities may be consolidated with ready access from Reception to suit service requirements. Swipe card access to reception may be provided to suit security requirements.
RECP-15 | Reception, 15m2 | 15 | 8am to 6pm, daily, dependent on function of unit | 2 staff |  | Sound attenuation level - Medium; acoustic privacy required. | A Reception area where visitors to a Unit can be received and directed to their destination or to a Waiting Area. The Reception may also act as an administrative base as well as controlling access to the Unit. Additional Considerations: Dependent on location, security glazing or shutters may be required. A recessed area for front access by wheelchair users is typically required. Depending on check in requirements at Reception, a set down area for visual access only may be sufficient. Adjustable height workstations may be provided in line with WHS policies and access for disabled staff working within the Reception area is to be considered. CCTV monitor, intercom and door release buttons linked to entry points may be required for access control depending on location. A desktop printer and labels printer may be provided as required. Printing, scanning and photocopying facilities may be consolidated with ready access from Reception to suit service requirements. Swipe card access to reception may be provided to suit security requirements.
RECP-20 | Reception, 20m2 | 20 | 8am to 6pm, daily, dependent on function of unit | 3 staff |  | Sound attenuation level - Medium; acoustic privacy required. | A Reception area where visitors to a Unit can be received and directed to their destination or to a Waiting Area. The Reception may also act as an administrative base as well as controlling access to the Unit. Additional Considerations: Dependent on location, security glazing or shutters may be required. A recessed area for front access by wheelchair users is typically required. Depending on check in requirements at Reception, a set down area for visual access only may be sufficient. Adjustable height workstations may be provided in line with WHS policies and access for disabled staff working within the Reception area is to be considered. CCTV monitor, intercom and door release buttons linked to entry points may be required for access control depending on location. A desktop printer and labels printer may be provided as required. Printing, scanning and photocopying facilities may be consolidated with ready access from Reception to suit service requirements. Swipe card access to reception may be provided to suit security requirements.
REPR | Reporting Room | 12 | Available 24 Hours | Up to 4 (6 to 8 for limited periods for image reviews by team) |  | Sound attenuation level high, acoustic privacy is required. | Room within a Clinical or Diagnostic Unit for reviewing and reporting of computerised patient imaging/PACS images. Film images may also be viewed from time to time.
REPW | Reporting Workstation | 6 | Available 24 Hours | 1 staff member, 1 supervisor on occasion |  | Acoustic privacy is required | Single workstation within a Clinical or Diagnostic unit for reviewing and reporting of computerised patient imaging/ PACS images. Intended for use as a component in spaces with more than one workstation.
SECR | Security Observation Room | 12 | 24 Hours | Up to 2 staff |  | Sound attenuation level medium; acoustic privacy is required. | The Security Observation Room is a lockable room for security personnel to observe nearby areas and monitor surveillance equipment. This room may be provided as a satellite room separate to the main Security Department (e.g., in the Emergency Unit) or as a reception area for the Security Department. Security staff in will be responsible maintaining order and de-escalating developing situations where aggression may become an issue. Staff identification and access control cards may be produced by security staff in this room. Security staff may also receive, receipt and record lost items of value, contraband and weapons, including reporting to police services in line with jurisdictional policies and procedures. Security Observation Rooms are typically located close to the main entrance, after hours entrances and/or overseeing the emergency unit waiting area. Additional Considerations: ICT requirements, including computer, camera and printer may be required for creating identification and access control cards for staff as well as passes for contractors and visitors, depending on facility requirements. Requirements for photography for identification card are to be confirmed based operational policies. Where provided, a camera will need to be located near a computer with unobstructed view through an observation window towards a blank wall in an area with adequate lighting. Lockdown buttons to entry ways may be provided within the room dependent on service requirements and will be determined based on overall security strategy for the facility. The provision and configuration of a weapons safe within the room will be depending on operational policies and jurisdictional requirements. Clear signage indicating the function of the room is required.
SHAC | Shower - Accessible | 4 | 24 Hours | 1 patient; 1-2 staff assisting as required |  | Sound attenuation level high, Acoustic privacy is required. | The Shower - Accessible provides a shower for use by patients, visitors or staff with disabilities (with or without assistance). If the Shower - Accessible is located in a patient care area the nurse call system should be included and emergency release door hardware should be installed to allow staff access in an emergency. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system. The entrance to the room should be flush with the adjoining floor for access in commode chairs/wheelchairs. Facilities for hand hygiene and PPE to be available nearby if this room is for patient use. The room shall comply with Australian Standard 1428.1. Signage with braille tactile nominating 'Accessible Shower' required
SHPT | Shower - Patient | 4 | 24 Hours | 1 patient; 1-2 staff assisting as required |  | Sound attenuation level high, Acoustic privacy is required. | The Shower - Patient provides a shower and basin for use by patients (with or without assistance). Emergency release door hardware should be installed to allow staff access in an emergency. The entrance to the room should be flush with the adjoining floor for access in commode chairs/wheelchairs. Signage with braille tactile nominating 'Patient Shower' required
SHST | Shower - Staff | 3 | Available 24 Hours | 1 staff |  | Sound attenuation level high, Acoustic privacy is required. | A shower and changing area for use by staff. Additional Design Considerations: Secure, staff-only access to change rooms may be required if located in a public or patient care area. The Staff Shower may be gender specific or unisex.
SPECT-CT | SPECT-CT Imaging Room | 50 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | The SPECT-CT Imaging Room provides the area and facilities for Single Photon Emission Computed Tomography (SPECT) scanning procedures with an adjacent Control Room and Equipment Room. The SPECT-CT Imaging Room will be located with ready access to patient dosing areas. Cable pathways for services between the Control Room, Equipment Room and Imaging Room must be established and coordinated prior to construction and developed based on vendor requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access. The SPECT-CT Imaging Room must be adequately shielded to ensure that radiation doses to surrounding areas do not exceed radiation dose limits and satisfy design/dose constraints. Where shielding significantly increases the weight of the doors, automatic doors may be considered. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. Additional Considerations: Configuration of medical gases is to be confirmed to meet clinical requirements such as the provision of anaesthetic services. Dimmable lights are recommended and are to be placed to avoid glare to patients. Back lit ceiling artwork may be installed for patient comfort. 'X-ray in Use' lights are required outside each entry door. Room temperature and humidity control is required. Temperature requirements for equipment are to be confirmed by the vendor, including requirement for circulated cooling water. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. UPS power is required to the SPECT-CT scanning and control equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. Some models of SPECT/CT require three phase power to be supplied. Structural assessment is required to ensure the floor will be rigid with no vibration and will accommodate the weight of the SPECT-CT unit. Consideration needs to be given to the clearances of equipment and replacement part delivery routes. Where a service is may be performing minor interventional procedures, a Type A handwashing basin will be required instead of Type B.
SPECT-CTCR | SPECT-CT Control Room | 14 | Available 24 hours | Up to 3 staff |  | Acoustic privacy required | The SPECT-CT Control room is located adjacent to the SPECT-CT Imaging Room, with direct visibility to the patient, for remote control of equipment with audio and visual contact with the patient via cameras, speakers and microphones. Each Imaging Room requires a Control Room. Cable pathways for services between the Control Room, Equipment Room and Imaging Room must be established and coordinated prior to construction and developed based on vendor requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. Additional Considerations: The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. UPS power is required to the SPECT-CT scanning and control equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. Some models of SPECT/CT require three phase power to be supplied. Cable management and the potential for computers to be consolidated into a rack (either within the room or in a dedicated, collocated air conditioned room/cupboard) to be determined at project level. Provision of a height adjustable workstation may be considered depending on local policies and staff accessibility requirements.
SRM-15 | Staff Room, 15m2 | 15 | Available 24 Hours | Up to 8 staff, intermittently |  | Sound attenuation level high, Acoustic privacy is required. | The Staff Room is a shared facility for staff to use for respite, rest and relaxation during meal breaks, especially where it is difficult for staff to use common facilities particularly at night or due to distance. The room may also be used for small meetings or tutorials and for the storage of staff resources or library materials. Facilities for food and beverage preparation and storage should be provided. Additional Design Considerations: An external outlook is desirable. Staff only access may be required if this room is provided in a public or patient care area. A dishwasher may be provided depending on local operational polices. A side approach to the sink for wheelchair access has been shown; front access may be required depending on local/jurisdictional policies and staff requirements. The use of thermal heat detectors for fire detection is recommended. A portable fire extinguisher and fire blanket must be provided.
SRM-35 | Staff Room, 35m2 | 35 | Available 24 Hours | Up to 24 staff, intermittently |  | Sound attenuation level high, Acoustic privacy is required. | The Staff Room is a shared facility for staff to use for respite, rest and relaxation during meal breaks, especially where it is difficult for staff to use common facilities particularly at night or due to distance. The room may also be used for small meetings or tutorials and for the storage of staff resources or library materials. Facilities for food and beverage preparation and storage should be provided. Additional Design Considerations: An external outlook is desirable. Staff only access may be required if this room is provided in a public or patient care area. A dishwasher may be provided depending on local operational polices. Front access to the sink for wheelchair access has been shown; a side approach may be acceptable depending on local/jurisdictional policies and staff requirements. The use of thermal heat detectors for fire detection is recommended. A portable fire extinguisher and fire blanket must be provided.
SSTN-10 | Staff Station, 10m2 | 10 | 24 Hours | 2 workstations |  | Sound attenuation level medium, Acoustic privacy is required. | The Staff Station is the operational and administrative base of a unit and enquiry point for patients and visitors. It provides for the co-ordination of patient care, patient observation, write-up of clinical notes, data entry and making/receiving telephone calls. The Staff Station may also accommodate a receptionist/ward clerk. The number of workspaces and area (m2) requirements will vary depending on the type of department, activity level, number of full and part-time staff, operational models and building layout. One computer has been shown as a downtime computer with UPS power; this requirement will need to be confirmed by the department. The downtime computer, if required by a department, may alternatively be provided in a Clinical Workroom. Provision of a downtime computer will need to be rationalised across the unit to avoid duplication (e.g. if a Clinical Workroom and/or multiple Staff Stations are provided). Additional Design Considerations: A set down area may be provided for side access by visitors and patients in wheelchairs. If the Staff Station also serves a reception function, a recessed area for front access by wheelchair users may be required. The room may be enclosed (e.g. with security glazing and lockable doors) or other physical design/planning solutions may be considered (e.g. increased counter depth, egress to secure area) for staff safety and security as required in high risk areas. Standing-height workstations may be provided. A desktop printer and labels printer may be provided as required. Printing/scanning/photocopying facilities should be consolidated across the unit to suit service requirements. Additional monitors may be required for patient monitoring or security purposes depending on service requirements. Placement and visibility of nurse annunciators in adjacent corridors needs to be considered and coordinated with other ceiling elements such as bulkheads. Any rooms that are accessed from behind the staff station will require additional area for appropriate door circulation clearance. Wall mounted fire services items may be included in the Staff Station or provided nearby in a conspicuous and readily accessible position including a mimic panel, warden intercom phone, manual call point, break glass alarm, portable fire extinguisher and fire blanket. A ceiling mounted visual alarm device may also be provided in the Staff Station or an adjacent corridor. Inclusion of these items in the Staff Station will be dependent on the overall department layout. Location, mounting heights and clearances to meet AS1670.1 and AS2444.
SSTN-14 | Staff Station, 14m2 | 14 | 24 Hours | 3 workstations |  | Sound attenuation level medium, Acoustic privacy is required. | The Staff Station is the operational and administrative base of a unit and enquiry point for patients and visitors. It provides for the co-ordination of patient care, patient observation, write-up of clinical notes, data entry and making/receiving telephone calls. The Staff Station may also accommodate a receptionist/ward clerk. The number of workspaces and area (m2) requirements will vary depending on the type of department, activity level, number of full and part-time staff, operational models and building layout. One computer has been shown as a downtime computer with UPS power; this requirement will need to be confirmed by the department. The downtime computer, if required by a department, may alternatively be provided in a Clinical Workroom. Provision of a downtime computer will need to be rationalised across the unit to avoid duplication (e.g. if a Clinical Workroom and/or multiple Staff Stations are provided). Additional Design Considerations: A set down area may be provided for side access by visitors and patients in wheelchairs. If the Staff Station also serves a reception function, a recessed area for front access by wheelchair users may be required. The room may be enclosed (e.g. with security glazing and lockable doors) or other physical design/planning solutions may be considered (e.g. increased counter depth, egress to secure area) for staff safety and security as required in high risk areas. Standing-height workstations may be provided. A desktop printer and labels printer may be provided as required. Printing/scanning/photocopying facilities should be consolidated across the unit to suit service requirements. Additional monitors may be required for patient monitoring or security purposes depending on service requirements. Placement and visibility of nurse annunciators in adjacent corridors needs to be considered and coordinated with other ceiling elements such as bulkheads. Any rooms that are accessed from behind the staff station will require additional area for appropriate door circulation clearance. Wall mounted fire services items may be included in the Staff Station or provided nearby in a conspicuous and readily accessible position including a mimic panel, warden intercom phone, manual call point, break glass alarm, portable fire extinguisher and fire blanket. A ceiling mounted visual alarm device may also be provided in the Staff Station or an adjacent corridor. Inclusion of these items in the Staff Station will be dependent on the overall department layout. Location, mounting heights and clearances to meet AS1670.1 and AS2444.
SSTN-20 | Staff Station, 20m2 | 20 | 24 Hours | 4 workstations |  | Sound attenuation level medium, Acoustic privacy is required. | The Staff Station is the operational and administrative base of a unit and enquiry point for patients and visitors. It provides for the co-ordination of patient care, patient observation, write-up of clinical notes, data entry and making/receiving telephone calls. The Staff Station may also accommodate a receptionist/ward clerk. The number of workspaces and area (m2) requirements will vary depending on the type of department, activity level, number of full and part-time staff, operational models and building layout. One computer has been shown as a downtime computer with UPS power; this requirement will need to be confirmed by the department. The downtime computer, if required by a department, may alternatively be provided in a Clinical Workroom. Provision of a downtime computer will need to be rationalised across the unit to avoid duplication (e.g. if a Clinical Workroom and/or multiple Staff Stations are provided). Additional Design Considerations: A set down area may be provided for side access by visitors and patients in wheelchairs. If the Staff Station also serves a reception function, a recessed area for front access by wheelchair users may be required. The room may be enclosed (e.g. with security glazing and lockable doors) or other physical design/planning solutions may be considered (e.g. increased counter depth, egress to secure area) for staff safety and security as required in high risk areas. Standing-height workstations may be provided. A desktop printer and labels printer may be provided as required. Printing/scanning/photocopying facilities should be consolidated across the unit to suit service requirements. Additional monitors may be required for patient monitoring or security purposes depending on service requirements. Placement and visibility of nurse annunciators in adjacent corridors needs to be considered and coordinated with other ceiling elements such as bulkheads. Any rooms that are accessed from behind the staff station will require additional area for appropriate door circulation clearance. Wall mounted fire services items may be included in the Staff Station or provided nearby in a conspicuous and readily accessible position including a mimic panel, warden intercom phone, manual call point, break glass alarm, portable fire extinguisher and fire blanket. A ceiling mounted visual alarm device may also be provided in the Staff Station or an adjacent corridor. Inclusion of these items in the Staff Station will be dependent on the overall department layout. Location, mounting heights and clearances to meet AS1670.1 and AS2444.
STAD | Store - Accountable Drugs | 7.5 | 24 Hours | 1 - 2 intermittently |  |  | The Store &#8211; Accountable Drugs is a strongroom located within the Pharmacy Unit for the storage of accountable (Schedule 8) drugs and medications and provides a secure caged area for assembling orders. This room is to comply with jurisdictional legislation and public health facility policy, including any special provisions for refrigerated medicines. The strong room should be designed in collaboration with an expert security consultant with assessment of the risks for the particular site. Additional Design Considerations: The strong room should be independent of the rest of the building structure with no beams or columns passing through the ceiling, walls or floor. The ceiling walls and floor should be built utilisiing suitable construction methods, e.g. reinforced concrete or steel plate modular forms (the wall thickness shown in this Standard Component is indicative only). A gate door is to be provided for security when the main door is open. The security door should be torch and drill resistant, fitted with an appropriate locking mechanism incorporating duress features and an internal release mechanism in case persons are trapped within. The strongroom will require adequate ventilation and a torch and drill resistant breather tube must be fitted to comply with fire regulations. Security monitoring and detection equipment should be incorporated into the alarm system, to detect any attack on the structure. A telephone should be installed for staff safety
STBK | Store - Bulk | 30 | Up to 24 Hours | 2-3 staff, accessed intermittently |  |  | The Store - Bulk is the primary storage area for all delivered supplies and equipment prior to distribution to hospital services / departments. The Bulk Store should be located with ready access to the loading dock area. Some departments, such as pharmacy, renal dialysis and pathology, will have bulk storage spaces that are managed independently. Additional Considerations: Type and extent of storage solutions for this room will depend on the departments to be serviced and operational models for stock and equipment delivery. Where a 'direct to imprest' model is used, consumable inventory is stored at a supply chain distribution centre and delivered to the hospital already packed into trolleys specific to the clinical area and immediately delivered to those areas. Long term storage is not required and less area is needed near the loading dock. A 'central bulk store' model refers to goods being delivered to the hospital in large quantities and stored in a central bulk store. The process involves the delivery of goods on pallets, goods are then checked, unloaded, and placed in a bulk store by hospital supply staff. Consumables will then be picked and packed from the central store and delivered to clinical units and other areas as required. Automated or motorised equipment may be considered to assist with the safe and efficient transfer of supply trolleys throughout the healthcare facility. Where tugs and AGVs are being considered, it is important to confirm clearances, door widths and turning circles within the room, in the adjacent corridor and along path of travel to all services / departments. Refrigerated storage requirements to be confirmed at project level. Wall protection is recommended to protect wall linings and finishes from damage from impact. The height of the wall protection will be dependent on equipment and trolleys frequently moved through the space; a height range of 900mm minimum up to 1500mm is recommended depending on trolley and equipment selection. Hold open door hardware may be required to allow for easy movement of pallet jacks and trolleys in and out of the room. Controlled, staff only access is required.
STCL | Store - Cleaner | 12 | Up to 24 Hours | 1-2 staff, accessed intermittently |  |  | The Store - Cleaner is a store provided for the central storage of consumables and bulk containers of cleaning chemicals. Large facilities may provide a separate room for the storage of chemicals depending on quantities, safety requirements and operational policies for centralised storage and decanting of chemicals to units. Additional Considerations: Hold open door hardware may be required to allow for easy movement of equipment and trolleys in and out of the room. Type and extent of storage solutions for this room will depend on the departments to be serviced and operational models for cleaning and restocking consumables. Type of storage provided for cleaning chemicals, particularly where the storage of significant quantities of chemicals containing flammable ingredients, to be assessed to meet safety and service requirements. A clean linen trolley, or equivalent, may be provided for storage of clean mop heads that are laundered either onsite or offsite. The size of trolley will vary depending on service model and size of facility. Controlled, staff only access is required.
STEQ-14 | Store - Equipment, 14m2 | 14 | Available 24 hours | 1-2 staff, intermittently |  |  | A secure room for the storage and recharging of mobile equipment when not in use. Clear space is required for parking mobile equipment including IV poles, wheelchairs, patient transfer equipment, trolleys and commode chairs. Hooks may be provided for hanging equipment and rails may be provided to mount equipment whilst recharging. Additional Design Considerations: Minimum and maximum shelving heights to suit local WHS and infection control policies, as well as fire safety requirements (e.g. the top of stored products is to be no nearer than 500mm to sprinkler deflectors).
STEQ-20 | Store - Equipment, 20m2 | 20 | Available 24 hours | 1-2 staff, intermittently |  |  | A secure room for the storage and recharging of mobile equipment when not in use. Clear space is required for parking mobile equipment including IV poles, wheelchairs, patient transfer equipment, trolleys and commode chairs. Hooks may be provided for hanging equipment and rails may be provided to mount equipment whilst recharging. Additional Design Considerations: Minimum and maximum shelving heights to suit local WHS and infection control policies, as well as fire safety requirements (e.g. the top of stored products is to be no nearer than 500mm to sprinkler deflectors).
STFS-10 | Store - Files, 10m2 | 10 | Available 24 hours | 1-2 staff intermittently |  |  | The Files Store is a secure room for the storage of departmental files or patient clinical records. Additional Considerations: Type, size and quantity of shelving to be selected to suit service requirements and WHS policies Fire rating to room to be in accordance with NCC and local jurisdiction Records Acts.
STFS-20 | Store - Files, 20m2 | 20 | Available 24 hours | 1-2 staff intermittently |  |  | The Files Store is a secure room for the storage of departmental files or patient clinical records. Additional Considerations: Type, size and quantity of shelving to be selected to suit service requirements and WHS policies. If compactus shelving is installed, the track may be surface mounted or recessed into the floor. Structural Engineer's advice to be sought on floor load bearing capacity and potential deflection effect on walls, doors, etc. Handwheel operation may be required to reduce operator strain depending on block load, refer to local WHS policies. Fire rating to room to be in accordance with NCC and local jurisdiction Records Acts.
STGN | Store - General | 9 | Available 24 hours | 1-2 staff, intermittently |  |  | A secure room for the storage of general supplies used within the department. Additional Design Considerations: Facilities to support restocking workflows, including the use of barcode systems, to be considered to suit operational requirements. Minimum and maximum shelving heights to suit local WHS and infection control policies, as well as fire safety requirements (e.g. the top of stored products is to be no nearer than 500mm to sprinkler deflectors).
STRT | Stress Testing | 20 | Available 24 Hours | 1 patient; 1 - 2 staff |  | Acoustic privacy is required | Room for undertaking patient diagnostic stress testing procedures. The Stress Testing room should be located close to patient toilet and change rooms.
TRMT | Treatment Room | 14 | 24 Hours | 1 patient; 1-2 staff |  | Sound attenuation level high, Acoustic privacy is required. | The Treatment Room provides a controlled environment, privacy and facilities for carrying out consultations, examinations, treatments and minor procedures.
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | 28 | up to 24 Hours | 2 patients, 2 -3 staff |  | Sound attenuation level - high; acoustic privacy is required | The Treatment Room - Hand Therapy/Splinting provides a controlled environment, privacy and facilities for carrying out treatment and rehabilitation performed by an occupational therapist or physiotherapist with patients that suffer from conditions affecting the hands and upper limbs, such as wounds, scars, burns, fractures or amputations. Additional Considerations: A second door may be required for staff safety and security; the location of the second door will be dependent on the availability an appropriate area to egress. Depending on operational model, the hand therapy treatment room may be an open area adjoining a space with other allied health treatment bays. If open, floor boxes may be required for power to height adjustable tables. Printing facilities to be shared within the overall department. Refrigeration requirements for some materials in warmer climates may need to be considered.
TRMT-POD | Treatment Room - Podiatry | 14 | up to 24 Hours | 1 patient, 1 visitor/carer, 1 staff |  | Sound attenuation level high, Acoustic privacy is required. | The Treatment Room - Podiatry provides space and facilities for podiatry consultation, examination and treatment. The room may be used for measuring and fitting prosthetic appliances but not for fabrication. Additional Considerations: Medication storage, clean-up facilities for instruments, sink and plaster trap facilities are to be provided nearby. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. An external outlook is desirable.
UTRM | Uptake Room | 10 | 8am to 6pm, daily | 1 patient, 1 staff |  |  | The Uptake Room is used to inject patients with a radiopharmaceutical and allow them to rest until 'uptake' has occurred before transfer to the scanning room. Patients need to be relaxed in a warm, quiet and dimmed environment to promote uptake. Uptake duration is typically 45 to 60 minutes. Additional Considerations: Visual and audio observation of the patient during this time is typically provided from the Control Room of the PET-CT Imaging Room and intercoms, cameras and display screens are to be provided to support patient observation. Injections of radiopharmaceuticals can be undertaken within the room or from outside the room via an automated dose delivery system, which will be accommodated outside each uptake room with a suitable lead-lined injection port in the wall that will allow the coil to be connected to both the machine and the patient (details should be confirmed with vendor) Radiation shielding is required to the room and is to be in accordance with a radiation shielding consultants advice. This will be based on the type of radiopharmaceutical, the number of patients seen on a daily basis and proximity to other services.
WAIT-10 | Waiting, 10m2 | 10 | 24 Hours | Up to 8 |  | Sound attenuation level - high | An area for visitors, families and patients to wait prior to or during visits to a department. A range of occupants will need to be accommodated including adults, children, able-bodied persons and people with disabilities. In many instances bariatric capacity will also be required. The size of waiting areas can vary depending on service requirements; add 1.2m2 per additional person and allow 1.5m2 for additional wheelchair spaces as required. Waiting areas may be shared between departments. Public amenities including baby change facilities should be located nearby. An external outlook is desirable. Additional Considerations: Consider furniture options including weight rating of all chairs and providing built-in furniture or securing furniture where there may be safety/security concerns. Wall protection is recommended to protect wall linings and finishes from damage due to movement of loose furniture. Acoustic treatments may be considered as required to ensure the noise level in the space is comfortable both in the waiting area and adjacent reception areas. Check-in kiosks may be located nearby and services will need to be confirmed based on system requirements. Provision of a hearing augmentation system may be required where a built in amplification system is provided, if a TV monitor is installed. Hand hygiene facilities, such as alcohol-based hand rub, should be provided in line with local infection control policies.
WAIT-20 | Waiting, 20m2 | 20 | 24 Hours | Up to 17 |  | Sound attenuation level - high | An area for visitors, families and patients to wait prior to or during visits to a department. A range of occupants will need to be accommodated including adults, children, able-bodied persons and people with disabilities. In many instances bariatric capacity will also be required. The size of waiting areas can vary depending on service requirements; add 1.2m2 per additional person and allow 1.5m2 for additional wheelchair spaces as required. Waiting areas may be shared between departments. Public amenities including baby change facilities should be located nearby. An external outlook is desirable. Additional Considerations: Consider furniture options including weight rating of all chairs and providing built-in furniture or securing furniture where there may be safety/security concerns. Wall protection is recommended to protect wall linings and finishes from damage due to movement of loose furniture. Acoustic treatments may be considered as required to ensure the noise level in the space is comfortable both in the waiting area and adjacent reception areas. Check-in kiosks may be located nearby and services will need to be confirmed based on system requirements. Provision of a hearing augmentation system may be required where a built in amplification system is provided, if a TV monitor is installed. Hand hygiene facilities, such as alcohol-based hand rub, should be provided in line with local infection control policies.
WAIT-30 | Waiting, 30m2 | 30 | 24 Hours | Up to 25 |  | Sound attenuation level - high. | An area for visitors, families and patients to wait prior to or during visits to a department. A range of occupants will need to be accommodated including adults, children, able-bodied persons and people with disabilities. In many instances bariatric capacity will also be required. The size of waiting areas can vary depending on service requirements; add 1.2m2 per additional person and allow 1.5m2 for additional wheelchair spaces as required. Waiting areas may be shared between departments. Public amenities including baby change facilities should be located nearby. An external outlook is desirable. Additional Considerations: Consider furniture options including weight rating of all chairs and providing built-in furniture or securing furniture where there may be safety/security concerns. Wall protection is recommended to protect wall linings and finishes from damage due to movement of loose furniture. Acoustic treatments may be considered as required to ensure the noise level in the space is comfortable both in the waiting area and adjacent reception areas. Check-in kiosks may be located nearby and services will need to be confirmed based on system requirements. In large waiting areas (e.g. Emergency Department, Main Entry) provision of facilities for phone charging may be considered. Provision of a hearing augmentation system may be required where a built in amplification system is provided, if a TV monitor is installed. Hand hygiene facilities, such as alcohol-based hand rub, should be provided in line with local infection control policies.
WAIT-S | Waiting, Sub | 5 | 24 Hours | Up to 5 patient/visitors |  |  | The Waiting, Sub provides a satellite area for visitors, families and patients to wait within a unit, close to procedure/imaging rooms, consult rooms or offices. The size of satellite waiting areas can vary depending on service requirements; add 1.2m2 per additional person and allow 1.5m2 for additional wheelchair spaces as required. Additional Considerations: Wall protection is recommended to protect wall linings and finishes from damage due to movement of loose furniture. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. Hand hygiene facilities, such as alcohol-based hand rub, should be provided in line with local infection control policies.
WCAC | Toilet - Accessible | 6 | 24 Hours | 1 |  | Sound attenuation level high, Acoustic privacy is required. | The Toilet - Accessible provides a toilet and basin for use by patients, visitors or staff with disabilities (with or without assistance). A baby change table may also be located in this room. If the Toilet - Accessible is located in a patient care area the nurse call system should be included and emergency release door hardware should be installed to allow staff access in an emergency. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system. The entrance to the room should be flush with the adjoining floor. The room shall comply with Australian Standard 1428.1. Signage with braille tactile nominating 'Unisex Accessible WC' required (including RH or LH dependent on layout).
WCPT | Toilet - Patient | 4 | 24 Hours | 1 patient; 1 staff assisting as required |  | Sound attenuation level high, Acoustic privacy is required. | The Toilet - Patient provides a toilet and basin for use by patients (with or without assistance). Emergency release door hardware should be installed to allow staff access in an emergency. The entrance to the room should be flush with the adjoining floor for access in commode chairs/wheelchairs. Signage with braille tactile nominating 'Unisex Patient WC' required
WCPU | Toilet - Public | 3 | 24 Hours | 1 |  | Sound attenuation level high, Acoustic privacy is required. | The Toilet - Public provides a toilet and basin for use by the general public. Signage with braille tactile nominating 'Unisex Public WC' required
WCST | Toilet - Staff | 3 | 24 Hours | 1 |  | Sound attenuation level high, Acoustic privacy is required. | The Toilet - Staff provides a toilet and basin for use by staff. Staff only access may be required if this room is provided in a public or patient care area. Signage with braille tactile nominating 'Unisex Staff WC' required

## Items
Standard Components - Items in room
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
1BR-CCU | 1 Bed Room - CCU | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
1BR-CCU | 1 Bed Room - CCU | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 11
1BR-CCU | 1 Bed Room - CCU | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
1BR-CCU | 1 Bed Room - CCU | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
1BR-CCU | 1 Bed Room - CCU | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 2 | to basin
1BR-CCU | 1 Bed Room - CCU | FA | WLPR-002 | WALL PROTECTION: bed protection wall panel, composite, laminate finish | 1
1BR-CCU | 1 Bed Room - CCU | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | ceiling tiles may also be acceptable dependent upon the acuity of patients to be accommodated and procedures to be performed within the room
1BR-CCU | 1 Bed Room - CCU | FA | DWWH-005 | WINDOW | 1 | external, natural light and outlook essential
1BR-CCU | 1 Bed Room - CCU | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | extent of glazing to suit service requirements
1BR-CCU | 1 Bed Room - CCU | FA | DWOB-001 | OBSERVATION: internal window | 1 | optional; extent of glazing to suit service requirements
1BR-CCU | 1 Bed Room - CCU | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
1BR-CCU | 1 Bed Room - CCU | FFE | MGFP-002 | FLOWMETER: medical air | 1
1BR-CCU | 1 Bed Room - CCU | FFE | FIHR-019 | BRACKET: for suction bottle | 2
1BR-CCU | 1 Bed Room - CCU | FFE | FIHR-016 | BRACKET: for sharps bin | 1 | optional, mobile sharps bin may be provided to be shared between multiple rooms
1BR-CCU | 1 Bed Room - CCU | FFE | MMGE-019 | CHAIR: patient | 1
1BR-CCU | 1 Bed Room - CCU | FFE | FQBS-030 | CHAIR: visitor, patient | 1
1BR-CCU | 1 Bed Room - CCU | FFE | DWWD-004 | BLIND: roller | 1 | to external window
1BR-CCU | 1 Bed Room - CCU | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | to door and observation window
1BR-CCU | 1 Bed Room - CCU | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
1BR-CCU | 1 Bed Room - CCU | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1 | to patient entertainment screen; configuration will be dependent on operational and ICT policies
1BR-CCU | 1 Bed Room - CCU | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
1BR-CCU | 1 Bed Room - CCU | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
1BR-CCU | 1 Bed Room - CCU | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
1BR-CCU | 1 Bed Room - CCU | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
1BR-CCU | 1 Bed Room - CCU | FFE | FQSN-051 | LOCKER: 2 drawers, open under, lockable, mobile | 1 | type and extent of storage for patient property to suit operational policies
1BR-CCU | 1 Bed Room - CCU | FFE | FQTA-013 | TABLE: overbed | 1
1BR-CCU | 1 Bed Room - CCU | FFE | MGFP-004 | FLOWMETER: oxygen | 2
1BR-CCU | 1 Bed Room - CCU | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1 | optional, mobile sharps bin may be provided to be shared between multiple rooms
1BR-CCU | 1 Bed Room - CCU | FFE | MMGE-002 | ADAPTER: suction | 2
1BR-CCU | 1 Bed Room - CCU | FFE | MMGE-100 | SUCTION BOTTLE | 2
1BR-CCU | 1 Bed Room - CCU | FFE | HYBA-010 | BASIN: handwash, Type B | 1
1BR-CCU | 1 Bed Room - CCU | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
1BR-CCU | 1 Bed Room - CCU | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1 | configuration will be dependent on operational and ICT policies
1BR-CCU | 1 Bed Room - CCU | FFE | MMBE-010 | BED: patient, electric with mattress | 1
1BR-CCU | 1 Bed Room - CCU | FFE | MMHA-005 | MONITOR: patient, cardiac | 1 | telemetry provision to suit service requirements
1BR-CCU | 1 Bed Room - CCU | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | optional, to patient locker; provision dependent on operational policies
1BR-CCU | 1 Bed Room - CCU | FFE | FIJO-209 | CUPBOARD: tall, single door, adj shelves and clothes rail, recessed | 1 | type and extent of storage for patient property to suit operational policies
1BR-CCU | 1 Bed Room - CCU | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1
1BR-CCU | 1 Bed Room - CCU | FFE | PPPP-130 | TRACK: curtain, privacy, S shaped, angled | 1 | to door and observation window
1BR-CCU | 1 Bed Room - CCU | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to patient monitor
1BR-CCU | 1 Bed Room - CCU | SE | ELPR-006 | RCD: residual current device | 3 | on services panel
1BR-CCU | 1 Bed Room - CCU | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment screen, may be replaced with data outlet in accordance with ICT policy
1BR-CCU | 1 Bed Room - CCU | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to patient entertainment system
1BR-CCU | 1 Bed Room - CCU | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel, with indicator button and light
1BR-CCU | 1 Bed Room - CCU | SE | LIFX-011 | LIGHT: reading light | 1 | switch to patient handset and services panel
1BR-CCU | 1 Bed Room - CCU | SE | LIFX-017 | LIGHT: night light | 1 | switch to patient handset and wall near entry
1BR-CCU | 1 Bed Room - CCU | SE | ELBO-015 | SWITCH: light, single | 3 | 1 to general room lighting, 1 to night lighting, 1 on services panel to reading light
1BR-CCU | 1 Bed Room - CCU | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2
1BR-CCU | 1 Bed Room - CCU | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
1BR-CCU | 1 Bed Room - CCU | SE | MGAS-062 | OUTLET: suction, on services panel | 2
1BR-CCU | 1 Bed Room - CCU | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on services panel, with indicator button and light
1BR-CCU | 1 Bed Room - CCU | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on services panel
1BR-CCU | 1 Bed Room - CCU | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | above services panel, with indicator button and light
1BR-CCU | 1 Bed Room - CCU | SE | ELGP-109 | GPO: single, wall mounted | 1 | to bedhead; low level for electric bed
1BR-CCU | 1 Bed Room - CCU | SE | ELGP-109 | GPO: single, wall mounted | 1 | adjacent to patient chair
1BR-CCU | 1 Bed Room - CCU | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | to patient monitor
1BR-CCU | 1 Bed Room - CCU | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
1BR-CCU | 1 Bed Room - CCU | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
1BR-CCU | 1 Bed Room - CCU | SE | ELBO-005 | EQUIPOTENTIAL JUNCTION BOX | 1
1BR-CCU | 1 Bed Room - CCU | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | quantity to suit service requirements
1BR-CCU | 1 Bed Room - CCU | SE | ELGP-122 | GPO: emergency power, single, on services panel | 9 | type and quantity to suit service requirements
ADLB | ADL Bathroom | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
ADLB | ADL Bathroom | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | significant change in colour from adjacent corridor to be avoided, to be suitable for bare feet, tiles are also acceptable
ADLB | ADL Bathroom | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
ADLB | ADL Bathroom | FA | WLWA-019 | WALL FINISH: vinyl, full height | 6 | tiles also acceptable
ADLB | ADL Bathroom | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1 | water resistant flush set ceiling is also acceptable
ADLB | ADL Bathroom | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 2
ADLB | ADL Bathroom | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1 | with privacy latch and emergency release
ADLB | ADL Bathroom | FA | FLGE-019 | FLOOR HOB: Shower | 1 | optional
ADLB | ADL Bathroom | FA | CLCN-009 | CORNICE: wall trim fixing | 1
ADLB | ADL Bathroom | FFE | MMGE-024 | CHAIR: shower | 1 | commode chair or combined shower/commode chair is also acceptable
ADLB | ADL Bathroom | FFE | FIHR-026 | HOOK: coat | 4
ADLB | ADL Bathroom | FFE | FIRT-034 | TRACK: curtain, shower, 90 deg | 2 | optional, dependent on infection control policy
ADLB | ADL Bathroom | FFE | FQGE-011 | CURTAIN: shower | 2 | optional, dependent on infection control policy.
ADLB | ADL Bathroom | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
ADLB | ADL Bathroom | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
ADLB | ADL Bathroom | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
ADLB | ADL Bathroom | FFE | FIDI-036 | DISPENSER: toilet paper | 1
ADLB | ADL Bathroom | FFE | FIRT-006 | GRAB RAIL: WC | 1 | 1 is optional, dependent on service requirements
ADLB | ADL Bathroom | FFE | FIRT-003 | GRAB RAIL: shower, horizontal | 1 | may be combined with rail for hand held shower
ADLB | ADL Bathroom | FFE | FIRT-008 | GRAB RAIL: WC, hinged | 1 | optional, dependent on service requirements and expected patient cohort
ADLB | ADL Bathroom | FFE | FIBM-006 | MIRROR: fixed, glass | 1 | to basin
ADLB | ADL Bathroom | FFE | FIBM-007 | MIRROR: fixed, glass, full length | 1 | dressing
ADLB | ADL Bathroom | FFE | FQCW-037 | BIN: waste, sanitary | 1 | optional, dependent on waste management policy
ADLB | ADL Bathroom | FFE | FIDI-051 | HOLDER: soap/shampoo, recessed, vertical format | 2
ADLB | ADL Bathroom | FFE | FQBS-008 | BENCH SEAT: freestanding, patient transfer | 1 | portable, for over bath
ADLB | ADL Bathroom | FFE | HYBA-025 | BATH: domestic | 1 | warm water
ADLB | ADL Bathroom | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1
ADLB | ADL Bathroom | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 2 | diverter required for shower over bath
ADLB | ADL Bathroom | FFE | HYTP-020 | SHOWER: hand-held | 1 | vertical rail to have grabrail strength and fixings
ADLB | ADL Bathroom | FFE | HYWC-006 | TOILET SUITE: patient | 1 | 100mm clearout access point optional
ADLB | ADL Bathroom | FFE | HYTP-021 | SHOWER: hand-held, over bath | 1 | vertical rail to have grabrail strength and fixings
ADLB | ADL Bathroom | FFE | HYBA-002 | BASIN: ensuite, with vanity unit | 1 | shelf for patient belongings may be integral or separate
ADLB | ADL Bathroom | FFE | HYTP-028 | SPOUT: bath, wall mounted | 1
ADLB | ADL Bathroom | FFE | FIRT-044 | GRAB RAIL: bath, vertical | 1
ADLB | ADL Bathroom | FFE | FIRT-048 | GRAB RAIL: L-shaped | 1
ADLB | ADL Bathroom | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
ADLB | ADL Bathroom | SE | HYDR-004 | DRAIN: floor waste | 2
ADLB | ADL Bathroom | SE | ELBO-015 | SWITCH: light, single | 1
ADLB | ADL Bathroom | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 3 | optional; + indicator button & light
ADLB | ADL Bathroom | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
ADLB | ADL Bathroom | SE | ELGP-209 | GPO: double, wall mounted, RCD protected | 1
ADLD | ADL Dining | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
ADLD | ADL Dining | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
ADLD | ADL Dining | FA | WLPR-006 | WALL PROTECTION: corner guards | 2
ADLD | ADL Dining | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
ADLD | ADL Dining | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 3
ADLD | ADL Dining | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
ADLD | ADL Dining | FA | DWWH-005 | WINDOW | 1 | desirable
ADLD | ADL Dining | FA | CLCN-009 | CORNICE: wall trim fixing | 1
ADLD | ADL Dining | FFE | FQGE-002 | ARTWORK | 1
ADLD | ADL Dining | FFE | FQBS-015 | CHAIR: dining | 3
ADLD | ADL Dining | FFE | DWWD-004 | BLIND: roller | 1
ADLD | ADL Dining | FFE | FQTA-085 | TABLE: square, 1100W, adjustable height | 1 | height adjustable between 750-850H
ADLD | ADL Dining | SE | ELPR-006 | RCD: residual current device | 1
ADLD | ADL Dining | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light, may be rationalised to central location if collocated with other ADL Rooms
ADLD | ADL Dining | SE | ELBO-015 | SWITCH: light, single | 1
ADLD | ADL Dining | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light, may be rationalised to central location if collocated with other ADL Rooms
ADLD | ADL Dining | SE | ELGP-208 | GPO: double, wall mounted | 2
ADLD | ADL Dining | SE | ITSE-006 | BUTTON: duress, fixed | 1
ADLK | ADL Kitchen | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent around sink
ADLK | ADL Kitchen | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 8
ADLK | ADL Kitchen | FA | WLWA-007 | WALL FINISH: splashback | 2
ADLK | ADL Kitchen | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
ADLK | ADL Kitchen | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | plasterboard flush set ceiling is also acceptable
ADLK | ADL Kitchen | FA | DWWH-005 | WINDOW | 1 | desirable
ADLK | ADL Kitchen | FA | CLCN-009 | CORNICE: wall trim fixing | 1
ADLK | ADL Kitchen | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 2
ADLK | ADL Kitchen | FFE | DWWD-004 | BLIND: roller | 1
ADLK | ADL Kitchen | FFE | FIJO-131 | CUPBOARD: tall, single door, adj shelves, lockable | 1
ADLK | ADL Kitchen | FFE | FQSM-040 | PEDESTAL: mobile, drawer unit, lockable, 660H | 1 | under bench
ADLK | ADL Kitchen | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 1
ADLK | ADL Kitchen | FFE | FIJO-156 | CUPBOARD: wall mount, single door, adj shelves | 2
ADLK | ADL Kitchen | FFE | FQCW-016 | BIN: waste, general, 20L | 1
ADLK | ADL Kitchen | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to sink
ADLK | ADL Kitchen | FFE | FIDI-030 | DISPENSER: soap | 1 | to sink
ADLK | ADL Kitchen | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
ADLK | ADL Kitchen | FFE | FQCA-046 | KETTLE: electric | 1
ADLK | ADL Kitchen | FFE | FQCA-067 | OVEN: microwave, domestic | 1 | on bench
ADLK | ADL Kitchen | FFE | FQCA-063 | OVEN: electric | 1 | retractable/side opening doors
ADLK | ADL Kitchen | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
ADLK | ADL Kitchen | FFE | FQCA-072 | RANGEHOOD | 1 | over cooktop
ADLK | ADL Kitchen | FFE | FQCA-031 | FRIDGE: domestic, 250L | 1
ADLK | ADL Kitchen | FFE | HYBA-041 | SINK: double, with drainer, half size rinse bowl | 1 | heat lined, warm water
ADLK | ADL Kitchen | FFE | FQCA-080 | TOASTER: domestic | 1
ADLK | ADL Kitchen | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 2
ADLK | ADL Kitchen | FFE | FIJO-205 | CUPBOARD: drawer unit, single drawer, lockable | 6
ADLK | ADL Kitchen | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead also acceptable
ADLK | ADL Kitchen | FFE | FIJO-213 | PANEL: joinery carcass end | 3
ADLK | ADL Kitchen | FFE | FIJO-223 | CUPBOARD: tall, oven cabinet | 1
ADLK | ADL Kitchen | FFE | FQCA-089 | COOKTOP: induction | 1 | controls to be located close to front
ADLK | ADL Kitchen | FFE | FIJO-224 | PANEL: joinery fixed panel | 5 | to sink and cooktop
ADLK | ADL Kitchen | SE | ELPR-006 | RCD: residual current device | 1
ADLK | ADL Kitchen | SE | ELBO-007 | SWITCH: emergency isolation | 1 | to cooktop and oven
ADLK | ADL Kitchen | SE | ELBO-015 | SWITCH: light, single | 1
ADLK | ADL Kitchen | SE | ELGP-109 | GPO: single, wall mounted | 2 | to range hood and fridge
ADLK | ADL Kitchen | SE | ELGP-208 | GPO: double, wall mounted | 4
ADLK | ADL Kitchen | SE | ELGE-001 | DIRECT CONNECTION: power | 2 | to cooktop and oven
ADLL | ADL Laundry | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
ADLL | ADL Laundry | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1
ADLL | ADL Laundry | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
ADLL | ADL Laundry | FA | WLWA-007 | WALL FINISH: splashback | 1
ADLL | ADL Laundry | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
ADLL | ADL Laundry | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | plasterboard flush set ceiling is also acceptable
ADLL | ADL Laundry | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
ADLL | ADL Laundry | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
ADLL | ADL Laundry | FA | CLCN-009 | CORNICE: wall trim fixing | 1
ADLL | ADL Laundry | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 2
ADLL | ADL Laundry | FFE | FIJO-156 | CUPBOARD: wall mount, single door, adj shelves | 1
ADLL | ADL Laundry | FFE | FQCW-016 | BIN: waste, general, 20L | 1
ADLL | ADL Laundry | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to sink
ADLL | ADL Laundry | FFE | FIDI-030 | DISPENSER: soap | 1 | to sink
ADLL | ADL Laundry | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
ADLL | ADL Laundry | FFE | FQSM-030 | TROLLEY: laundry, basket | 1 | optional
ADLL | ADL Laundry | FFE | FQGE-015 | DRYER: clothes, domestic | 1
ADLL | ADL Laundry | FFE | FQGE-028 | IRON: clothes, electric | 1
ADLL | ADL Laundry | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
ADLL | ADL Laundry | FFE | HYBA-064 | SINK: single, laundry | 1 | with overflow drain, warm water
ADLL | ADL Laundry | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1 | with inset laundry tub
ADLL | ADL Laundry | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead also acceptable
ADLL | ADL Laundry | FFE | MMAH-039 | PLINTH: laundry | 2
ADLL | ADL Laundry | FFE | FIJO-222 | CUPBOARD: tall, single door, fixed shelf, ADL Laundry | 1
ADLL | ADL Laundry | FFE | FQCW-063 | MACHINE: washing, domestic, front loading | 1
ADLL | ADL Laundry | SE | ELPR-006 | RCD: residual current device | 1 | type 1 10mA in accordance with AS/NZS 3003
ADLL | ADL Laundry | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light, may be rationalised to central location if collocated with other ADL Rooms
ADLL | ADL Laundry | SE | HYTP-008 | OUTLET: cold water | 1 | to washing machine
ADLL | ADL Laundry | SE | HYDR-004 | DRAIN: floor waste | 1
ADLL | ADL Laundry | SE | ELBO-015 | SWITCH: light, single | 1
ADLL | ADL Laundry | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light, may be rationalised to central location if collocated with other ADL Rooms
ADLL | ADL Laundry | SE | ELGP-109 | GPO: single, wall mounted | 2 | to washing machine and dryer
ADLL | ADL Laundry | SE | ELGP-208 | GPO: double, wall mounted | 1 | 30mA RCD protection
AHBF | After Hours Blood Fridge | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
AHBF | After Hours Blood Fridge | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
AHBF | After Hours Blood Fridge | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
AHBF | After Hours Blood Fridge | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
AHBF | After Hours Blood Fridge | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
AHBF | After Hours Blood Fridge | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | optional, to laboratory if adjacent
AHBF | After Hours Blood Fridge | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1
AHBF | After Hours Blood Fridge | FA | CLCN-009 | CORNICE: wall trim fixing | 1
AHBF | After Hours Blood Fridge | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
AHBF | After Hours Blood Fridge | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
AHBF | After Hours Blood Fridge | FFE | ITNE-003 | BAR CODING: reader | 1
AHBF | After Hours Blood Fridge | FFE | FIJO-017 | BENCH: 600D, open under | 1 | laboratory grade
AHBF | After Hours Blood Fridge | FFE | MMLA-053 | FRIDGE: blood products, single door | 1 | sections may be lockable; BMS monitored, audible alarm 24hr; size is to suit service requirements
AHBF | After Hours Blood Fridge | FFE | FIJO-213 | PANEL: joinery carcass end | 1
AHBF | After Hours Blood Fridge | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to fridge for temperature monitoring
AHBF | After Hours Blood Fridge | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
AHBF | After Hours Blood Fridge | SE | ITBU-001 | ALARM: equipment monitoring, temperature | 1 | may be integral to fridge
AHBF | After Hours Blood Fridge | SE | ELBO-015 | SWITCH: light, single | 1
AHBF | After Hours Blood Fridge | SE | ELGP-208 | GPO: double, wall mounted | 2
AHBF | After Hours Blood Fridge | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to fridge
AHBF | After Hours Blood Fridge | SE | ITSE-038 | READER: access control, general | 2 | 1 to entry, 1 is optional to laboratory door if provided
AHDR | After Hours Drug Store | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
AHDR | After Hours Drug Store | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
AHDR | After Hours Drug Store | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
AHDR | After Hours Drug Store | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
AHDR | After Hours Drug Store | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
AHDR | After Hours Drug Store | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | from corridor
AHDR | After Hours Drug Store | FA | DWGL-007 | DOOR: 910mm c/o, single leaf, observation panel, solid core/safety glass, paint | 1 | optional, if room located with direct access from pharmacy
AHDR | After Hours Drug Store | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
AHDR | After Hours Drug Store | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
AHDR | After Hours Drug Store | FFE | FIJO-141 | CUPBOARD: underbench, single door, adj shelves, lockable | 1
AHDR | After Hours Drug Store | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 2
AHDR | After Hours Drug Store | FFE | ITSE-012 | CAMERA: CCTV | 1 | optional, provision and placement dependent on local policies
AHDR | After Hours Drug Store | FFE | MMGE-036 | FRIDGE: drugs, underbench | 1
AHDR | After Hours Drug Store | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
AHDR | After Hours Drug Store | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
AHDR | After Hours Drug Store | FFE | FQSM-045 | SHELVING: storage solution, 400W x 600D x 2100H nom., wire, mobile | 2 | storage solution/mix of storage options to be tailored to suit project requirements
AHDR | After Hours Drug Store | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to drug fridge, for temperature monitoring
AHDR | After Hours Drug Store | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
AHDR | After Hours Drug Store | SE | ELBO-015 | SWITCH: light, single | 1
AHDR | After Hours Drug Store | SE | ELGP-208 | GPO: double, wall mounted | 2
AHDR | After Hours Drug Store | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to drug fridge
AHDR | After Hours Drug Store | SE | ITSE-006 | BUTTON: duress, fixed | 1 | optional
AHDR | After Hours Drug Store | SE | ITSE-038 | READER: access control, general | 2
AIRLE-12 | Airlock - Entry, 12m2 | FA | FLGE-016 | FLOOR MAT: recessed, composite slip resistant | 1
AIRLE-12 | Airlock - Entry, 12m2 | FA | FLSK-008 | SKIRTING: to suit glazing system over | 6
AIRLE-12 | Airlock - Entry, 12m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
AIRLE-12 | Airlock - Entry, 12m2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
AIRLE-12 | Airlock - Entry, 12m2 | FA | DWWH-005 | WINDOW | 4 | safety glass
AIRLE-12 | Airlock - Entry, 12m2 | FA | DWGL-050 | DOOR: 1600mm c/o, double leaf, fully glazed, alum/safety glass, powdercoat/paint | 2 | automated doors, provision to hold open doors via key switch to be considered
AIRLE-12 | Airlock - Entry, 12m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
AIRLE-12 | Airlock - Entry, 12m2 | FFE | FIHR-002 | BRACKET: for CCTV camera, ceiling mounted | 1
AIRLE-12 | Airlock - Entry, 12m2 | FFE | ITSE-012 | CAMERA: CCTV | 1 | as required for access and security to suit operational policies
AIRLE-12 | Airlock - Entry, 12m2 | FFE | MEEX-018 | AIR CURTAIN: external door | 1 | optional
AIRLE-12 | Airlock - Entry, 12m2 | FFE | FIDI-055 | DISPENSER: umbrella cover | 1 | optional
AIRLE-12 | Airlock - Entry, 12m2 | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 1 | to camera
AIRLE-12 | Airlock - Entry, 12m2 | SE | ITSE-024 | DOOR RELEASE: access control, emergency override | 1
AIRLE-12 | Airlock - Entry, 12m2 | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1 | to camera
AIRLE-12 | Airlock - Entry, 12m2 | SE | ITSE-049 | INTERCOM: external, audiovisual, wall mounted | 1 | as required for access and security to suit operational policies, may be located to internal door
AIRLE-12 | Airlock - Entry, 12m2 | SE | ITSE-050 | READER: access control, external | 1 | as required for access and security to suit operational policies
AIRLE-12 | Airlock - Entry, 12m2 | SE | ITSE-051 | BUTTON: push to exit | 1 | may be within airlock to external door or nearby adjacent to internal door
AIRLE-6 | Airlock - Entry, 6m2 | FA | FLGE-016 | FLOOR MAT: recessed, composite slip resistant | 1
AIRLE-6 | Airlock - Entry, 6m2 | FA | FLSK-008 | SKIRTING: to suit glazing system over | 6
AIRLE-6 | Airlock - Entry, 6m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
AIRLE-6 | Airlock - Entry, 6m2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
AIRLE-6 | Airlock - Entry, 6m2 | FA | DWWH-005 | WINDOW | 2 | safety glass
AIRLE-6 | Airlock - Entry, 6m2 | FA | DWGL-026 | DOOR: 1200mm c/o, single leaf, sliding, fixed panel, fully glazed | 2 | automated doors, provision to hold open doors via key switch to be considered
AIRLE-6 | Airlock - Entry, 6m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
AIRLE-6 | Airlock - Entry, 6m2 | FFE | FIHR-002 | BRACKET: for CCTV camera, ceiling mounted | 1
AIRLE-6 | Airlock - Entry, 6m2 | FFE | ITSE-012 | CAMERA: CCTV | 1 | as required for access and security to suit operational policies
AIRLE-6 | Airlock - Entry, 6m2 | FFE | MEEX-018 | AIR CURTAIN: external door | 1 | optional
AIRLE-6 | Airlock - Entry, 6m2 | FFE | FIDI-055 | DISPENSER: umbrella cover | 1 | optional
AIRLE-6 | Airlock - Entry, 6m2 | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 1 | to camera
AIRLE-6 | Airlock - Entry, 6m2 | SE | ITSE-024 | DOOR RELEASE: access control, emergency override | 1
AIRLE-6 | Airlock - Entry, 6m2 | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1 | to camera
AIRLE-6 | Airlock - Entry, 6m2 | SE | ITSE-049 | INTERCOM: external, audiovisual, wall mounted | 1 | as required for access and security to suit operational policies, may be located to internal door
AIRLE-6 | Airlock - Entry, 6m2 | SE | ITSE-050 | READER: access control, external | 1 | as required for access and security to suit operational policies
AIRLE-6 | Airlock - Entry, 6m2 | SE | ITSE-051 | BUTTON: push to exit | 1 | may be within airlock to external door or nearby adjacent to internal door
AUD-CR | Audiology Control Room | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
AUD-CR | Audiology Control Room | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 6
AUD-CR | Audiology Control Room | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 2
AUD-CR | Audiology Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
AUD-CR | Audiology Control Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
AUD-CR | Audiology Control Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 2 | acoustic seals to suit testing requirements
AUD-CR | Audiology Control Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
AUD-CR | Audiology Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
AUD-CR | Audiology Control Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
AUD-CR | Audiology Control Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
AUD-CR | Audiology Control Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
AUD-CR | Audiology Control Room | FFE | FIDI-021 | DISPENSER: mask | 1
AUD-CR | Audiology Control Room | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1
AUD-CR | Audiology Control Room | FFE | MMAH-005 | BOOTH: audiometric | 1 | one way window (non-reflective); acoustic seals to suit testing requirements
AUD-CR | Audiology Control Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | optional
AUD-CR | Audiology Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
AUD-CR | Audiology Control Room | FFE | FQDW-031 | WORKSTATION: straight, adj height, 1650W x 750D x 750H nom. | 1
AUD-CR | Audiology Control Room | FFE | MMAH-036 | CONSOLE: control, audiometric booth | 1 | optional, dependent on service requirements
AUD-CR | Audiology Control Room | FFE | MMAH-042 | DEVICE: tympanometer | 1
AUD-CR | Audiology Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
AUD-CR | Audiology Control Room | SE | ELPR-006 | RCD: residual current device | 1
AUD-CR | Audiology Control Room | SE | ELGP-208 | GPO: double, wall mounted | 3 | provision of emergency power to be determined based on service requirements
AUD-CR | Audiology Control Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
AUD-CR | Audiology Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
AUD-CR | Audiology Control Room | SE | MMSP-032 | SERVICES PANEL: instrument jack | 1 | optional
AUD-TR | Audiology Testing Room | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
AUD-TR | Audiology Testing Room | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
AUD-TR | Audiology Testing Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | acoustic seals to suit testing requirements
AUD-TR | Audiology Testing Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
AUD-TR | Audiology Testing Room | FFE | FQBS-012 | CHAIR: child | 1
AUD-TR | Audiology Testing Room | FFE | FQBS-014 | CHAIR: client | 2
AUD-TR | Audiology Testing Room | FFE | FQBS-039 | LOUNGE: 1 seat | 1
AUD-TR | Audiology Testing Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
AUD-TR | Audiology Testing Room | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 1
AUD-TR | Audiology Testing Room | FFE | FQDW-001 | DESK: adj height, 1500W x 750D x 750H nom. | 1
AUD-TR | Audiology Testing Room | FFE | FIHR-029 | HOOK: equipment | 2 | for headphones
AUD-TR | Audiology Testing Room | FFE | FQTA-003 | TABLE: children's play | 1
AUD-TR | Audiology Testing Room | FFE | MMAH-005 | BOOTH: audiometric | 1 | with one way window (non-reflective)
AUD-TR | Audiology Testing Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
AUD-TR | Audiology Testing Room | FFE | MMAH-036 | CONSOLE: control, audiometric booth | 1
AUD-TR | Audiology Testing Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
AUD-TR | Audiology Testing Room | FFE | ITAV-061 | SPEAKER: audiology | 2 | wall mounted, calibrated for testing space
AUD-TR | Audiology Testing Room | FFE | MMAH-043 | DEVICE: visual reinforcement audiometry | 2 | wall mounted, may be provided as a screen or puppet box, inclusion is dependent on service requirements
AUD-TR | Audiology Testing Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
AUD-TR | Audiology Testing Room | SE | ELPR-006 | RCD: residual current device | 1
AUD-TR | Audiology Testing Room | SE | ELGP-208 | GPO: double, wall mounted | 6 | provision of emergency power to be determined based on service requirements
AUD-TR | Audiology Testing Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
AUD-TR | Audiology Testing Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
AUD-TR | Audiology Testing Room | SE | MMSP-032 | SERVICES PANEL: instrument jack | 4 | optional, as required for remote operation of equipment
BATH | Bathroom | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
BATH | Bathroom | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
BATH | Bathroom | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
BATH | Bathroom | FA | WLWA-019 | WALL FINISH: vinyl, full height | 6 | tiles are also acceptable
BATH | Bathroom | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
BATH | Bathroom | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1 | privacy latch
BATH | Bathroom | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
BATH | Bathroom | FFE | MMGE-024 | CHAIR: shower | 1
BATH | Bathroom | FFE | FIHR-026 | HOOK: coat | 2
BATH | Bathroom | FFE | FIRT-034 | TRACK: curtain, shower, 90 deg | 1 | optional, refer to local infection control policy
BATH | Bathroom | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | to entry door
BATH | Bathroom | FFE | FQGE-011 | CURTAIN: shower | 1 | optional, refer to local infection control policy
BATH | Bathroom | FFE | FQCW-016 | BIN: waste, general, 20L | 1
BATH | Bathroom | FFE | FQSM-033 | TROLLEY: shower | 1 | height adjustable bath may be provided
BATH | Bathroom | FFE | FIDI-014 | DISPENSER: disposable glove | 1
BATH | Bathroom | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
BATH | Bathroom | FFE | FIDI-030 | DISPENSER: soap | 1
BATH | Bathroom | FFE | FIDI-036 | DISPENSER: toilet paper | 1 | to drop down grab rail
BATH | Bathroom | FFE | FIRT-008 | GRAB RAIL: WC, hinged | 2 | one with toilet paper dispenser, one with nurse call pendant
BATH | Bathroom | FFE | FIBM-006 | MIRROR: fixed, glass | 1
BATH | Bathroom | FFE | FQCW-037 | BIN: waste, sanitary | 1
BATH | Bathroom | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 1 | or soap dispenser
BATH | Bathroom | FFE | FIRT-016 | RAIL: towel | 1 | with grab rail strength/fixings
BATH | Bathroom | FFE | HYBA-005 | BASIN: handwash, accessible | 1
BATH | Bathroom | FFE | HYWC-001 | TOILET SUITE: accessible | 1 | with backrest
BATH | Bathroom | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 2
BATH | Bathroom | FFE | HYTP-020 | SHOWER: hand-held | 1 | to shower
BATH | Bathroom | FFE | HYTP-035 | TAPWARE: mixer, basin, accessible | 1
BATH | Bathroom | FFE | HYTP-045 | SHOWER: hand-held, accessible | 1
BATH | Bathroom | FFE | FIRT-045 | GRAB RAIL: shower, horizontal and vertical | 1
BATH | Bathroom | FFE | FIGE-011 | BACKREST: accessible toilet | 1
BATH | Bathroom | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 1
BATH | Bathroom | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | IPX4 rating required
BATH | Bathroom | SE | HYDR-004 | DRAIN: floor waste | 2 | special floor drain required if height adjustable bath provided
BATH | Bathroom | SE | ELBO-015 | SWITCH: light, single | 1
BATH | Bathroom | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | IPX4 rating required
BATH | Bathroom | SE | ITCL-031 | PENDANT: nurse call, patient to staff | 1 | to toilet, IPX4 rating required, additional required for height adjustable bath if provided
BATH | Bathroom | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | IPX4 rating required
BATH | Bathroom | SE | ELGP-208 | GPO: double, wall mounted | 1 | IPX4 rating required
BBEV | Bay - Beverage | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10/pendulum P3 or agreed equivalent around sink; coordinated with adjacent area, significant change in colour to be avoided
BBEV | Bay - Beverage | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | coordinated with adjacent area
BBEV | Bay - Beverage | FA | WLPR-010 | WALL PROTECTION: corner guards, to 900 AFFL | 1 | coordinated with adjacent wall protection
BBEV | Bay - Beverage | FA | WLWA-007 | WALL FINISH: splashback | 2 | to back of bench and return wall
BBEV | Bay - Beverage | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BBEV | Bay - Beverage | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | coordinated with adjacent area
BBEV | Bay - Beverage | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BBEV | Bay - Beverage | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
BBEV | Bay - Beverage | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 1 | optional; configuration of joinery to meet storage requirements
BBEV | Bay - Beverage | FFE | FIJO-104 | CUPBOARD: drawer unit, 3 equal | 2
BBEV | Bay - Beverage | FFE | FQCW-019 | BIN: waste, general, 50L | 1 | size, type and quantity of bins dependent on local waste management policies
BBEV | Bay - Beverage | FFE | HYGE-003 | BOILING/CHILLED WATER UNIT: Combination | 1 | provide grille in joinery for ventilation
BBEV | Bay - Beverage | FFE | HYTP-029 | TAPWARE: boiling/chilled water unit (comb) | 1 | instant hot/cold water tap to sink
BBEV | Bay - Beverage | FFE | FQCA-015 | DISPENSER: ice, potable | 1 | optional; item to comply with infection control requirements
BBEV | Bay - Beverage | FFE | FQCA-067 | OVEN: microwave, domestic | 1
BBEV | Bay - Beverage | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
BBEV | Bay - Beverage | FFE | FQCA-032 | FRIDGE: domestic, 300L | 1 | clear area for fridge to allow for adequate ventilation and full opening of door
BBEV | Bay - Beverage | FFE | FIJO-069 | BENCH TOP: 750D, joinery under, laminate finish | 1 | minimum depth 600mm; where an ice machine is provided the depth may need to increase to accommodate
BBEV | Bay - Beverage | FFE | FIJO-163 | CUPBOARD: tall, single microwave | 1
BBEV | Bay - Beverage | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead is also acceptable
BBEV | Bay - Beverage | FFE | FIJO-212 | BULKHEAD: joinery | 1 | optional, provide if wall mounted joinery provided, plasterboard bulkhead is also acceptable
BBEV | Bay - Beverage | FFE | FQCW-062 | BIN: waste, recycling, comingled, 50L | 1 | size, type and quantity of bins dependent on local waste management policies
BBEV | Bay - Beverage | FFE | HYBA-080 | SINK: single, with drainer, accessible | 1
BBEV | Bay - Beverage | SE | HYTP-008 | OUTLET: cold water | 1 | to boiling/ chilled water unit
BBEV | Bay - Beverage | SE | ELBO-015 | SWITCH: light, single | 1 | optional, lighting and controls to be coordinated with adjacent area
BBEV | Bay - Beverage | SE | ELGP-109 | GPO: single, wall mounted | 3 | 1 to fridge, 1 to microwave, 1 to boiling/ chilled water unit
BBEV | Bay - Beverage | SE | ELGP-208 | GPO: double, wall mounted | 2 | over bench
BBLD | Bay - Blood Fridge | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BBLD | Bay - Blood Fridge | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4 | coordinated with adjacent area
BBLD | Bay - Blood Fridge | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
BBLD | Bay - Blood Fridge | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BBLD | Bay - Blood Fridge | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | coordinated with adjacent area
BBLD | Bay - Blood Fridge | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1 | coordinated with adjacent area
BBLD | Bay - Blood Fridge | FFE | MMLA-053 | FRIDGE: blood products, single door | 2 | sections may be lockable; BMS monitored, audible alarm 24hr; quantity and size is to suit service requirements
BBLD | Bay - Blood Fridge | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | for temperature monitoring
BBLD | Bay - Blood Fridge | SE | ITBU-001 | ALARM: equipment monitoring, temperature | 2 | may be integral to fridge
BBLD | Bay - Blood Fridge | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 2
BBW | Bay - Blanket / Fluid Warmer | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BBW | Bay - Blanket / Fluid Warmer | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | coordinated with adjacent area
BBW | Bay - Blanket / Fluid Warmer | FA | WLPR-006 | WALL PROTECTION: corner guards | 1 | coordinated with adjacent wall protection
BBW | Bay - Blanket / Fluid Warmer | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BBW | Bay - Blanket / Fluid Warmer | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area, flush set ceiling is also acceptable
BBW | Bay - Blanket / Fluid Warmer | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BBW | Bay - Blanket / Fluid Warmer | FFE | MMGE-010 | CABINET: blanket/ fluid warming | 1 | for fluids only, blankets only or a combination of fluids and blankets; configuration to suit service requirements
BBW | Bay - Blanket / Fluid Warmer | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | optional; provide for connection to BMS where temperature monitoring required for fluids; alternative connection may be required depending on BMS or equipment requirements
BBW | Bay - Blanket / Fluid Warmer | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
BES | Bay - Emergency Shower | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
BES | Bay - Emergency Shower | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1
BES | Bay - Emergency Shower | FA | WLPR-011 | WALL PROTECTION: corner guards, to 900 AFFL min. | 1
BES | Bay - Emergency Shower | FA | WLWA-019 | WALL FINISH: vinyl, full height | 1
BES | Bay - Emergency Shower | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
BES | Bay - Emergency Shower | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
BES | Bay - Emergency Shower | FFE | HYTP-023 | SHOWER: safety, eyewash combination | 1
BES | Bay - Emergency Shower | SE | HYDR-004 | DRAIN: floor waste | 1
BHW | Bay - Height/ Weight | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BHW | Bay - Height/ Weight | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | coordinated with adjacent area
BHW | Bay - Height/ Weight | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
BHW | Bay - Height/ Weight | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BHW | Bay - Height/ Weight | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 3 | height dependent on type and height of mobile equipment, crash rails are also acceptable
BHW | Bay - Height/ Weight | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area
BHW | Bay - Height/ Weight | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BHW | Bay - Height/ Weight | FFE | MMGE-089 | SCALES: electronic, chair | 1 | optional, dependent on service requirements
BHW | Bay - Height/ Weight | FFE | MMGE-188 | SCALES: electronic, standing, with stadiometer | 1 | bariatric capacity will be dependent on service requirements; stadiometer may be separate wall mounted or floor standing item; scales may include integral grab rails for support
BHW | Bay - Height/ Weight | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional, dependent on equipment requirements
BHW | Bay - Height/ Weight | SE | ELGP-208 | GPO: double, wall mounted | 1
BHWS-A | Bay - Handwashing, Type A | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4
BHWS-A | Bay - Handwashing, Type A | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | if bay is recessed
BHWS-A | Bay - Handwashing, Type A | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | floor finish of adjacent area must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin. significant change in colour to be avoided
BHWS-A | Bay - Handwashing, Type A | FA | WLWA-007 | WALL FINISH: splashback | 3
BHWS-A | Bay - Handwashing, Type A | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BHWS-A | Bay - Handwashing, Type A | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | ceiling tiles are also acceptable, ceiling to be coordinated with adjacent area
BHWS-A | Bay - Handwashing, Type A | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
BHWS-A | Bay - Handwashing, Type A | FFE | FQCW-016 | BIN: waste, general, 20L | 1
BHWS-A | Bay - Handwashing, Type A | FFE | FIDI-014 | DISPENSER: disposable glove | 1
BHWS-A | Bay - Handwashing, Type A | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
BHWS-A | Bay - Handwashing, Type A | FFE | FIDI-030 | DISPENSER: soap | 1
BHWS-A | Bay - Handwashing, Type A | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1
BHWS-A | Bay - Handwashing, Type A | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
BHWS-A | Bay - Handwashing, Type A | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
BHWS-B | Bay - Handwashing, Type B | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
BHWS-B | Bay - Handwashing, Type B | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | if bay is recessed
BHWS-B | Bay - Handwashing, Type B | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | floor finish of adjacent area must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin. significant change in colour to be avoided
BHWS-B | Bay - Handwashing, Type B | FA | WLWA-007 | WALL FINISH: splashback | 3
BHWS-B | Bay - Handwashing, Type B | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BHWS-B | Bay - Handwashing, Type B | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | ceiling tiles are also acceptable, ceiling to be coordinated with adjacent area
BHWS-B | Bay - Handwashing, Type B | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
BHWS-B | Bay - Handwashing, Type B | FFE | FQCW-016 | BIN: waste, general, 20L | 1
BHWS-B | Bay - Handwashing, Type B | FFE | FIDI-014 | DISPENSER: disposable glove | 1
BHWS-B | Bay - Handwashing, Type B | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
BHWS-B | Bay - Handwashing, Type B | FFE | FIDI-030 | DISPENSER: soap | 1
BHWS-B | Bay - Handwashing, Type B | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | optional, dependent on service requirements
BHWS-B | Bay - Handwashing, Type B | FFE | HYBA-010 | BASIN: handwash, Type B | 1
BHWS-B | Bay - Handwashing, Type B | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
BIRM-A | Birthing Room, Without Pool | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
BIRM-A | Birthing Room, Without Pool | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
BIRM-A | Birthing Room, Without Pool | FA | WLPR-006 | WALL PROTECTION: corner guards | 3
BIRM-A | Birthing Room, Without Pool | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
BIRM-A | Birthing Room, Without Pool | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BIRM-A | Birthing Room, Without Pool | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FA | WLPR-002 | WALL PROTECTION: bed protection wall panel, composite, laminate finish | 1
BIRM-A | Birthing Room, Without Pool | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
BIRM-A | Birthing Room, Without Pool | FA | DWWH-005 | WINDOW | 1
BIRM-A | Birthing Room, Without Pool | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | with hold open device to support bed movement as required
BIRM-A | Birthing Room, Without Pool | FA | DWGE-015 | DOOR: concertina folding leaf | 1 | 2000 wide
BIRM-A | Birthing Room, Without Pool | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
BIRM-A | Birthing Room, Without Pool | FFE | MGFP-002 | FLOWMETER: medical air | 2 | 1 to bedhead, 1 to resuscitaire
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-001 | BASSINET | 1 | optional
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | 1 to bedhead cupboard, 1 to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-016 | BRACKET: for sharps bin | 1 | mounted on trolley
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | FFE | DWWD-004 | BLIND: roller | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-156 | CUPBOARD: wall mount, single door, adj shelves | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-019 | DISPENSER: gown/apron | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-014 | DISPENSER: disposable glove | 2 | 1 to basin, 1 to room
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQSN-051 | LOCKER: 2 drawers, open under, lockable, mobile | 1 | type and extent of storage for patient property to suit operational policies
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-044 | LOUNGE: 2 seat, sofa bed | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQTA-013 | TABLE: overbed | 1
BIRM-A | Birthing Room, Without Pool | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | 1 to bedhead, 1 to resuscitaire
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-011 | MIXER: oxygen and nitrous oxide | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-194 | SHELF: laminate | 2 | to equipment store cupboard, type and quantity of storage options to be determined based on service requirements
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-051 | STOOL: mobile, adj height | 1 | for suturing and small procedures
BIRM-A | Birthing Room, Without Pool | FFE | MMGE-002 | ADAPTER: suction | 2 | 1 to bedhead cupboard, 1 to resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | FFE | MMGE-100 | SUCTION BOTTLE | 2 | 1 to bedhead cupboard, 1 to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | FFE | FQSM-016 | TROLLEY | 1 | to equipment store cupboard
BIRM-A | Birthing Room, Without Pool | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
BIRM-A | Birthing Room, Without Pool | FFE | MMBE-005 | BED: patient, birthing, electric | 1
BIRM-A | Birthing Room, Without Pool | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
BIRM-A | Birthing Room, Without Pool | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | to staff write-up, dual monitors may be provided as required to suit operational, EMR and ICT requirements; a workstation on wheels is also acceptable
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-015 | RESUSCITAIRE: infant, mobile | 1 | Mobile resuscitaire shown. Item may be wall mounted where suitable product is available.
BIRM-A | Birthing Room, Without Pool | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
BIRM-A | Birthing Room, Without Pool | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | optional, to patient locker
BIRM-A | Birthing Room, Without Pool | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-129 | CUPBOARD: tall, single door, adj shelves | 1 | type and extent of storage to suit service requirements
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-099 | CAC: cable access cap | 1 | to staff write-up bench
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-162 | CUPBOARD: tall, single door, to conceal services | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-209 | CUPBOARD: tall, single door, adj shelves and clothes rail, recessed | 1 | wardrobe, type and extent of storage for patient property to suit operational policies
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-212 | BULKHEAD: joinery | 3
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-213 | PANEL: joinery carcass end | 1 | below FIJO-156 to separate dirty linen trolley from hand wash, storage for clean and dirty linen to suit local operational policies
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-068 | BRACKET: for alcohol based hand rub, attached to bed | 1
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-031 | DEVICE: patient, fetal, cardiotocography (CTG), on trolley | 1 | to equipment store cupboard
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-043 | SUPPORT FIXTURE: birthing, hook, ceiling mounted | 1 | optional, provision dependent on birthing aids/supportive equipment to be used
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-233 | CUPBOARD: tall, infant resuscitaire | 1 | Configuration of joinery to suit resuscitare equipment selected, including consideration of maneuverability and provision of sufficient space for emergency resuscitation procedures. Confirm design with users at project level.
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-041 | SUPPORT FIXTURE: birthing, ladder, wall mounted | 1
BIRM-A | Birthing Room, Without Pool | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to staff write-up for wall mounted phone
BIRM-A | Birthing Room, Without Pool | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | SE | ELPR-006 | RCD: residual current device | 3 | 2 to bedhead, 1 to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELPR-006 | RCD: residual current device | 2 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment system, may be replaced with data outlet depending on ICT policies
BIRM-A | Birthing Room, Without Pool | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to patient entertainment system
BIRM-A | Birthing Room, Without Pool | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | with indicator button and light, to bedhead
BIRM-A | Birthing Room, Without Pool | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | LIFX-011 | LIGHT: reading light | 1 | to bedhead
BIRM-A | Birthing Room, Without Pool | SE | LIFX-017 | LIGHT: night light | 1 | to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ELBO-015 | SWITCH: light, single | 1 | to bedhead MSP for reading light
BIRM-A | Birthing Room, Without Pool | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | with indicator button and light, to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light, to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level for patient bed
BIRM-A | Birthing Room, Without Pool | SE | ELGP-208 | GPO: double, wall mounted | 3 | 2 to staff write-up, 1 to lounge
BIRM-A | Birthing Room, Without Pool | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
BIRM-A | Birthing Room, Without Pool | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
BIRM-A | Birthing Room, Without Pool | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
BIRM-A | Birthing Room, Without Pool | SE | ITSE-006 | BUTTON: duress, fixed | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | SE | MMSP-026 | SERVICES PANEL: wall mounted | 2 | 1 to bedhead, 1 inside bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELBO-012 | SWITCH: light, dimmer | 2 | 1 to entry door, 1 to staff write up for general room lights
BIRM-A | Birthing Room, Without Pool | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | 1 to bedhead, 1 to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELGP-122 | GPO: emergency power, single, on services panel | 6 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELGP-122 | GPO: emergency power, single, on services panel | 8 | 5 to bedhead, 3 to bedhead cupboard
BIRM-B | Birthing Room, With Pool | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
BIRM-B | Birthing Room, With Pool | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
BIRM-B | Birthing Room, With Pool | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to birthing pool
BIRM-B | Birthing Room, With Pool | FA | WLPR-006 | WALL PROTECTION: corner guards | 2
BIRM-B | Birthing Room, With Pool | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
BIRM-B | Birthing Room, With Pool | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BIRM-B | Birthing Room, With Pool | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin
BIRM-B | Birthing Room, With Pool | FA | WLPR-002 | WALL PROTECTION: bed protection wall panel, composite, laminate finish | 1
BIRM-B | Birthing Room, With Pool | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
BIRM-B | Birthing Room, With Pool | FA | DWWH-005 | WINDOW | 2
BIRM-B | Birthing Room, With Pool | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | with hold open device to support bed movement as required
BIRM-B | Birthing Room, With Pool | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
BIRM-B | Birthing Room, With Pool | FA | DWGE-017 | DOOR: concertina folding leaf | 1 | 1800 wide
BIRM-B | Birthing Room, With Pool | FFE | MGFP-002 | FLOWMETER: medical air | 2 | 1 to bedhead, 1 to resuscitaire
BIRM-B | Birthing Room, With Pool | FFE | FQBS-001 | BASSINET | 1 | optional
BIRM-B | Birthing Room, With Pool | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1
BIRM-B | Birthing Room, With Pool | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | 1 to bedhead, 1 to resuscitaire
BIRM-B | Birthing Room, With Pool | FFE | FIHR-016 | BRACKET: for sharps bin | 1 | mounted on trolley
BIRM-B | Birthing Room, With Pool | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1 | to staff write-up
BIRM-B | Birthing Room, With Pool | FFE | DWWD-004 | BLIND: roller | 2
BIRM-B | Birthing Room, With Pool | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
BIRM-B | Birthing Room, With Pool | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
BIRM-B | Birthing Room, With Pool | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 1 | to staff write-up
BIRM-B | Birthing Room, With Pool | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
BIRM-B | Birthing Room, With Pool | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
BIRM-B | Birthing Room, With Pool | FFE | FIDI-019 | DISPENSER: gown/apron | 1
BIRM-B | Birthing Room, With Pool | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
BIRM-B | Birthing Room, With Pool | FFE | FIDI-014 | DISPENSER: disposable glove | 2 | 1 to basin, 1 to room
BIRM-B | Birthing Room, With Pool | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
BIRM-B | Birthing Room, With Pool | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
BIRM-B | Birthing Room, With Pool | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
BIRM-B | Birthing Room, With Pool | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
BIRM-B | Birthing Room, With Pool | FFE | FQSN-051 | LOCKER: 2 drawers, open under, lockable, mobile | 1 | type and extent of storage for patient property to suit operational policies
BIRM-B | Birthing Room, With Pool | FFE | FQBS-044 | LOUNGE: 2 seat, sofa bed | 1
BIRM-B | Birthing Room, With Pool | FFE | FQTA-013 | TABLE: overbed | 1
BIRM-B | Birthing Room, With Pool | FFE | MGFP-004 | FLOWMETER: oxygen | 3 | 1 to bedhead cupboard, 1 to resuscitaire cupboard, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | FFE | MMMA-011 | MIXER: oxygen and nitrous oxide | 1 | to bedhead cupboard
BIRM-B | Birthing Room, With Pool | FFE | FIJO-194 | SHELF: laminate | 2 | to equipment store cupboard, type and quantity of storage options to be determined based on service requirements
BIRM-B | Birthing Room, With Pool | FFE | FQBS-051 | STOOL: mobile, adj height | 1 | for suturing and small procedures
BIRM-B | Birthing Room, With Pool | FFE | MMGE-002 | ADAPTER: suction | 2 | 1 to bedhead cupboard, 1 to resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | FFE | MMGE-100 | SUCTION BOTTLE | 2 | 1 to bedhead cupboard, 1 to resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | FFE | FQSM-016 | TROLLEY | 1 | to equipment store cupboard
BIRM-B | Birthing Room, With Pool | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
BIRM-B | Birthing Room, With Pool | FFE | MMBE-005 | BED: patient, birthing, electric | 1
BIRM-B | Birthing Room, With Pool | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
BIRM-B | Birthing Room, With Pool | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
BIRM-B | Birthing Room, With Pool | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | to staff write-up, dual monitors may be provided as required to suit operational, EMR and ICT requirements; a workstation on wheels is also acceptable
BIRM-B | Birthing Room, With Pool | FFE | MMMA-015 | RESUSCITAIRE: infant, mobile | 1
BIRM-B | Birthing Room, With Pool | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
BIRM-B | Birthing Room, With Pool | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 2 | 1 to shower, 1 to spout for birthing pool
BIRM-B | Birthing Room, With Pool | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | optional, to patient locker
BIRM-B | Birthing Room, With Pool | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1 | to staff write-up
BIRM-B | Birthing Room, With Pool | FFE | FIJO-129 | CUPBOARD: tall, single door, adj shelves | 1 | type and extent of storage to suit service requirements
BIRM-B | Birthing Room, With Pool | FFE | HYTP-021 | SHOWER: hand-held, over bath | 1
BIRM-B | Birthing Room, With Pool | FFE | HYBA-023 | BATH: birthing | 1 | Birthing pools can vary significantly in shape and the pool shown is indicative only. Final selection will have implications for required area and clearances. Ergonomic assessment of birthing pool options should be undertaken prior to selection.
BIRM-B | Birthing Room, With Pool | FFE | FIJO-099 | CAC: cable access cap | 1 | to staff write-up bench
BIRM-B | Birthing Room, With Pool | FFE | FIJO-162 | CUPBOARD: tall, single door, to conceal services | 1
BIRM-B | Birthing Room, With Pool | FFE | FIJO-209 | CUPBOARD: tall, single door, adj shelves and clothes rail, recessed | 1 | wardrobe, type and extent of storage for patient property to suit operational policies
BIRM-B | Birthing Room, With Pool | FFE | FIJO-212 | BULKHEAD: joinery | 5
BIRM-B | Birthing Room, With Pool | FFE | FIJO-213 | PANEL: joinery carcass end | 2 | 1 below FIJO-156 to separate dirty linen trolley from hand wash, storage for clean and dirty linen to suit local operational policies, 1 below write-up bench
BIRM-B | Birthing Room, With Pool | FFE | FIHR-068 | BRACKET: for alcohol based hand rub, attached to bed | 1
BIRM-B | Birthing Room, With Pool | FFE | MMMA-031 | DEVICE: patient, fetal, cardiotocography (CTG), on trolley | 1 | to equipment store cupboard
BIRM-B | Birthing Room, With Pool | FFE | MMMA-043 | SUPPORT FIXTURE: birthing, hook, ceiling mounted | 1 | optional, provision dependent on birthing aids/supportive equipment to be used
BIRM-B | Birthing Room, With Pool | FFE | FIJO-233 | CUPBOARD: tall, infant resuscitaire | 1 | Configuration of joinery to suit resuscitare equipment selected, including consideration of maneuverability and provision of sufficient space for emergency resuscitation procedures. Confirm design with users at project level.
BIRM-B | Birthing Room, With Pool | FFE | MMMA-041 | SUPPORT FIXTURE: birthing, ladder, wall mounted | 1
BIRM-B | Birthing Room, With Pool | FFE | HYTP-048 | SPOUT: birthing pool, high flow, wall mounted | 1
BIRM-B | Birthing Room, With Pool | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to staff write-up for wall mounted phone
BIRM-B | Birthing Room, With Pool | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to staff write-up
BIRM-B | Birthing Room, With Pool | SE | ELPR-006 | RCD: residual current device | 3 | 2 to bedhead, 1 to bedhead cupboard
BIRM-B | Birthing Room, With Pool | SE | ELPR-006 | RCD: residual current device | 2 | to infant resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment system, may be replaced with data outlet depending on ICT policies
BIRM-B | Birthing Room, With Pool | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to patient entertainment system
BIRM-B | Birthing Room, With Pool | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | with indicator button and light, 1 to bedhead, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | SE | MGAS-053 | OUTLET: scavenge, on services panel | 2 | 1 to bedhead cupboard, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | SE | HYDR-004 | DRAIN: floor waste | 1 | to birthing pool area
BIRM-B | Birthing Room, With Pool | SE | LIFX-011 | LIGHT: reading light | 1 | to bedhead
BIRM-B | Birthing Room, With Pool | SE | LIFX-017 | LIGHT: night light | 1 | to bedhead
BIRM-B | Birthing Room, With Pool | SE | ELBO-015 | SWITCH: light, single | 1 | to bedhead for reading light
BIRM-B | Birthing Room, With Pool | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | to infant resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 3 | 2 to bedhead cupboard, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 2 | 1 to bedhead cupboard, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | to infant resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | to bedhead cupboard
BIRM-B | Birthing Room, With Pool | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | to infant resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | to bedhead cupboard
BIRM-B | Birthing Room, With Pool | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | with indicator button and light
BIRM-B | Birthing Room, With Pool | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | to bedhead
BIRM-B | Birthing Room, With Pool | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light, 1 to bedhead, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level for patient bed
BIRM-B | Birthing Room, With Pool | SE | ELGP-208 | GPO: double, wall mounted | 3 | 2 to staff write-up, 1 to lounge
BIRM-B | Birthing Room, With Pool | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
BIRM-B | Birthing Room, With Pool | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
BIRM-B | Birthing Room, With Pool | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
BIRM-B | Birthing Room, With Pool | SE | ITSE-006 | BUTTON: duress, fixed | 2 | 1 to staff write-up, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | SE | MMSP-026 | SERVICES PANEL: wall mounted | 4 | 2 to bedhead, 1 to infant resuscitaire cupboard, 1 to birthing pool
BIRM-B | Birthing Room, With Pool | SE | ELBO-012 | SWITCH: light, dimmer | 2 | 1 to entry door, 1 to staff write up for general room lights
BIRM-B | Birthing Room, With Pool | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 | to infant resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | 1 to bedhead, 1 to bedhead cupboard
BIRM-B | Birthing Room, With Pool | SE | ELGP-122 | GPO: emergency power, single, on services panel | 6 | to infant resuscitaire cupboard
BIRM-B | Birthing Room, With Pool | SE | ELGP-122 | GPO: emergency power, single, on services panel | 8 | 5 to bedhead, 3 to bedhead cupboard
BIRM-B | Birthing Room, With Pool | SE | HYGE-028 | DEVICE: backflow prevention | 1 | to birthing pool area
BLIN | Bay - Linen | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BLIN | Bay - Linen | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | coordinated with adjacent area
BLIN | Bay - Linen | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
BLIN | Bay - Linen | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BLIN | Bay - Linen | FA | WLWA-021 | WALL FINISH: vinyl, to 1800 AFFL | 1
BLIN | Bay - Linen | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area, flush set ceiling is also acceptable
BLIN | Bay - Linen | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BLIN | Bay - Linen | FFE | FIHR-029 | HOOK: equipment | 3 | optional, for storage of linen bags
BLIN | Bay - Linen | FFE | MMGE-109 | TROLLEY: clean linen | 1 | with cover as required to suit infection prevention and control policies
BLIN | Bay - Linen | FFE | FIJO-129 | CUPBOARD: tall, single door, adj shelves | 1 | optional, for storage of additional items such as pillows, privacy curtains, shower curtains, etc.
BLIN | Bay - Linen | FFE | FIJO-212 | BULKHEAD: joinery | 1 | optional, to joinery cupboard if provided, plasterboard bulkhead is also acceptable
BLIN | Bay - Linen | SE | ELBO-015 | SWITCH: light, single | 1 | optional, lighting and controls to be coordinated with adjacent area, may be provided if bay is enclosed
BMEQ | Bay - Mobile Equipment | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BMEQ | Bay - Mobile Equipment | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4 | coordinated with adjacent area
BMEQ | Bay - Mobile Equipment | FA | WLPR-006 | WALL PROTECTION: corner guards | 1 | coordinated with adjacent wall protection
BMEQ | Bay - Mobile Equipment | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BMEQ | Bay - Mobile Equipment | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 3 | height dependent on type and height of mobile equipment, crash rails are also acceptable
BMEQ | Bay - Mobile Equipment | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area
BMEQ | Bay - Mobile Equipment | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BMEQ | Bay - Mobile Equipment | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1 | optional
BMEQ | Bay - Mobile Equipment | FFE | MMBE-042 | SLIDE: patient | 1 | optional
BMEQ | Bay - Mobile Equipment | SE | ELGP-208 | GPO: double, wall mounted | 2 | type, quantity and configuration of outlets to be determined based on equipment requirements
BMFD-3 | Bay - Multifunction Device, 3m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
BMFD-3 | Bay - Multifunction Device, 3m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
BMFD-3 | Bay - Multifunction Device, 3m2 | FA | WLPR-006 | WALL PROTECTION: corner guards | 2
BMFD-3 | Bay - Multifunction Device, 3m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BMFD-3 | Bay - Multifunction Device, 3m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
BMFD-3 | Bay - Multifunction Device, 3m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
BMFD-3 | Bay - Multifunction Device, 3m2 | FFE | FQCW-038 | BIN: waste, secure | 1 | type and size of bins is to suit waste management policies and may be rationalised with bins provided in adjacent area
BMFD-3 | Bay - Multifunction Device, 3m2 | FFE | FQCW-029 | BIN: waste, recycling, paper, 240L | 1 | type and size of bins is to suit waste management policies and may be rationalised with bins provided in adjacent area
BMFD-3 | Bay - Multifunction Device, 3m2 | FFE | ITNE-021 | MULTIFUNCTION DEVICE: floorstanding | 1 | recommended manufacturer clearances zones to be considered, bins and other moveable items may be within clearance zones
BMFD-3 | Bay - Multifunction Device, 3m2 | FFE | FIJO-129 | CUPBOARD: tall, single door, adj shelves | 1 | stationary storage to suit service requirements, mobile is storage also acceptable
BMFD-3 | Bay - Multifunction Device, 3m2 | FFE | FIJO-212 | BULKHEAD: joinery | 1
BMFD-3 | Bay - Multifunction Device, 3m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
BMFD-3 | Bay - Multifunction Device, 3m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
BMFD-7 | Bay - Multifunction Device, 7m2 | FA | WLPR-006 | WALL PROTECTION: corner guards | 2
BMFD-7 | Bay - Multifunction Device, 7m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
BMFD-7 | Bay - Multifunction Device, 7m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 2
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FQCW-038 | BIN: waste, secure | 1 | type and size of bins is to suit waste management policies and may be rationalised with bins provided in adjacent area
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 2
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FQCW-029 | BIN: waste, recycling, paper, 240L | 1 | type and size of bins is to suit waste management policies and may be rationalised with bins provided in adjacent area
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | ITNE-021 | MULTIFUNCTION DEVICE: floorstanding | 1 | recommended manufacturer clearances zones to be considered, bins and other moveable items may be within clearance zones
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FIJO-129 | CUPBOARD: tall, single door, adj shelves | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | FFE | FIJO-212 | BULKHEAD: joinery | 2
BMFD-7 | Bay - Multifunction Device, 7m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
BMFD-7 | Bay - Multifunction Device, 7m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
BMT | Bay - Meal Trolley | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BMT | Bay - Meal Trolley | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | coordinated with adjacent area
BMT | Bay - Meal Trolley | FA | WLPR-006 | WALL PROTECTION: corner guards | 1 | coordinated with adjacent wall protection
BMT | Bay - Meal Trolley | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BMT | Bay - Meal Trolley | FA | WLWA-021 | WALL FINISH: vinyl, to 1800 AFFL | 1
BMT | Bay - Meal Trolley | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area, flush set ceiling is also acceptable
BMT | Bay - Meal Trolley | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BMT | Bay - Meal Trolley | FFE | FQCA-082 | TROLLEY: food delivery, hot/cold | 2 | type of meal trolley will be dependent on food service delivery model
BMT | Bay - Meal Trolley | SE | ELGP-109 | GPO: single, wall mounted | 2 | power requirements will be dependent on type of meal trolley
BMT | Bay - Meal Trolley | SE | ELGP-140 | GPO: three phase, single | 2 | optional, power requirements will be dependent on type of meal trolley
BPATH | Bay - Pathology Point of Care Testing (POCT) | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BPATH | Bay - Pathology Point of Care Testing (POCT) | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9 | coordinated with adjacent area
BPATH | Bay - Pathology Point of Care Testing (POCT) | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
BPATH | Bay - Pathology Point of Care Testing (POCT) | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BPATH | Bay - Pathology Point of Care Testing (POCT) | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area
BPATH | Bay - Pathology Point of Care Testing (POCT) | FA | WLWA-009 | WALL FINISH: splashback, laminate | 3 | coved, integral to bench
BPATH | Bay - Pathology Point of Care Testing (POCT) | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FIHR-016 | BRACKET: for sharps bin | 1
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FIJO-141 | CUPBOARD: underbench, single door, adj shelves, lockable | 1 | lockable; extent of storage to suit testing requirements
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | quantity and type of bins to be based on local waste management policies
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 1 | lockable; extent of storage to suit testing requirements
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1 | with wide opening
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1 | quantity and type of bins to be based on local waste management policies
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | MMLA-049 | DEVICE: point of care | 2 | quantity and type to suit service requirements
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FQGE-056 | PRINTER: desktop, labels | 1
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | MMLA-055 | FRIDGE: laboratory, underbench | 1 | lockable; temperature monitored
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | may be wall mounted
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | ITNE-003 | BAR CODING: reader | 1 | optional, dependent on software requirements
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FIJO-069 | BENCH TOP: 750D, joinery under, laminate finish | 1 | laboratory grade
BPATH | Bay - Pathology Point of Care Testing (POCT) | FFE | FIJO-099 | CAC: cable access cap | 1
BPATH | Bay - Pathology Point of Care Testing (POCT) | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to fridge
BPATH | Bay - Pathology Point of Care Testing (POCT) | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | quantity to suit project specific equipment requirements
BPATH | Bay - Pathology Point of Care Testing (POCT) | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | to computer
BPATH | Bay - Pathology Point of Care Testing (POCT) | SE | ITBU-001 | ALARM: equipment monitoring, temperature | 1 | to fridge, dependent on requirements, may be integral to fridge
BPATH | Bay - Pathology Point of Care Testing (POCT) | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 5 | quantity to suit project specific equipment requirements, confirm UPS requirements
BPATH | Bay - Pathology Point of Care Testing (POCT) | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to fridge
BPROP | Bay - Property, Staff | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BPROP | Bay - Property, Staff | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3 | coordinated with adjacent area
BPROP | Bay - Property, Staff | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | align height with locker doors
BPROP | Bay - Property, Staff | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 2
BPROP | Bay - Property, Staff | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | coordinated with adjacent area
BPROP | Bay - Property, Staff | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BPROP | Bay - Property, Staff | FFE | FQSN-048 | LOCKER: 2 tier (half size) | 2 | quantity and size of lockers to be based on staffing and service requirements
BPROP | Bay - Property, Staff | FFE | FQSN-050 | LOCKER: 4 tier (quarter size) | 6 | quantity and size of lockers to be based on staffing and service requirements
BPROP | Bay - Property, Staff | FFE | FIJO-212 | BULKHEAD: joinery | 1 | optional; plasterboard bulkhead is also acceptable
BPTS | Bay - Pneumatic Tube Station | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BPTS | Bay - Pneumatic Tube Station | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4 | coordinated with adjacent area
BPTS | Bay - Pneumatic Tube Station | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
BPTS | Bay - Pneumatic Tube Station | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BPTS | Bay - Pneumatic Tube Station | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area
BPTS | Bay - Pneumatic Tube Station | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BPTS | Bay - Pneumatic Tube Station | FFE | FQSN-001 | BASKET: drop in, pneumatic station | 1 | may be lockable depending on location, service and security requirements
BPTS | Bay - Pneumatic Tube Station | FFE | MMLA-072 | RACK: canister storage (pneumatic tube) | 1 | quantity to suit service requirements
BPTS | Bay - Pneumatic Tube Station | FFE | MMGE-097 | STATION: pneumatic tube | 1 | with audible alarm
BPTS | Bay - Pneumatic Tube Station | SE | LIFX-015 | LIGHT: pneumatic tube alert light | 1
BPTS | Bay - Pneumatic Tube Station | SE | ELGE-001 | DIRECT CONNECTION: power | 1 | direct connection to station
BRES | Bay - Resuscitation Trolley | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BRES | Bay - Resuscitation Trolley | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4 | coordinated with adjacent area
BRES | Bay - Resuscitation Trolley | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
BRES | Bay - Resuscitation Trolley | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BRES | Bay - Resuscitation Trolley | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 3 | crash rails are also acceptable
BRES | Bay - Resuscitation Trolley | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area
BRES | Bay - Resuscitation Trolley | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BRES | Bay - Resuscitation Trolley | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1 | optional
BRES | Bay - Resuscitation Trolley | FFE | MMBE-042 | SLIDE: patient | 1 | optional
BRES | Bay - Resuscitation Trolley | FFE | MMHA-027 | TROLLEY: resuscitation, adult | 1 | type dependent on service requirements
BRES | Bay - Resuscitation Trolley | FFE | MMHA-016 | SUCTION: TWIN-O-VAC UNIT | 1
BRES | Bay - Resuscitation Trolley | FFE | MMHA-002 | DEFIBRILLATOR | 1
BRES | Bay - Resuscitation Trolley | FFE | FIHR-090 | BRACKET: for oxygen cylinder, trolley mounted | 1 | optional
BRES | Bay - Resuscitation Trolley | FFE | MMHA-033 | CYLINDER: gas, oxygen | 1 | optional
BRES | Bay - Resuscitation Trolley | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2
BVM | Bay - Vending Machine | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BVM | Bay - Vending Machine | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4 | coordinated with adjacent area
BVM | Bay - Vending Machine | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
BVM | Bay - Vending Machine | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BVM | Bay - Vending Machine | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | coordinated with adjacent area
BVM | Bay - Vending Machine | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BVM | Bay - Vending Machine | FFE | FQCA-055 | MACHINE: vending, cold, drinks | 1 | quantity and type of machine dependent on service requirements
BVM | Bay - Vending Machine | FFE | FQCA-057 | MACHINE: vending, snacks | 1 | quantity and type of machine dependent on service requirements
BVM | Bay - Vending Machine | FFE | FIJO-212 | BULKHEAD: joinery | 1 | optional
BVM | Bay - Vending Machine | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 to each vending machine
BVM | Bay - Vending Machine | SE | ELGP-109 | GPO: single, wall mounted | 2 | 1 to each vending machine
BWC | Bay - Wheelchair Park | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area
BWC | Bay - Wheelchair Park | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4 | coordinated with adjacent area
BWC | Bay - Wheelchair Park | FA | WLPR-006 | WALL PROTECTION: corner guards | 1 | coordinated with adjacent wall protection
BWC | Bay - Wheelchair Park | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
BWC | Bay - Wheelchair Park | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 3 | crash rails are also acceptable
BWC | Bay - Wheelchair Park | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | coordinated with adjacent area
BWC | Bay - Wheelchair Park | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BWC | Bay - Wheelchair Park | FFE | MMBE-045 | WHEELCHAIR | 2 | quantity to suit service requirements
BWC | Bay - Wheelchair Park | FFE | MMBE-050 | WHEELCHAIR: bariatric | 1 | quantity to suit service requirements
BWC | Bay - Wheelchair Park | SE | ELGP-208 | GPO: double, wall mounted | 1 | optional
BWTR | Bay - Water Fountain | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
BWTR | Bay - Water Fountain | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | floor finish of adjacent area must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under water dispenser. significant change in colour to be avoided.
BWTR | Bay - Water Fountain | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BWTR | Bay - Water Fountain | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | coordinated with adjacent area
BWTR | Bay - Water Fountain | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 1
BWTR | Bay - Water Fountain | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
BWTR | Bay - Water Fountain | FFE | FQCW-016 | BIN: waste, general, 20L | 1
BWTR | Bay - Water Fountain | FFE | HYGE-011 | CHILLED WATER UNIT: fixed | 1
BWTR | Bay - Water Fountain | FFE | FQCA-014 | DISPENSER: cup | 1
BWTR | Bay - Water Fountain | SE | ELGP-111 | GPO: single, wall mounted, RCD protected | 1 | optional, for refrigerated units only
CHPT | Change Cubicle - Patient | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CHPT | Change Cubicle - Patient | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
CHPT | Change Cubicle - Patient | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
CHPT | Change Cubicle - Patient | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
CHPT | Change Cubicle - Patient | FA | CLFS-002 | CEILING: fibre cement, flush set, suspended, paint, washable | 1 | drop-in tile ceiling is also acceptable
CHPT | Change Cubicle - Patient | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | outward opening, with privacy latch; cubicle partition system or curtain track with privacy curtain are also acceptable
CHPT | Change Cubicle - Patient | FA | CLCN-009 | CORNICE: wall trim fixing | 1
CHPT | Change Cubicle - Patient | FFE | FIGE-003 | BENCH SEAT: fixed, indoor | 1 | colour of seating to contrast with flooring and walls as per AS1428 luminance contrast requirements
CHPT | Change Cubicle - Patient | FFE | FIHR-026 | HOOK: coat | 2
CHPT | Change Cubicle - Patient | FFE | FIRT-006 | GRAB RAIL: WC | 1 | L-shaped grabrail
CHPT | Change Cubicle - Patient | FFE | FIBM-005 | MIRROR: fixed, glass, 1200H nom. | 1
CHPT | Change Cubicle - Patient | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | with indicator button and light
CHPT | Change Cubicle - Patient | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | with indicator button and light, low for access from seated position or floor
CHPT | Change Cubicle - Patient | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
CHPT | Change Cubicle - Patient | SE | ITSE-044 | SENSOR: motion detection, for lighting control | 1 | light switch is also acceptable
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | Default= vinyl, carpet also acceptable
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | Feather edged vinyl if carpeted
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | To be decided at project level.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | WLWA-007 | WALL FINISH: splashback | 1 | To basin. Selection of material to comply with Part D Infection Prevention and Control AusHFG.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | CLCN-002 | CORNICE: aluminium, 24mm shadow angle, powdercoat | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | optional, 910mm outward opening, privacy latch, with emergency release
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIGE-003 | BENCH SEAT: fixed, indoor | 1 | To change cubicle. Or chair, consider weight capacity.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FQBS-030 | CHAIR: visitor, patient | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | WLPT-001 | PARTITION: change cubicle | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1 | Optional
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIHR-026 | HOOK: coat | 2
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIDI-030 | DISPENSER: soap | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIDI-036 | DISPENSER: toilet paper | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | MMGE-122 | TROLLEY: gowns | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIRT-006 | GRAB RAIL: WC | 2 | Ambulant provision.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FQSN-048 | LOCKER: 2 tier (half size) | 4 | 16 half lockers or quarter lockers.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FQSN-050 | LOCKER: 4 tier (quarter size) | 2
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIBM-006 | MIRROR: fixed, glass | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FQCW-037 | BIN: waste, sanitary | 1 | To Female Change Rooms. Optional, depends on Waste Management Policy.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | FIJO-191 | SHELF: above basin | 1 | Or integrated with basin.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | WLPT-005 | PARTITION: toilet | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | HYBA-011 | BASIN: handwash, Type C | 1 | With vanity, or ensuite basin with integral shelf for patient belongings.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1 | to basin
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | FFE | HYWC-006 | TOILET SUITE: patient | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | SE | HYDR-004 | DRAIN: floor waste | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | SE | ELBO-015 | SWITCH: light, single | 1
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 2 | + indicator button & light. WC and change cubicle.
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | SE | ELGP-208 | GPO: double, wall mounted | 1 | Optional.
CHPT-AC | Change Cubicle - Patient, Accessible | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CHPT-AC | Change Cubicle - Patient, Accessible | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
CHPT-AC | Change Cubicle - Patient, Accessible | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
CHPT-AC | Change Cubicle - Patient, Accessible | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
CHPT-AC | Change Cubicle - Patient, Accessible | FA | CLFS-002 | CEILING: fibre cement, flush set, suspended, paint, washable | 1 | drop-in tile ceiling is also acceptable
CHPT-AC | Change Cubicle - Patient, Accessible | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | outward opening, with accessible privacy latch; cubicle system or curtain track with privacy curtain are also acceptable
CHPT-AC | Change Cubicle - Patient, Accessible | FA | CLCN-009 | CORNICE: wall trim fixing | 1
CHPT-AC | Change Cubicle - Patient, Accessible | FFE | FIGE-003 | BENCH SEAT: fixed, indoor | 1 | colour of seating to contrast with flooring and walls as per AS1428.1 luminance contrast requirements
CHPT-AC | Change Cubicle - Patient, Accessible | FFE | FIHR-026 | HOOK: coat | 2 | 1 mounted at accessible height to AS1428.1 requirements
CHPT-AC | Change Cubicle - Patient, Accessible | FFE | FIRT-006 | GRAB RAIL: WC | 1 | L-shaped grabrail
CHPT-AC | Change Cubicle - Patient, Accessible | FFE | FIBM-006 | MIRROR: fixed, glass | 1
CHPT-AC | Change Cubicle - Patient, Accessible | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | with indicator button and light
CHPT-AC | Change Cubicle - Patient, Accessible | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | with indicator button and light, low for access from seated position or floor
CHPT-AC | Change Cubicle - Patient, Accessible | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
CHPT-AC | Change Cubicle - Patient, Accessible | SE | ITSE-044 | SENSOR: motion detection, for lighting control | 1 | light switch is also acceptable
CHST-10 | Change - Staff, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 12 | to be tiles if tiled floor
CHST-10 | Change - Staff, 10m2 | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
CHST-10 | Change - Staff, 10m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
CHST-10 | Change - Staff, 10m2 | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
CHST-10 | Change - Staff, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 6
CHST-10 | Change - Staff, 10m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1 | water resistant ceiling tiles are also acceptable
CHST-10 | Change - Staff, 10m2 | FA | DWSC-001 | DOOR: 850mm c/o, single leaf, solid core, paint | 1 | with closer as required for privacy
CHST-10 | Change - Staff, 10m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CHST-10 | Change - Staff, 10m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 4 | to shower; tiles are also acceptable
CHST-10 | Change - Staff, 10m2 | FFE | FIGE-003 | BENCH SEAT: fixed, indoor | 2 | 1 x to shower, material to be suitable for wet areas
CHST-10 | Change - Staff, 10m2 | FFE | FIHR-026 | HOOK: coat | 4 | 2 x to changing area optional; quantity to be based on staffing and service requirements
CHST-10 | Change - Staff, 10m2 | FFE | FIRT-033 | TRACK: curtain, shower | 1 | optional
CHST-10 | Change - Staff, 10m2 | FFE | FQGE-011 | CURTAIN: shower | 1 | optional
CHST-10 | Change - Staff, 10m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CHST-10 | Change - Staff, 10m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CHST-10 | Change - Staff, 10m2 | FFE | FIDI-030 | DISPENSER: soap | 2 | 1 x to basin, 1 x to shower
CHST-10 | Change - Staff, 10m2 | FFE | FIDI-036 | DISPENSER: toilet paper | 1
CHST-10 | Change - Staff, 10m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
CHST-10 | Change - Staff, 10m2 | FFE | FIBM-006 | MIRROR: fixed, glass | 1 | to basin
CHST-10 | Change - Staff, 10m2 | FFE | FQCW-037 | BIN: waste, sanitary | 1 | optional, to toilet in female change rooms
CHST-10 | Change - Staff, 10m2 | FFE | WLPT-004 | PARTITION: shower | 1
CHST-10 | Change - Staff, 10m2 | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 1 | optional, to shower
CHST-10 | Change - Staff, 10m2 | FFE | WLPT-005 | PARTITION: toilet | 1 | lift off hinges to inward opening sanitary compartment doors
CHST-10 | Change - Staff, 10m2 | FFE | HYBA-012 | BASIN: handwash, Type C, with vanity unit | 1
CHST-10 | Change - Staff, 10m2 | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1
CHST-10 | Change - Staff, 10m2 | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 1
CHST-10 | Change - Staff, 10m2 | FFE | HYTP-018 | SHOWER: fixed rose | 1
CHST-10 | Change - Staff, 10m2 | FFE | HYWC-008 | TOILET SUITE: standard | 1 | with 100mm clearout access point
CHST-10 | Change - Staff, 10m2 | FFE | FIJO-212 | BULKHEAD: joinery | 1 | optional, over lockers, plasterboard bulkhead is also acceptable
CHST-10 | Change - Staff, 10m2 | FFE | FQSN-054 | LOCKER: 2 tier with shoe storage | 2
CHST-10 | Change - Staff, 10m2 | FFE | FQSN-055 | LOCKER: 4 tier with shoe storage | 2
CHST-10 | Change - Staff, 10m2 | SE | HYDR-004 | DRAIN: floor waste | 1
CHST-10 | Change - Staff, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
CHST-10 | Change - Staff, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
CHST-10 | Change - Staff, 10m2 | SE | ITSE-038 | READER: access control, general | 1 | to entry door from corridor
CHST-35 | Change - Staff, 35m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 33 | to be tiles if tiled floor
CHST-35 | Change - Staff, 35m2 | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
CHST-35 | Change - Staff, 35m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 2
CHST-35 | Change - Staff, 35m2 | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basins
CHST-35 | Change - Staff, 35m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 16
CHST-35 | Change - Staff, 35m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1 | water resistant ceiling tiles are also acceptable
CHST-35 | Change - Staff, 35m2 | FA | DWSC-001 | DOOR: 850mm c/o, single leaf, solid core, paint | 2 | with closer as required for privacy
CHST-35 | Change - Staff, 35m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CHST-35 | Change - Staff, 35m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 8 | to shower; tiles are also acceptable
CHST-35 | Change - Staff, 35m2 | FFE | ITNE-052 | TELEPHONE: handset, wall mount | 1 | optional
CHST-35 | Change - Staff, 35m2 | FFE | FIGE-003 | BENCH SEAT: fixed, indoor | 2 | to showers, material to be suitable for wet areas
CHST-35 | Change - Staff, 35m2 | FFE | FQBS-007 | BENCH SEAT: freestanding, indoor | 1
CHST-35 | Change - Staff, 35m2 | FFE | FIHR-026 | HOOK: coat | 16 | optional, quantity and configuration to be determined to suit operational requirements
CHST-35 | Change - Staff, 35m2 | FFE | FIRT-033 | TRACK: curtain, shower | 2 | optional
CHST-35 | Change - Staff, 35m2 | FFE | FQGE-011 | CURTAIN: shower | 2 | optional
CHST-35 | Change - Staff, 35m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 2 | to basins
CHST-35 | Change - Staff, 35m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basins
CHST-35 | Change - Staff, 35m2 | FFE | FIDI-030 | DISPENSER: soap | 4 | 2 x to basins, 2 x to showers
CHST-35 | Change - Staff, 35m2 | FFE | FIDI-036 | DISPENSER: toilet paper | 3
CHST-35 | Change - Staff, 35m2 | FFE | FIRT-006 | GRAB RAIL: WC | 2 | to ambulant toilet
CHST-35 | Change - Staff, 35m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 2
CHST-35 | Change - Staff, 35m2 | FFE | MMGE-109 | TROLLEY: clean linen | 1
CHST-35 | Change - Staff, 35m2 | FFE | FIBM-006 | MIRROR: fixed, glass | 1 | to basins
CHST-35 | Change - Staff, 35m2 | FFE | FIBM-007 | MIRROR: fixed, glass, full length | 1
CHST-35 | Change - Staff, 35m2 | FFE | FQCW-037 | BIN: waste, sanitary | 3 | optional, to toilets in female change rooms
CHST-35 | Change - Staff, 35m2 | FFE | WLPT-004 | PARTITION: shower | 2
CHST-35 | Change - Staff, 35m2 | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 2 | optional, to showers
CHST-35 | Change - Staff, 35m2 | FFE | WLPT-005 | PARTITION: toilet | 2 | lift off hinges to inward opening sanitary compartment doors
CHST-35 | Change - Staff, 35m2 | FFE | HYBA-012 | BASIN: handwash, Type C, with vanity unit | 2
CHST-35 | Change - Staff, 35m2 | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 2
CHST-35 | Change - Staff, 35m2 | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 2
CHST-35 | Change - Staff, 35m2 | FFE | HYTP-018 | SHOWER: fixed rose | 2
CHST-35 | Change - Staff, 35m2 | FFE | HYWC-008 | TOILET SUITE: standard | 2 | with 100mm clearout access point
CHST-35 | Change - Staff, 35m2 | FFE | HYWC-010 | TOILET SUITE: ambulant | 1 | with 100mm clearout access point
CHST-35 | Change - Staff, 35m2 | FFE | FIJO-212 | BULKHEAD: joinery | 3 | optional, over lockers, plasterboard bulkhead is also acceptable
CHST-35 | Change - Staff, 35m2 | FFE | WLPT-016 | PARTITION: toilet, ambulant | 1
CHST-35 | Change - Staff, 35m2 | FFE | FQSN-054 | LOCKER: 2 tier with shoe storage | 6
CHST-35 | Change - Staff, 35m2 | FFE | FQSN-055 | LOCKER: 4 tier with shoe storage | 17
CHST-35 | Change - Staff, 35m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional, to wall phone if provided
CHST-35 | Change - Staff, 35m2 | SE | HYDR-004 | DRAIN: floor waste | 3
CHST-35 | Change - Staff, 35m2 | SE | ELBO-015 | SWITCH: light, single | 2
CHST-35 | Change - Staff, 35m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
CHST-35 | Change - Staff, 35m2 | SE | ITSE-038 | READER: access control, general | 2 | to entry door from corridor
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 4
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | WLWA-019 | WALL FINISH: vinyl, full height | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | at 3000 mm AFFL
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to control room
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | DWGL-006 | DOOR: 910mm c/o, single leaf, observation panel, solid core/lead glass, lead shielding, paint | 2 | staff access to sterile stock/ scrub areas
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MGFP-002 | FLOWMETER: medical air | 3 | on pendant/s
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIHR-019 | BRACKET: for suction bottle | 4 | on pendant/s
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | wall mounted
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-119 | TROLLEY: equipment | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1 | for slave monitor
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-122 | TROLLEY: gowns | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-045 | IV STAND: mobile | 2
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MGFP-004 | FLOWMETER: oxygen | 4 | on pendant/s
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 1 | under bench
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMHA-027 | TROLLEY: resuscitation, adult | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-088 | SAFE: dangerous drugs, small | 1 | optional
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-142 | TROLLEY: sharps bin | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FQBS-052 | STOOL: mobile, high, adj height | 2
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-002 | ADAPTER: suction | 4 | on pendant/s
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-100 | SUCTION BOTTLE | 4 | on pendant mounted brackets
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIJO-108 | CUPBOARD: drawer unit, baskets, 3 x Wire (400mm wide) | 3
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMOP-040 | MACHINE: anaesthetic with integral monitors | 1 | optional
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMSP-015 | PENDANT: single, anaesthetic (fixed or adjustable height) | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | ITCL-021 | CLOCK: clock system slave dual time / elapsed time | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 2
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMHA-002 | DEFIBRILLATOR | 1 | on resus trolley
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMSP-029 | LIGHT: operating, single, 550 diam nom. | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-032 | DISPLAY SCREEN: slave, for patient monitor | 1 | on wall bracket
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMSP-018 | PENDANT: single, wet and dry combination (fixed or adjustable height) | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-329 | DEVICE: imaging, ultrasound scanning unit | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 2
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-100 | INJECTOR: contrast, CT | 1 | optionally can be on table or ceiling mounted
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIJO-212 | BULKHEAD: joinery | 1 | over cupboards
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-377 | WARMER: contrast | 1 | on bench
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-385 | SYSTEM: imaging, cath lab, dual plane, ceiling | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-389 | TABLE: operating, fixed, narrow table plate, cath lab, dual plane | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-390 | SHIELD: radiation, upper body, mobile, cath lab | 2 | ceiling mounted
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-396 | DISPLAY SCREEN: cath lab, multi modality | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIJO-220 | CUPBOARD: tall, double door, glazed, catheter hanging | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-398 | SYSTEM: electrophysiology, cath lab | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | MMMI-400 | STAND: ceiling, with traverse bridge, cath lab | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | FFE | FIJO-221 | CUPBOARD: tall, double door, glazed | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITIN-022 | OUTLET: data, double RJ45, in floorbox | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room, outside room, to control room, delete from room list if listed in adjacent accessway
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELPR-006 | RCD: residual current device | 4 | 1 to each pendant, 1 for wall mounted outlets and 1 for floor box outlets
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 1 | under table
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELBO-003 | CONTROL PANEL: lighting | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITCL-008 | BUTTON: nurse call, staff assist | 2 | to wall, + indicator button & light
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELBO-001 | BUTTON: emergency stop | 1 | optional, for equipment shut-off
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | LIFX-008 | LIGHT: task light, built in | 2 | over bench
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 3 | outside all entry points. Delete from room list if listed in adjacent room
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | MGFP-001 | ALARM PANEL: medical gases | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITCL-004 | BUTTON: nurse call, emergency | 2 | to wall, + indicator button & light
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | 2 above bench
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 9
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELPR-004 | MONITOR: line isolation | 2 | 1 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITCL-023 | CONTROL PANEL: master clock system | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELBO-005 | EQUIPOTENTIAL JUNCTION BOX | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELBO-012 | SWITCH: light, dimmer | 4
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ITIN-024 | OUTLET: data, double RJ45, on pendant | 2 | on pendant, 1 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | MGAS-021 | OUTLET: medical air (MA), on pendant | 3 | on pendant, 2 to EP pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | MGAS-031 | OUTLET: nitrous oxide (N2O), on pendant | 2 | on pendant, 1 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | MGAS-041 | OUTLET: oxygen (O2), on pendant | 4 | on pendant, 2 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | MGAS-061 | OUTLET: suction, on pendant | 4 | on pendant, 2 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | MGAS-051 | OUTLET: scavenge, on pendant | 2 | on pendant, 1 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELGP-132 | GPO: UPS, single, on pendant | 4 | on pendant, 2 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELGP-121 | GPO: emergency power, single, on pendant | 16 | on pendant, 8 to each pendant
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELGP-225 | GPO: emergency power, double, in floorbox | 3
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | MEEX-016 | RETURN AIR: grille, low level | 4
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | SE | ELPR-007 | UPS: status indicator | 2
CLAB-I | Catheter Laboratory - Interventional | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CLAB-I | Catheter Laboratory - Interventional | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 18
CLAB-I | Catheter Laboratory - Interventional | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
CLAB-I | Catheter Laboratory - Interventional | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 4
CLAB-I | Catheter Laboratory - Interventional | FA | WLWA-005 | WALL FINISH: paint, scrub resistant | 5
CLAB-I | Catheter Laboratory - Interventional | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | at 3000 mm AFFL
CLAB-I | Catheter Laboratory - Interventional | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to control room
CLAB-I | Catheter Laboratory - Interventional | FA | DWGL-006 | DOOR: 910mm c/o, single leaf, observation panel, solid core/lead glass, lead shielding, paint | 2 | staff access to sterile stock/ scrub areas, with sliding cover for privacy
CLAB-I | Catheter Laboratory - Interventional | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
CLAB-I | Catheter Laboratory - Interventional | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLAB-I | Catheter Laboratory - Interventional | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 20
CLAB-I | Catheter Laboratory - Interventional | FFE | MGFP-002 | FLOWMETER: medical air | 1 | on pendant/s
CLAB-I | Catheter Laboratory - Interventional | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | on pendant/s
CLAB-I | Catheter Laboratory - Interventional | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2
CLAB-I | Catheter Laboratory - Interventional | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | wall mounted
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-119 | TROLLEY: equipment | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1 | for slave monitor
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-122 | TROLLEY: gowns | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-045 | IV STAND: mobile | 2
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | on pendant/s
CLAB-I | Catheter Laboratory - Interventional | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 1 | under bench
CLAB-I | Catheter Laboratory - Interventional | FFE | MMHA-027 | TROLLEY: resuscitation, adult | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-088 | SAFE: dangerous drugs, small | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-142 | TROLLEY: sharps bin | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | FQBS-052 | STOOL: mobile, high, adj height | 2
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-002 | ADAPTER: suction | 2 | on pendant/s
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-100 | SUCTION BOTTLE | 2 | on pendant mounted brackets
CLAB-I | Catheter Laboratory - Interventional | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | FIJO-108 | CUPBOARD: drawer unit, baskets, 3 x Wire (400mm wide) | 3
CLAB-I | Catheter Laboratory - Interventional | FFE | MMSP-015 | PENDANT: single, anaesthetic (fixed or adjustable height) | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 2
CLAB-I | Catheter Laboratory - Interventional | FFE | MMHA-002 | DEFIBRILLATOR | 1 | on resus trolley
CLAB-I | Catheter Laboratory - Interventional | FFE | MMSP-029 | LIGHT: operating, single, 550 diam nom. | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-032 | DISPLAY SCREEN: slave, for patient monitor | 1 | on wall bracket
CLAB-I | Catheter Laboratory - Interventional | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-329 | DEVICE: imaging, ultrasound scanning unit | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 2
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-108 | INJECTOR: contrast, table mounted | 1 | optionally can be mobile or ceiling mounted
CLAB-I | Catheter Laboratory - Interventional | FFE | FIJO-212 | BULKHEAD: joinery | 1 | over cupboards
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-377 | WARMER: contrast | 1 | on bench
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-387 | SYSTEM: imaging, cath lab, single plane, ceiling | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-388 | TABLE: operating, fixed, narrow, cath lab, single plane | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-390 | SHIELD: radiation, upper body, mobile, cath lab | 1 | ceiling mounted
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-396 | DISPLAY SCREEN: cath lab, multi modality | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | FIJO-220 | CUPBOARD: tall, double door, glazed, catheter hanging | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | MMMI-400 | STAND: ceiling, with traverse bridge, cath lab | 1
CLAB-I | Catheter Laboratory - Interventional | FFE | FIJO-221 | CUPBOARD: tall, double door, glazed | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
CLAB-I | Catheter Laboratory - Interventional | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3
CLAB-I | Catheter Laboratory - Interventional | SE | ITIN-022 | OUTLET: data, double RJ45, in floorbox | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room
CLAB-I | Catheter Laboratory - Interventional | SE | ELPR-006 | RCD: residual current device | 3 | 1 to pendant, 1 for wall mounted outlets, 1 for floor box outlets
CLAB-I | Catheter Laboratory - Interventional | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 1 | under table
CLAB-I | Catheter Laboratory - Interventional | SE | ELBO-003 | CONTROL PANEL: lighting | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ITCL-008 | BUTTON: nurse call, staff assist | 2 | to wall, + indicator button & light
CLAB-I | Catheter Laboratory - Interventional | SE | ELBO-001 | BUTTON: emergency stop | 1 | optional, for equipment shut-off
CLAB-I | Catheter Laboratory - Interventional | SE | LIFX-008 | LIGHT: task light, built in | 2 | over bench
CLAB-I | Catheter Laboratory - Interventional | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 3 | outside all entry points, delete from room list if listed in adjacent room
CLAB-I | Catheter Laboratory - Interventional | SE | MGFP-001 | ALARM PANEL: medical gases | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ITCL-004 | BUTTON: nurse call, emergency | 2 | to wall, + indicator button & light
CLAB-I | Catheter Laboratory - Interventional | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 3 | 2 above bench
CLAB-I | Catheter Laboratory - Interventional | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 9
CLAB-I | Catheter Laboratory - Interventional | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ELPR-004 | MONITOR: line isolation | 1 | 1 to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
CLAB-I | Catheter Laboratory - Interventional | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified
CLAB-I | Catheter Laboratory - Interventional | SE | ELBO-005 | EQUIPOTENTIAL JUNCTION BOX | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ELBO-012 | SWITCH: light, dimmer | 4
CLAB-I | Catheter Laboratory - Interventional | SE | ITIN-024 | OUTLET: data, double RJ45, on pendant | 1 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | MGAS-021 | OUTLET: medical air (MA), on pendant | 1 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | MGAS-031 | OUTLET: nitrous oxide (N2O), on pendant | 1 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | MGAS-041 | OUTLET: oxygen (O2), on pendant | 2 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | MGAS-061 | OUTLET: suction, on pendant | 2 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | MGAS-051 | OUTLET: scavenge, on pendant | 1 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | ELGP-132 | GPO: UPS, single, on pendant | 2 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | ELGP-121 | GPO: emergency power, single, on pendant | 8 | to pendant
CLAB-I | Catheter Laboratory - Interventional | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1
CLAB-I | Catheter Laboratory - Interventional | SE | ELGP-225 | GPO: emergency power, double, in floorbox | 3
CLAB-I | Catheter Laboratory - Interventional | SE | MEEX-016 | RETURN AIR: grille, low level | 4
CLAB-I | Catheter Laboratory - Interventional | SE | ELPR-007 | UPS: status indicator | 1
CLCR-EP | Catheter Laboratory EP Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CLCR-EP | Catheter Laboratory EP Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
CLCR-EP | Catheter Laboratory EP Control Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
CLCR-EP | Catheter Laboratory EP Control Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
CLCR-EP | Catheter Laboratory EP Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CLCR-EP | Catheter Laboratory EP Control Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
CLCR-EP | Catheter Laboratory EP Control Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to procedure room
CLCR-EP | Catheter Laboratory EP Control Room | FA | DWGL-002 | DOOR: 910mm c/o, single leaf, half glazed, solid core/glass, paint | 1
CLCR-EP | Catheter Laboratory EP Control Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 4
CLCR-EP | Catheter Laboratory EP Control Room | FFE | DWWD-004 | BLIND: roller | 1 | to observation window for patient privacy; refer to local infection control policies
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FQSN-031 | CABINET: short, single door, tambour, adj shelves, lockable | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 2
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITBU-013 | DISPLAY SCREEN: patient flow system | 1 | patient journey board
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITNE-023 | PRINTER: networked | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1 | optional
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2 | on bench, 1 cordless optional
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 2 | part of communications system integrated with imaging system
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 2
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FIJO-099 | CAC: cable access cap | 6
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | for reporting station
CLCR-EP | Catheter Laboratory EP Control Room | FFE | MMMI-369 | DISPLAY SCREEN: contrast injector | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 2 | for reporting station
CLCR-EP | Catheter Laboratory EP Control Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | MMMI-386 | CONTROL UNIT: c/w display, control box and keyboard, Cath Lab | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | MMMI-392 | DISPLAY SCREEN: DCS, cath lab | 1 | haemodynamic monitoring display
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | ITNE-069 | DISPLAY SCREEN: PACS, portrait | 2
CLCR-EP | Catheter Laboratory EP Control Room | FFE | MMMI-397 | COMPUTER: control, c/w dual display & keyboard, 3D mapping | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | MMMI-399 | SYSTEM: electrophysiology recording, c/w display screens | 1
CLCR-EP | Catheter Laboratory EP Control Room | FFE | MMGE-182 | MACHINE: Electrophysiology, cardiac pacing | 1
CLCR-EP | Catheter Laboratory EP Control Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 4
CLCR-EP | Catheter Laboratory EP Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 6 | 2 outlets for telephone
CLCR-EP | Catheter Laboratory EP Control Room | SE | ELPR-006 | RCD: residual current device | 4
CLCR-EP | Catheter Laboratory EP Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1 | optional
CLCR-EP | Catheter Laboratory EP Control Room | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | for control room temperature, optional
CLCR-EP | Catheter Laboratory EP Control Room | SE | MGFP-001 | ALARM PANEL: medical gases | 1
CLCR-EP | Catheter Laboratory EP Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 7
CLCR-EP | Catheter Laboratory EP Control Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
CLCR-EP | Catheter Laboratory EP Control Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 3
CLCR-EP | Catheter Laboratory EP Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
CLCR-EP | Catheter Laboratory EP Control Room | SE | ELGP-130 | GPO: UPS, single | 15
CLCR-I | Catheter Laboratory Interventional Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
CLCR-I | Catheter Laboratory Interventional Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CLCR-I | Catheter Laboratory Interventional Control Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to procedure room
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWGL-002 | DOOR: 910mm c/o, single leaf, half glazed, solid core/glass, paint | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWGL-006 | DOOR: 910mm c/o, single leaf, observation panel, solid core/lead glass, lead shielding, paint | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 3
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | DWWD-004 | BLIND: roller | 1 | to observation window for patient privacy; refer to local infection control policies
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITBU-013 | DISPLAY SCREEN: patient flow system | 1 | patient journey board
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-009 | COMPUTER: hi performance, for clinical / PACS use | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-023 | PRINTER: networked | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1 | optional
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2 | on bench, 1 cordless optional
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 2 | part of communications system integrated with imaging system
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 2
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-099 | CAC: cable access cap | 5
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | for reporting station
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | MMMI-369 | DISPLAY SCREEN: contrast injector | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 3
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | MMMI-386 | CONTROL UNIT: c/w display, control box and keyboard, Cath Lab | 2
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 3
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-069 | DISPLAY SCREEN: PACS, portrait | 2
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 6 | 2 outlets for telephone
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELPR-006 | RCD: residual current device | 3
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1 | optional
CLCR-I | Catheter Laboratory Interventional Control Room | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | for control room temperature, optional
CLCR-I | Catheter Laboratory Interventional Control Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | MGFP-001 | ALARM PANEL: medical gases | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 8
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-130 | GPO: UPS, single | 7
CLN-10 | Clean Store, 10m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CLN-10 | Clean Store, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
CLN-10 | Clean Store, 10m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
CLN-10 | Clean Store, 10m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
CLN-10 | Clean Store, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CLN-10 | Clean Store, 10m2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
CLN-10 | Clean Store, 10m2 | FA | DWGL-003 | DOOR: 910mm c/o, single leaf, half glazed, solid core/safety glass, paint | 1
CLN-10 | Clean Store, 10m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLN-10 | Clean Store, 10m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 5 | extent dependent on type and height storage solution
CLN-10 | Clean Store, 10m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CLN-10 | Clean Store, 10m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
CLN-10 | Clean Store, 10m2 | FFE | FIJO-108 | CUPBOARD: drawer unit, baskets, 3 x Wire (400mm wide) | 2 | remove if desk provided in lieu of bench
CLN-10 | Clean Store, 10m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1 | desk is also acceptable
CLN-10 | Clean Store, 10m2 | FFE | FQSM-046 | SHELVING: storage solution, 600W x 400D x 2100H nom., wire, mobile | 5 | type and quantity of storage options to be determined based on service requirements
CLN-10 | Clean Store, 10m2 | FFE | FQSM-048 | SHELVING: storage solution, shelving, 600W x 400D x 2100H nom., wire, mobile | 2 | type and quantity of storage options to be determined based on service requirements
CLN-10 | Clean Store, 10m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
CLN-10 | Clean Store, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
CLN-10 | Clean Store, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 3 | extent dependent on type and height storage solution
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | DWGL-003 | DOOR: 910mm c/o, single leaf, half glazed, solid core/safety glass, paint | 2
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 3 | extent dependent on type and height storage solution
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | MMGE-087 | SAFE: dangerous drugs, medium | 1 | size and quantity to be based on service requirements
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIJO-108 | CUPBOARD: drawer unit, baskets, 3 x Wire (400mm wide) | 8
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | ITSE-012 | CAMERA: CCTV | 1 | optional, provision and placement dependent on local policies
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | MMGE-034 | FRIDGE: drugs, large | 1 | size and quantity to be based on service requirements
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FQSM-046 | SHELVING: storage solution, 600W x 400D x 2100H nom., wire, mobile | 5 | type and quantity of storage options to be determined based on service requirements
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FQCW-061 | BIN: waste, pharmaceutical | 1
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FQSM-048 | SHELVING: storage solution, shelving, 600W x 400D x 2100H nom., wire, mobile | 4 | type and quantity of storage options to be determined based on service requirements
CLN-MED-20 | Clean Store / Medication Room, 20m2 | FFE | FIDI-052 | DISPENSER: sticker roll | 1 | optional
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to drug fridge, for temperature monitoring
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | 2 x optional, for future provision of automated dispensing cabinets
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | LIFX-008 | LIGHT: task light, built in | 2 | under wall mounted joinery
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ELBO-015 | SWITCH: light, single | 1 | to task light
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ELBO-010 | SWITCH: light, 2-way | 2
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | optional, for future provision of automated dispensing cabinets
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to drug fridge
CLN-MED-20 | Clean Store / Medication Room, 20m2 | SE | ITSE-038 | READER: access control, general | 2
CLN-MED-S | Clean Store / Medication Room, Sub | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
CLN-MED-S | Clean Store / Medication Room, Sub | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 11
CLN-MED-S | Clean Store / Medication Room, Sub | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
CLN-MED-S | Clean Store / Medication Room, Sub | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CLN-MED-S | Clean Store / Medication Room, Sub | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 4 | extent dependent on type and height storage solution
CLN-MED-S | Clean Store / Medication Room, Sub | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FA | DWGL-003 | DOOR: 910mm c/o, single leaf, half glazed, solid core/safety glass, paint | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 1 | extent dependent on type and height storage solution
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIHR-016 | BRACKET: for sharps bin | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIJO-140 | CUPBOARD: underbench, single door, adj shelves | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIJO-157 | CUPBOARD: wall mount, single door, adj shelves, lockable | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | MMGE-087 | SAFE: dangerous drugs, medium | 1 | size and quantity to be based on service requirements
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIJO-108 | CUPBOARD: drawer unit, baskets, 3 x Wire (400mm wide) | 2
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | HYBA-010 | BASIN: handwash, Type B | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | ITSE-012 | CAMERA: CCTV | 1 | optional, provision and placement dependent on local policies
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | MMGE-036 | FRIDGE: drugs, underbench | 1 | size and quantity to be based on service requirements
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FQSM-046 | SHELVING: storage solution, 600W x 400D x 2100H nom., wire, mobile | 3 | type and quantity of storage options to be determined based on service requirements
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FQCW-061 | BIN: waste, pharmaceutical | 1
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FQSM-048 | SHELVING: storage solution, shelving, 600W x 400D x 2100H nom., wire, mobile | 2 | type and quantity of storage options to be determined based on service requirements
CLN-MED-S | Clean Store / Medication Room, Sub | FFE | FIDI-052 | DISPENSER: sticker roll | 1 | optional
CLN-MED-S | Clean Store / Medication Room, Sub | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to drug fridge, for temperature monitoring
CLN-MED-S | Clean Store / Medication Room, Sub | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | 1x optional, for future provision of automated dispensing cabinets
CLN-MED-S | Clean Store / Medication Room, Sub | SE | LIFX-008 | LIGHT: task light, built in | 1 | under wall mounted joinery
CLN-MED-S | Clean Store / Medication Room, Sub | SE | ELBO-015 | SWITCH: light, single | 2 | 1 x general room lighting, 1 x task light
CLN-MED-S | Clean Store / Medication Room, Sub | SE | ELGP-208 | GPO: double, wall mounted | 2
CLN-MED-S | Clean Store / Medication Room, Sub | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | optional, for future provision of automated dispensing cabinets
CLN-MED-S | Clean Store / Medication Room, Sub | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to drug fridge
CLN-MED-S | Clean Store / Medication Room, Sub | SE | ITSE-038 | READER: access control, general | 1
CLRM | Cleaner's Room | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1
CLRM | Cleaner's Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
CLRM | Cleaner's Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
CLRM | Cleaner's Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
CLRM | Cleaner's Room | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
CLRM | Cleaner's Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | lockable or secure staff-only access may be provided, hold open hardware may be required, optional vision panel
CLRM | Cleaner's Room | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 1
CLRM | Cleaner's Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
CLRM | Cleaner's Room | FFE | FQCW-045 | BUCKET: cleaner | 1 | on cleaner's trolley
CLRM | Cleaner's Room | FFE | FIHR-040 | RACK: broom/mop | 1
CLRM | Cleaner's Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CLRM | Cleaner's Room | FFE | FQSM-021 | TROLLEY: cleaner | 1
CLRM | Cleaner's Room | FFE | FIHR-026 | HOOK: coat | 2
CLRM | Cleaner's Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CLRM | Cleaner's Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
CLRM | Cleaner's Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CLRM | Cleaner's Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CLRM | Cleaner's Room | FFE | FIDI-021 | DISPENSER: mask | 1 | optional, type of PPE dispensers to be provided to suit service requirements
CLRM | Cleaner's Room | FFE | FQCW-058 | MOP | 1
CLRM | Cleaner's Room | FFE | FIJO-196 | SHELF: stainless steel, overbench, wall mount | 1 | type and extent of storage to suit service requirements and operational policies for the decanting and storage of chemicals and consumables
CLRM | Cleaner's Room | FFE | HYBA-011 | BASIN: handwash, Type C | 1
CLRM | Cleaner's Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to cleaner's sink
CLRM | Cleaner's Room | FFE | FIDI-010 | DISPENSER: detergent/chemical mixer | 1
CLRM | Cleaner's Room | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1 | to basin
CLRM | Cleaner's Room | FFE | HYBA-040 | SINK: cleaner | 1
CLRM | Cleaner's Room | FFE | FIDI-057 | DISPENSER: eye wash bottles | 1 | optional, provide as required depending chemical safety data sheets and WHS policies
CLRM | Cleaner's Room | SE | HYTP-008 | OUTLET: cold water | 1 | optional, provision dependent on detergent/chemical mixer requirements
CLRM | Cleaner's Room | SE | HYDR-004 | DRAIN: floor waste | 1
CLRM | Cleaner's Room | SE | ELBO-015 | SWITCH: light, single | 1
CLRM | Cleaner's Room | SE | ELGP-208 | GPO: double, wall mounted | 1 | quantity to suit equipment charging requirements
CLRM | Cleaner's Room | SE | ITSE-038 | READER: access control, general | 1 | optional, dependent on security requirements
CLUP-7 | Clean-Up Room, 7m2 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1
CLUP-7 | Clean-Up Room, 7m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
CLUP-7 | Clean-Up Room, 7m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
CLUP-7 | Clean-Up Room, 7m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
CLUP-7 | Clean-Up Room, 7m2 | FA | WLWA-005 | WALL FINISH: paint, scrub resistant | 4
CLUP-7 | Clean-Up Room, 7m2 | FA | WLWA-021 | WALL FINISH: vinyl, to 1800 AFFL | 1
CLUP-7 | Clean-Up Room, 7m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
CLUP-7 | Clean-Up Room, 7m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
CLUP-7 | Clean-Up Room, 7m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | with wet edge
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIJO-196 | SHELF: stainless steel, overbench, wall mount | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FQSM-016 | TROLLEY | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FQCW-029 | BIN: waste, recycling, paper, 240L | 1 | optional
CLUP-7 | Clean-Up Room, 7m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | elbow action levers
CLUP-7 | Clean-Up Room, 7m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | 1 to sink, 1 to basin
CLUP-7 | Clean-Up Room, 7m2 | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | HYBA-052 | SINK: single, 450 x 350 x 250D | 1 | hot water
CLUP-7 | Clean-Up Room, 7m2 | FFE | FIJO-100 | BENCH LEGS: stainless steel, for bench top | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1
CLUP-7 | Clean-Up Room, 7m2 | FFE | FQCW-060 | BIN: waste, cytotoxic, 50L | 1 | optional
CLUP-7 | Clean-Up Room, 7m2 | FFE | MMIC-122 | DISINFECTOR: ultrasound probes | 1 | optional
CLUP-7 | Clean-Up Room, 7m2 | SE | HYDR-004 | DRAIN: floor waste | 1
CLUP-7 | Clean-Up Room, 7m2 | SE | ELBO-015 | SWITCH: light, single | 1
CLUP-7 | Clean-Up Room, 7m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
COMM | Communications Room | FA | FLVY-001 | FLOOR FINISH: vinyl, seamless, coved, anti-static | 1
COMM | Communications Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
COMM | Communications Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
COMM | Communications Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
COMM | Communications Room | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
COMM | Communications Room | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1
COMM | Communications Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
COMM | Communications Room | FFE | ITNE-052 | TELEPHONE: handset, wall mount | 1
COMM | Communications Room | FFE | ITNE-005 | COMMS BOARD | 1 | to suit project requirements
COMM | Communications Room | FFE | ITIN-123 | RACK: communications | 1 | to suit project requirements
COMM | Communications Room | FFE | ITIN-128 | RACK: voice, intermediate distribution frame (IDF) | 1 | optional, to suit project requirements
COMM | Communications Room | FFE | ITIN-139 | COMPUTER: mainframe | 1 | to suit project requirements
COMM | Communications Room | FFE | ITNE-011 | COMPUTER: server | 1 | to suit project requirements
COMM | Communications Room | FFE | FQGE-095 | STEP: safety | 1
COMM | Communications Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to suit project requirements
COMM | Communications Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | to suit project requirements
COMM | Communications Room | SE | ELBO-015 | SWITCH: light, single | 1
COMM | Communications Room | SE | ELGP-208 | GPO: double, wall mounted | 3 | to suit project requirements
COMM | Communications Room | SE | ITSE-046 | SENSOR: motion detection, for staff alert | 1
COMM | Communications Room | SE | ITSE-038 | READER: access control, general | 1
COMM | Communications Room | SE | ITSE-020 | CONTROL PANEL: duress | 1 | optional, to suit project requirements
COMM | Communications Room | SE | ITBU-015 | HEAD END: IPTV | 1 | optional, to suit project requirements
COMM | Communications Room | SE | ITBU-016 | HEAD END: MATV | 1 | optional, to suit project requirements
COMM | Communications Room | SE | ELGP-235 | GPO: UPS, double, wall mounted | 4 | to suit project requirements, power may be wall mounted or delivered via ceiling mounted umbilicals for access depending on server selection.
CONS | Consult Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
CONS | Consult Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
CONS | Consult Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
CONS | Consult Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
CONS | Consult Room | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
CONS | Consult Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CONS | Consult Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 2 | to exam couch
CONS | Consult Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
CONS | Consult Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | 1 1/2 leaf door is also acceptable as required
CONS | Consult Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 1 | to patient/visitor chairs
CONS | Consult Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
CONS | Consult Room | FA | DWGL-077 | DOOR: 850mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional, second door for egress
CONS | Consult Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | optional, provide if medical air required
CONS | Consult Room | FFE | FIDI-003 | DISPENSER: bedroll | 1 | may be provided as an attachment to exam couch
CONS | Consult Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1 | optional, provide if suction required
CONS | Consult Room | FFE | FIHR-016 | BRACKET: for sharps bin | 1
CONS | Consult Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
CONS | Consult Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CONS | Consult Room | FFE | FIHR-026 | HOOK: coat | 2
CONS | Consult Room | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1
CONS | Consult Room | FFE | FQGE-010 | CURTAIN: privacy screen | 1
CONS | Consult Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CONS | Consult Room | FFE | MMGE-030 | DIAGNOSTIC SET: portable | 1 | may be wall mounted/mobile on stand
CONS | Consult Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
CONS | Consult Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CONS | Consult Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CONS | Consult Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
CONS | Consult Room | FFE | FIBM-006 | MIRROR: fixed, glass | 1
CONS | Consult Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1 | optional, provide if oxygen required
CONS | Consult Room | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
CONS | Consult Room | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
CONS | Consult Room | FFE | MMGE-002 | ADAPTER: suction | 1 | optional, provide if suction required
CONS | Consult Room | FFE | MMGE-100 | SUCTION BOTTLE | 1 | optional, provide if suction required
CONS | Consult Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
CONS | Consult Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
CONS | Consult Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
CONS | Consult Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | high resolution monitor for PACS viewing
CONS | Consult Room | FFE | MMBE-023 | COUCH: examination, electric, mobile | 1 | hi/low operation
CONS | Consult Room | FFE | LIFX-026 | LIGHT: examination, wall mounted | 1 | may be ceiling mounted
CONS | Consult Room | FFE | ITNE-023 | PRINTER: networked | 1 | optional
CONS | Consult Room | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1
CONS | Consult Room | FFE | FQDW-037 | DESK: adj height, 1500W x 750D x 750H nom. | 1
CONS | Consult Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
CONS | Consult Room | FFE | FQBS-058 | CHAIR: visitor, patient, with arms | 3 | consider weight capacity of chairs
CONS | Consult Room | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | to desk
CONS | Consult Room | SE | ELPR-006 | RCD: residual current device | 3
CONS | Consult Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
CONS | Consult Room | SE | ELBO-015 | SWITCH: light, single | 2 | 1 to exam light
CONS | Consult Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | optional
CONS | Consult Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | optional
CONS | Consult Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | optional
CONS | Consult Room | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1
CONS | Consult Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
CONS | Consult Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level, for electric couch
CONS | Consult Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | for exam light
CONS | Consult Room | SE | ELGP-208 | GPO: double, wall mounted | 3 | 2x below desk, 1x above desk
CONS | Consult Room | SE | ITSE-006 | BUTTON: duress, fixed | 1 | in discreet location easily accessible by staff
CONS | Consult Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
CONS | Consult Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
CONS | Consult Room | SE | ELGP-205 | GPO: double, on services panel | 1
CONS | Consult Room | SE | ELGP-222 | GPO: emergency power, double, on services panel | 1
CONS-UN | Consult Room - Universal Access | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
CONS-UN | Consult Room - Universal Access | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
CONS-UN | Consult Room - Universal Access | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
CONS-UN | Consult Room - Universal Access | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
CONS-UN | Consult Room - Universal Access | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
CONS-UN | Consult Room - Universal Access | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
CONS-UN | Consult Room - Universal Access | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1 | to exam couch
CONS-UN | Consult Room - Universal Access | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
CONS-UN | Consult Room - Universal Access | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional, second door for egress
CONS-UN | Consult Room - Universal Access | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
CONS-UN | Consult Room - Universal Access | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 1 | to patient/visitor chairs
CONS-UN | Consult Room - Universal Access | FA | CLCN-009 | CORNICE: wall trim fixing | 1
CONS-UN | Consult Room - Universal Access | FFE | MGFP-002 | FLOWMETER: medical air | 1 | optional, provide if medical air required
CONS-UN | Consult Room - Universal Access | FFE | FIDI-003 | DISPENSER: bedroll | 1 | may be provided as an attachment to exam couch
CONS-UN | Consult Room - Universal Access | FFE | FIHR-019 | BRACKET: for suction bottle | 1 | optional, provide if suction required
CONS-UN | Consult Room - Universal Access | FFE | FIHR-016 | BRACKET: for sharps bin | 1
CONS-UN | Consult Room - Universal Access | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
CONS-UN | Consult Room - Universal Access | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CONS-UN | Consult Room - Universal Access | FFE | FIHR-026 | HOOK: coat | 2
CONS-UN | Consult Room - Universal Access | FFE | FIRT-031 | TRACK: curtain, privacy | 1
CONS-UN | Consult Room - Universal Access | FFE | FQGE-010 | CURTAIN: privacy screen | 2
CONS-UN | Consult Room - Universal Access | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CONS-UN | Consult Room - Universal Access | FFE | MMBE-054 | LIFTER: patient, ceiling mounted, sling component | 1 | optional, for patient lifter motor
CONS-UN | Consult Room - Universal Access | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
CONS-UN | Consult Room - Universal Access | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CONS-UN | Consult Room - Universal Access | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CONS-UN | Consult Room - Universal Access | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
CONS-UN | Consult Room - Universal Access | FFE | FIBM-007 | MIRROR: fixed, glass, full length | 1
CONS-UN | Consult Room - Universal Access | FFE | MGFP-004 | FLOWMETER: oxygen | 1 | optional, provide if oxygen required
CONS-UN | Consult Room - Universal Access | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
CONS-UN | Consult Room - Universal Access | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
CONS-UN | Consult Room - Universal Access | FFE | MMGE-002 | ADAPTER: suction | 1 | optional, provide if suction required
CONS-UN | Consult Room - Universal Access | FFE | MMGE-100 | SUCTION BOTTLE | 1 | optional, provide if suction required
CONS-UN | Consult Room - Universal Access | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
CONS-UN | Consult Room - Universal Access | FFE | HYBA-010 | BASIN: handwash, Type B | 1
CONS-UN | Consult Room - Universal Access | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
CONS-UN | Consult Room - Universal Access | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | high resolution monitor for PACS viewing
CONS-UN | Consult Room - Universal Access | FFE | MMGE-031 | DIAGNOSTIC SET: wall mounted | 1 | may be portable/mobile on stand
CONS-UN | Consult Room - Universal Access | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | may be wall mounted
CONS-UN | Consult Room - Universal Access | FFE | ITNE-023 | PRINTER: networked | 1 | optional
CONS-UN | Consult Room - Universal Access | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1
CONS-UN | Consult Room - Universal Access | FFE | MMBE-052 | LIFTER: patient, ceiling mounted, cassette motor for track | 1 | for patient lifter track
CONS-UN | Consult Room - Universal Access | FFE | FQDW-037 | DESK: adj height, 1500W x 750D x 750H nom. | 1
CONS-UN | Consult Room - Universal Access | FFE | MMBE-056 | LIFTER: patient, ceiling mounted, track, linear | 1
CONS-UN | Consult Room - Universal Access | FFE | MMBE-064 | COUCH: examination, electric, mobile, bariatric | 1 | hi/low operation
CONS-UN | Consult Room - Universal Access | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
CONS-UN | Consult Room - Universal Access | FFE | FQBS-058 | CHAIR: visitor, patient, with arms | 3 | consider weight capacity of chairs
CONS-UN | Consult Room - Universal Access | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | to desk
CONS-UN | Consult Room - Universal Access | SE | ELPR-006 | RCD: residual current device | 3
CONS-UN | Consult Room - Universal Access | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | for exam light
CONS-UN | Consult Room - Universal Access | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
CONS-UN | Consult Room - Universal Access | SE | ELBO-015 | SWITCH: light, single | 2
CONS-UN | Consult Room - Universal Access | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | optional
CONS-UN | Consult Room - Universal Access | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | optional
CONS-UN | Consult Room - Universal Access | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | optional
CONS-UN | Consult Room - Universal Access | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1
CONS-UN | Consult Room - Universal Access | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
CONS-UN | Consult Room - Universal Access | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level, for electric couch
CONS-UN | Consult Room - Universal Access | SE | ELGP-208 | GPO: double, wall mounted | 3 | 2x below desk, 1x above desk
CONS-UN | Consult Room - Universal Access | SE | ITSE-006 | BUTTON: duress, fixed | 1 | in discreet location easily accessible by staff
CONS-UN | Consult Room - Universal Access | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
CONS-UN | Consult Room - Universal Access | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
CONS-UN | Consult Room - Universal Access | SE | ELGP-105 | GPO: single, on services panel | 1 | for diagnostic set
CONS-UN | Consult Room - Universal Access | SE | ELGP-205 | GPO: double, on services panel | 1
CONS-UN | Consult Room - Universal Access | SE | ELGP-222 | GPO: emergency power, double, on services panel | 1
CORR | Corridor - Patient Treatment Areas | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CORR | Corridor - Patient Treatment Areas | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 16
CORR | Corridor - Patient Treatment Areas | FA | WLPR-006 | WALL PROTECTION: corner guards | 6 | coordinated with adjacent wall protection
CORR | Corridor - Patient Treatment Areas | FA | FIRT-010 | HANDRAIL | 4 | provided to at least one side of the corridor, to be continuous wherever practical; may be combined with crash rail if provided
CORR | Corridor - Patient Treatment Areas | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
CORR | Corridor - Patient Treatment Areas | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
CORR | Corridor - Patient Treatment Areas | FA | CLCN-009 | CORNICE: wall trim fixing | 1
CORR | Corridor - Patient Treatment Areas | FA | WLPR-026 | WALL PROTECTION: rigid wall sheeting | 12 | height dependent on service requirements, crash rails are also acceptable
CORR | Corridor - Patient Treatment Areas | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 | quantity to suit layout and infection control policies
CORR | Corridor - Patient Treatment Areas | SE | ITCL-001 | ANNUNCIATOR: nurse call | 1 | ceiling mounted; quantity to suit layout and service requirements
CORR | Corridor - Patient Treatment Areas | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 | quantity and location to comply with mandatory requirements.
CTPR | CT Planning Room | FA | FLVY-001 | FLOOR FINISH: vinyl, seamless, coved, anti-static | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
CTPR | CT Planning Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
CTPR | CT Planning Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
CTPR | CT Planning Room | FA | WLWA-007 | WALL FINISH: splashback | 1
CTPR | CT Planning Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
CTPR | CT Planning Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 7
CTPR | CT Planning Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
CTPR | CT Planning Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | safety glass, to control room
CTPR | CT Planning Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1 | to control room
CTPR | CT Planning Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
CTPR | CT Planning Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
CTPR | CT Planning Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
CTPR | CT Planning Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CTPR | CT Planning Room | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 3
CTPR | CT Planning Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CTPR | CT Planning Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
CTPR | CT Planning Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CTPR | CT Planning Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CTPR | CT Planning Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
CTPR | CT Planning Room | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1 | for patient slide
CTPR | CT Planning Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
CTPR | CT Planning Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1
CTPR | CT Planning Room | FFE | MMBE-042 | SLIDE: patient | 1 | patient transfer devices to be based on operational requirements and WHS policies
CTPR | CT Planning Room | FFE | MMGE-139 | TROLLEY: procedure | 1
CTPR | CT Planning Room | FFE | MMGE-142 | TROLLEY: sharps bin | 1
CTPR | CT Planning Room | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 2
CTPR | CT Planning Room | FFE | MMGE-002 | ADAPTER: suction | 1
CTPR | CT Planning Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
CTPR | CT Planning Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
CTPR | CT Planning Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
CTPR | CT Planning Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
CTPR | CT Planning Room | FFE | ITCL-012 | CAMERA: observation | 2 | 1 x optional, for patient observation or visualisation of additional monitoring equipment, pan, tilt and zoom (PTZ) capability is required, quantity to be based on equipment and service requirements
CTPR | CT Planning Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
CTPR | CT Planning Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
CTPR | CT Planning Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
CTPR | CT Planning Room | FFE | MMMI-105 | TABLE: patient, fixed, CT | 1
CTPR | CT Planning Room | FFE | MMMI-107 | INJECTOR: contrast, CT, ceiling mounted | 1 | may be mobile
CTPR | CT Planning Room | FFE | MMMI-109 | DEVICE: imaging, CT, wide bore | 1
CTPR | CT Planning Room | FFE | FIJO-166 | CUPBOARD: tall, shelving, horizontal patient device storage, radiotherapy | 2
CTPR | CT Planning Room | FFE | FIJO-167 | CUPBOARD: tall, shelving, vertical patient device storage, radiotherapy | 1
CTPR | CT Planning Room | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead is also acceptable
CTPR | CT Planning Room | FFE | MMMI-377 | WARMER: contrast | 1 | may be located in control room
CTPR | CT Planning Room | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 1
CTPR | CT Planning Room | FFE | FIHR-078 | BRACKET: for observation camera, wall mounted | 1
CTPR | CT Planning Room | FFE | MMRA-008 | DEVICE: surface guided radiation therapy system | 1 | as required for respiratory gating / motion management / surface guidance systems, to suit service requirements
CTPR | CT Planning Room | FFE | MMRA-014 | LASER: radiotherapy, bridge | 1 | wall and ceiling mounted lasers are also acceptable
CTPR | CT Planning Room | FFE | MMRA-020 | OVEN: thermoplastic warming | 1
CTPR | CT Planning Room | FFE | FIHR-092 | HOOK: anchor, patient ladder | 1
CTPR | CT Planning Room | FFE | MMRA-019 | LASER: radiotherapy, couch tracking | 1
CTPR | CT Planning Room | FFE | FIHR-089 | BRACKET: surface guided radiation therapy system, ceiling mounted | 1
CTPR | CT Planning Room | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 2 | 1 x optional, to additional patient observation camera if provided
CTPR | CT Planning Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
CTPR | CT Planning Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
CTPR | CT Planning Room | SE | ELPR-006 | RCD: residual current device | 3
CTPR | CT Planning Room | SE | ELGP-101 | GPO: single, ceiling mounted | 2 | 1 x optional, to additional patient observation camera if provided
CTPR | CT Planning Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
CTPR | CT Planning Room | SE | ELBO-001 | BUTTON: emergency stop | 1
CTPR | CT Planning Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 2 | outside room doors, auto-on when modality in use
CTPR | CT Planning Room | SE | ELBO-015 | SWITCH: light, single | 1
CTPR | CT Planning Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
CTPR | CT Planning Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1
CTPR | CT Planning Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
CTPR | CT Planning Room | SE | ELGP-109 | GPO: single, wall mounted | 2 | to clock
CTPR | CT Planning Room | SE | ELGP-208 | GPO: double, wall mounted | 4
CTPR | CT Planning Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
CTPR | CT Planning Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
CTPR | CT Planning Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors, shielding to be reviewed and confirmed by Radiation Consultant
CTPR | CT Planning Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
CTPR | CT Planning Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
CTPR | CT Planning Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3
CTPR | CT Planning Room | SE | LIFX-038 | LIGHT: backlit feature panels | 2 | desirable
DEN-MLB | Dental Laboratory, Minor | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent under basin and around sink
DEN-MLB | Dental Laboratory, Minor | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
DEN-MLB | Dental Laboratory, Minor | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DEN-MLB | Dental Laboratory, Minor | FA | WLWA-007 | WALL FINISH: splashback | 1
DEN-MLB | Dental Laboratory, Minor | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DEN-MLB | Dental Laboratory, Minor | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 2
DEN-MLB | Dental Laboratory, Minor | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
DEN-MLB | Dental Laboratory, Minor | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
DEN-MLB | Dental Laboratory, Minor | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DEN-MLB | Dental Laboratory, Minor | FFE | ITNE-052 | TELEPHONE: handset, wall mount | 1
DEN-MLB | Dental Laboratory, Minor | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 2
DEN-MLB | Dental Laboratory, Minor | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
DEN-MLB | Dental Laboratory, Minor | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DEN-MLB | Dental Laboratory, Minor | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
DEN-MLB | Dental Laboratory, Minor | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DEN-MLB | Dental Laboratory, Minor | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DEN-MLB | Dental Laboratory, Minor | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 2
DEN-MLB | Dental Laboratory, Minor | FFE | FIJO-197 | SHELF: stainless steel, underbench | 4 | some shelves may be removable or omitted to allow for storage of bulk items on mobile dolly
DEN-MLB | Dental Laboratory, Minor | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 5
DEN-MLB | Dental Laboratory, Minor | FFE | FQBS-052 | STOOL: mobile, high, adj height | 1
DEN-MLB | Dental Laboratory, Minor | FFE | FQSN-006 | BIN: storage bins, for wall mounted panel | 10 | various sizes
DEN-MLB | Dental Laboratory, Minor | FFE | MMDN-004 | DEVICE: dental, micromotor | 1 | optional
DEN-MLB | Dental Laboratory, Minor | FFE | MMDN-005 | DEVICE: dental, model trimmer | 1 | plumbed to sink
DEN-MLB | Dental Laboratory, Minor | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DEN-MLB | Dental Laboratory, Minor | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DEN-MLB | Dental Laboratory, Minor | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
DEN-MLB | Dental Laboratory, Minor | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | optional; to workstation if provided
DEN-MLB | Dental Laboratory, Minor | FFE | HYDR-008 | TRAP: plaster | 1
DEN-MLB | Dental Laboratory, Minor | FFE | MMDN-009 | LATHE: dental, polishing | 1 | extraction is required if not integral to equipment
DEN-MLB | Dental Laboratory, Minor | FFE | MMDN-013 | MOULDER: dental, forming machine | 1
DEN-MLB | Dental Laboratory, Minor | FFE | HYBA-052 | SINK: single, 450 x 350 x 250D | 1
DEN-MLB | Dental Laboratory, Minor | FFE | FIJO-100 | BENCH LEGS: stainless steel, for bench top | 2
DEN-MLB | Dental Laboratory, Minor | FFE | ITNE-003 | BAR CODING: reader | 1 | optional; to workstation if provided
DEN-MLB | Dental Laboratory, Minor | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead is also acceptable
DEN-MLB | Dental Laboratory, Minor | FFE | FQDW-052 | DESK: adj height, 1000W x 600D x 750H nom. | 1 | optional; to workstation if provided
DEN-MLB | Dental Laboratory, Minor | FFE | MMDN-023 | DEVICE: dental extraction unit | 1
DEN-MLB | Dental Laboratory, Minor | FFE | MMDN-024 | DEVICE: dental, mixing vibrator | 1 | a vacuum mixer may also be considered
DEN-MLB | Dental Laboratory, Minor | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
DEN-MLB | Dental Laboratory, Minor | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
DEN-MLB | Dental Laboratory, Minor | SE | HYDR-004 | DRAIN: floor waste | 1
DEN-MLB | Dental Laboratory, Minor | SE | LIFX-008 | LIGHT: task light, built in | 5
DEN-MLB | Dental Laboratory, Minor | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 2 x task light
DEN-MLB | Dental Laboratory, Minor | SE | MGAS-008 | OUTLET: compressed air, wall mounted | 1 | to forming machine
DEN-MLB | Dental Laboratory, Minor | SE | ELGP-109 | GPO: single, wall mounted | 1
DEN-MLB | Dental Laboratory, Minor | SE | ELGP-208 | GPO: double, wall mounted | 8 | 2 x optional; to workstation if provided
DENSR-1 | Dental Surgery, Type 1 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
DENSR-1 | Dental Surgery, Type 1 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
DENSR-1 | Dental Surgery, Type 1 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DENSR-1 | Dental Surgery, Type 1 | FA | WLWA-007 | WALL FINISH: splashback | 1
DENSR-1 | Dental Surgery, Type 1 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DENSR-1 | Dental Surgery, Type 1 | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
DENSR-1 | Dental Surgery, Type 1 | FA | DWWH-005 | WINDOW | 1 | desirable
DENSR-1 | Dental Surgery, Type 1 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
DENSR-1 | Dental Surgery, Type 1 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | DWWD-007 | BLIND: venetian, integral | 1 | to external window if provided
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-019 | BRACKET: for suction bottle | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1 | below sink
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 2 | optional, extent of storage dependent on service requirements and approach to consumables storage
DENSR-1 | Dental Surgery, Type 1 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-014 | DISPENSER: disposable glove | 2 | 1 x to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-021 | DISPENSER: mask | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MGFP-004 | FLOWMETER: oxygen | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQBS-051 | STOOL: mobile, adj height | 2
DENSR-1 | Dental Surgery, Type 1 | FFE | MMGE-002 | ADAPTER: suction | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMGE-100 | SUCTION BOTTLE | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-016 | TROLLEY: dental | 2
DENSR-1 | Dental Surgery, Type 1 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | optional, to sink if provided
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-002 | CHAIR: dental, electric | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-018 | X-RAY UNIT: dental, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-020 | LIGHT: dental, ceiling mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 2
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-099 | CAC: cable access cap | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-212 | BULKHEAD: joinery | 2 | optional, over wall mounted joinery if provided; plasterboard bulkhead is also acceptable
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-213 | PANEL: joinery carcass end | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | HYBA-079 | SINK: single, dental | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-091 | RACK: lead apron, single | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-022 | CONSOLE: dentist, chair mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-021 | CONSOLE: dental assistant, chair mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-025 | MONITOR: dentist, chair mounted | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-088 | BRACKET: for monitor, desk mounted, articulated, adj height | 1 | articulated with smooth action, reach to be sufficient for left handed and right handed setup
DENSR-1 | Dental Surgery, Type 1 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ELPR-006 | RCD: residual current device | 2
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | SE | HYTP-008 | OUTLET: cold water | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | HYDR-001 | DRAIN: special requirements | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | LIFX-008 | LIGHT: task light, built in | 2 | optional, under wall mounted joinery if provided
DENSR-1 | Dental Surgery, Type 1 | SE | ELBO-015 | SWITCH: light, single | 1 | 1 x general room lighting, 2 x optional, to task lighting if provided
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-062 | OUTLET: suction, on services panel | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | optional, depending on service requirements
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-208 | GPO: double, wall mounted | 6
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to x-ray unit
DENSR-1 | Dental Surgery, Type 1 | SE | ITSE-006 | BUTTON: duress, fixed | 1
DENSR-1 | Dental Surgery, Type 1 | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-020 | OUTLET: medical air (MA) | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-090 | OUTLET: vacuum | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | ITIN-010 | OUTLET: data, single RJ45 | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | ELGE-004 | DIRECT CONNECTION: power, emergency | 1 | to dental chair
DENSR-2 | Dental Surgery, Type 2 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
DENSR-2 | Dental Surgery, Type 2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 8
DENSR-2 | Dental Surgery, Type 2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DENSR-2 | Dental Surgery, Type 2 | FA | WLWA-007 | WALL FINISH: splashback | 1
DENSR-2 | Dental Surgery, Type 2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DENSR-2 | Dental Surgery, Type 2 | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | optional
DENSR-2 | Dental Surgery, Type 2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
DENSR-2 | Dental Surgery, Type 2 | FA | DWWH-005 | WINDOW | 1 | desirable
DENSR-2 | Dental Surgery, Type 2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
DENSR-2 | Dental Surgery, Type 2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | DWWD-007 | BLIND: venetian, integral | 1 | to external window if provided
DENSR-2 | Dental Surgery, Type 2 | FFE | FIHR-019 | BRACKET: for suction bottle | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 2 | optional, extent of storage dependent on service requirements and approach to consumables storage
DENSR-2 | Dental Surgery, Type 2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DENSR-2 | Dental Surgery, Type 2 | FFE | FIDI-014 | DISPENSER: disposable glove | 2 | 1 x to basin
DENSR-2 | Dental Surgery, Type 2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DENSR-2 | Dental Surgery, Type 2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DENSR-2 | Dental Surgery, Type 2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DENSR-2 | Dental Surgery, Type 2 | FFE | FIDI-021 | DISPENSER: mask | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MGFP-004 | FLOWMETER: oxygen | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FQBS-051 | STOOL: mobile, adj height | 2
DENSR-2 | Dental Surgery, Type 2 | FFE | MMGE-002 | ADAPTER: suction | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MMGE-100 | SUCTION BOTTLE | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MMDN-016 | TROLLEY: dental | 2
DENSR-2 | Dental Surgery, Type 2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | optional, to sink if provided
DENSR-2 | Dental Surgery, Type 2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MMDN-002 | CHAIR: dental, electric | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MMDN-018 | X-RAY UNIT: dental, wall mounted | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MMDN-020 | LIGHT: dental, ceiling mounted | 1 | may be chair mounted
DENSR-2 | Dental Surgery, Type 2 | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIJO-099 | CAC: cable access cap | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | FIJO-212 | BULKHEAD: joinery | 2 | optional, over wall mounted joinery if provided; plasterboard bulkhead is also acceptable
DENSR-2 | Dental Surgery, Type 2 | FFE | FIJO-213 | PANEL: joinery carcass end | 2
DENSR-2 | Dental Surgery, Type 2 | FFE | HYBA-079 | SINK: single, dental | 1 | optional
DENSR-2 | Dental Surgery, Type 2 | FFE | FIHR-091 | RACK: lead apron, single | 1 | optional
DENSR-2 | Dental Surgery, Type 2 | FFE | MMDN-022 | CONSOLE: dentist, chair mounted | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MMDN-021 | CONSOLE: dental assistant, chair mounted | 1
DENSR-2 | Dental Surgery, Type 2 | FFE | MMDN-025 | MONITOR: dentist, chair mounted | 1 | optional
DENSR-2 | Dental Surgery, Type 2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
DENSR-2 | Dental Surgery, Type 2 | SE | ELPR-006 | RCD: residual current device | 2
DENSR-2 | Dental Surgery, Type 2 | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional
DENSR-2 | Dental Surgery, Type 2 | SE | HYTP-008 | OUTLET: cold water | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-2 | Dental Surgery, Type 2 | SE | HYDR-001 | DRAIN: special requirements | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-2 | Dental Surgery, Type 2 | SE | LIFX-008 | LIGHT: task light, built in | 2 | optional, under wall mounted joinery if provided
DENSR-2 | Dental Surgery, Type 2 | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 2 x optional, to task lighting if provided
DENSR-2 | Dental Surgery, Type 2 | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
DENSR-2 | Dental Surgery, Type 2 | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional
DENSR-2 | Dental Surgery, Type 2 | SE | MGAS-062 | OUTLET: suction, on services panel | 1
DENSR-2 | Dental Surgery, Type 2 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | optional, depending on service requirements
DENSR-2 | Dental Surgery, Type 2 | SE | ELGP-208 | GPO: double, wall mounted | 6
DENSR-2 | Dental Surgery, Type 2 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
DENSR-2 | Dental Surgery, Type 2 | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to x-ray unit
DENSR-2 | Dental Surgery, Type 2 | SE | ITSE-006 | BUTTON: duress, fixed | 2
DENSR-2 | Dental Surgery, Type 2 | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
DENSR-2 | Dental Surgery, Type 2 | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1
DENSR-2 | Dental Surgery, Type 2 | SE | MGAS-020 | OUTLET: medical air (MA) | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-2 | Dental Surgery, Type 2 | SE | MGAS-090 | OUTLET: vacuum | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-2 | Dental Surgery, Type 2 | SE | ITIN-010 | OUTLET: data, single RJ45 | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-2 | Dental Surgery, Type 2 | SE | ELGE-004 | DIRECT CONNECTION: power, emergency | 1 | to dental chair
DINR | Dining - Patient | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10/pendulum P3 or agreed equivalent to extend under basin; significant change in colour to be avoided
DINR | Dining - Patient | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
DINR | Dining - Patient | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DINR | Dining - Patient | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
DINR | Dining - Patient | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
DINR | Dining - Patient | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
DINR | Dining - Patient | FA | DWWH-005 | WINDOW | 2 | natural light and external outlook is essential
DINR | Dining - Patient | FA | DWGL-031 | DOOR: 1400mm c/o, 1 1/2 leaf, half glazed, solid core/glass, paint | 1 | with hold open device to support trolley movement as required
DINR | Dining - Patient | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 2 | to meal trolley parking area
DINR | Dining - Patient | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DINR | Dining - Patient | FFE | FQGE-002 | ARTWORK | 1
DINR | Dining - Patient | FFE | FQBS-015 | CHAIR: dining | 16 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
DINR | Dining - Patient | FFE | DWWD-004 | BLIND: roller | 2
DINR | Dining - Patient | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2
DINR | Dining - Patient | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DINR | Dining - Patient | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
DINR | Dining - Patient | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DINR | Dining - Patient | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DINR | Dining - Patient | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DINR | Dining - Patient | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1 | optional
DINR | Dining - Patient | FFE | FIJO-104 | CUPBOARD: drawer unit, 3 equal | 1
DINR | Dining - Patient | FFE | FQTA-059 | TABLE: square, 1100W, dining height | 4 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
DINR | Dining - Patient | FFE | FQCW-019 | BIN: waste, general, 50L | 1 | size, type and quantity of bins dependent on local waste management policies
DINR | Dining - Patient | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1 | optional
DINR | Dining - Patient | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DINR | Dining - Patient | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DINR | Dining - Patient | FFE | FQCA-082 | TROLLEY: food delivery, hot/cold | 1 | as required to suit food service model
DINR | Dining - Patient | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
DINR | Dining - Patient | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
DINR | Dining - Patient | FFE | FQCW-062 | BIN: waste, recycling, comingled, 50L | 1 | size, type and quantity of bins dependent on local waste management policies
DINR | Dining - Patient | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to wall mounted telephone
DINR | Dining - Patient | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | optional, to television, MATV outlet may also be provided depending on local ICT policies
DINR | Dining - Patient | SE | ITCL-008 | BUTTON: nurse call, staff assist | 2 | with indicator button and light
DINR | Dining - Patient | SE | ELBO-015 | SWITCH: light, single | 1
DINR | Dining - Patient | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 2 | optional, dependent on patient cohort, provide in areas requiring high staff observation; with indicator button and light
DINR | Dining - Patient | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | optional, dependent on patient cohort, provide in areas requiring high staff observation; with indicator button and light
DINR | Dining - Patient | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
DINR | Dining - Patient | SE | ELGP-109 | GPO: single, wall mounted | 1 | to meal trolley parking area; power requirements dependent on type of meal trolley
DINR | Dining - Patient | SE | ELGP-208 | GPO: double, wall mounted | 3
DINR | Dining - Patient | SE | ELGP-208 | GPO: double, wall mounted | 1 | optional, to television
DINR | Dining - Patient | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
DISP-10 | Disposal Room, 10m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
DISP-10 | Disposal Room, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
DISP-10 | Disposal Room, 10m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
DISP-10 | Disposal Room, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
DISP-10 | Disposal Room, 10m2 | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
DISP-10 | Disposal Room, 10m2 | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | lockable or secure staff access, hold open hardware may be required
DISP-10 | Disposal Room, 10m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 1
DISP-10 | Disposal Room, 10m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DISP-10 | Disposal Room, 10m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DISP-10 | Disposal Room, 10m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DISP-10 | Disposal Room, 10m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin, locate with alcohol based hand rub if basin is located outside of room
DISP-10 | Disposal Room, 10m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DISP-10 | Disposal Room, 10m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DISP-10 | Disposal Room, 10m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DISP-10 | Disposal Room, 10m2 | FFE | FIJO-194 | SHELF: laminate | 1 | optional, for storage of small sharps bins
DISP-10 | Disposal Room, 10m2 | FFE | MMGE-143 | TROLLEY: soiled linen collection | 1
DISP-10 | Disposal Room, 10m2 | FFE | FQCW-029 | BIN: waste, recycling, paper, 240L | 1 | size, type and quantity of bins dependent on service requirements and local waste management policies
DISP-10 | Disposal Room, 10m2 | FFE | FQCW-006 | BIN: waste, clinical, 240L | 1 | size, type and quantity of bins dependent on service requirements and local waste management policies
DISP-10 | Disposal Room, 10m2 | FFE | FQCW-022 | BIN: waste, general, skip 660 litre | 1 | size, type and quantity of bins dependent on service requirements and local waste management policies
DISP-10 | Disposal Room, 10m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | may be provided outside of room depending on local infection control policies
DISP-10 | Disposal Room, 10m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DISP-10 | Disposal Room, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
DISP-15 | Disposal Room, 15m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
DISP-15 | Disposal Room, 15m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
DISP-15 | Disposal Room, 15m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 4 | 2 optional to optional door
DISP-15 | Disposal Room, 15m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
DISP-15 | Disposal Room, 15m2 | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
DISP-15 | Disposal Room, 15m2 | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 2 | 1 door optional, lockable or secure staff access, hold open hardware may be required, dual access may be provided depending on overall planning arrangement
DISP-15 | Disposal Room, 15m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 1
DISP-15 | Disposal Room, 15m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DISP-15 | Disposal Room, 15m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DISP-15 | Disposal Room, 15m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DISP-15 | Disposal Room, 15m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin, locate with alcohol based hand rub if basin located outside of room
DISP-15 | Disposal Room, 15m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DISP-15 | Disposal Room, 15m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DISP-15 | Disposal Room, 15m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DISP-15 | Disposal Room, 15m2 | FFE | FIJO-194 | SHELF: laminate | 1 | optional, for storage of small sharps bins
DISP-15 | Disposal Room, 15m2 | FFE | MMGE-143 | TROLLEY: soiled linen collection | 1
DISP-15 | Disposal Room, 15m2 | FFE | FQCW-017 | BIN: waste, general, 240L | 1 | size, type and quantity of bins dependent on service requirements and local waste management policies
DISP-15 | Disposal Room, 15m2 | FFE | FQCW-029 | BIN: waste, recycling, paper, 240L | 1 | size, type and quantity of bins dependent on service requirements and local waste management policies
DISP-15 | Disposal Room, 15m2 | FFE | FQCW-006 | BIN: waste, clinical, 240L | 2 | size, type and quantity of bins dependent on service requirements and local waste management policies
DISP-15 | Disposal Room, 15m2 | FFE | FQCW-022 | BIN: waste, general, skip 660 litre | 2 | size, type and quantity of bins dependent on service requirements and local waste management policies
DISP-15 | Disposal Room, 15m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | may be provided outside of room depending on local infection control policies
DISP-15 | Disposal Room, 15m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DISP-15 | Disposal Room, 15m2 | SE | ELBO-015 | SWITCH: light, single | 1
DTUR-10 | Dirty Utility, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
DTUR-10 | Dirty Utility, 10m2 | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1
DTUR-10 | Dirty Utility, 10m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
DTUR-10 | Dirty Utility, 10m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DTUR-10 | Dirty Utility, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DTUR-10 | Dirty Utility, 10m2 | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
DTUR-10 | Dirty Utility, 10m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | pull handle/push plate, closer
DTUR-10 | Dirty Utility, 10m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 5
DTUR-10 | Dirty Utility, 10m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DTUR-10 | Dirty Utility, 10m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | with wet edge
DTUR-10 | Dirty Utility, 10m2 | FFE | FIHR-038 | RACK: bowl/basin | 1 | quantity to be based on service requirements
DTUR-10 | Dirty Utility, 10m2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
DTUR-10 | Dirty Utility, 10m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DTUR-10 | Dirty Utility, 10m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DTUR-10 | Dirty Utility, 10m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
DTUR-10 | Dirty Utility, 10m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DTUR-10 | Dirty Utility, 10m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DTUR-10 | Dirty Utility, 10m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DTUR-10 | Dirty Utility, 10m2 | FFE | FIHR-029 | HOOK: equipment | 2 | optional, for clean linen bags
DTUR-10 | Dirty Utility, 10m2 | FFE | FIHR-025 | HOOK: catheter bag | 2 | optional
DTUR-10 | Dirty Utility, 10m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 3
DTUR-10 | Dirty Utility, 10m2 | FFE | FIHR-053 | RACK: pan & urinal | 1 | quantity to be based on service requirements
DTUR-10 | Dirty Utility, 10m2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
DTUR-10 | Dirty Utility, 10m2 | FFE | FIJO-196 | SHELF: stainless steel, overbench, wall mount | 1
DTUR-10 | Dirty Utility, 10m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1
DTUR-10 | Dirty Utility, 10m2 | FFE | FQCW-017 | BIN: waste, general, 240L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
DTUR-10 | Dirty Utility, 10m2 | FFE | FQCW-006 | BIN: waste, clinical, 240L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
DTUR-10 | Dirty Utility, 10m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DTUR-10 | Dirty Utility, 10m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DTUR-10 | Dirty Utility, 10m2 | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
DTUR-10 | Dirty Utility, 10m2 | FFE | MMIC-095 | WASHER SANITISER: pan flusher | 1 | combined utensil/pan sanitiser may be provided
DTUR-10 | Dirty Utility, 10m2 | FFE | HYBA-052 | SINK: single, 450 x 350 x 250D | 1 | hot water
DTUR-10 | Dirty Utility, 10m2 | FFE | MMLA-013 | ANALYSER: urinalysis | 1 | optional
DTUR-10 | Dirty Utility, 10m2 | FFE | MMIC-097 | WASHER SANITISER: utensil washer | 1 | optional, combined utensil/pan sanitiser may be provided
DTUR-10 | Dirty Utility, 10m2 | FFE | FIJO-100 | BENCH LEGS: stainless steel, for bench top | 1
DTUR-10 | Dirty Utility, 10m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional, for urinalysis machine
DTUR-10 | Dirty Utility, 10m2 | SE | ELPR-003 | ISOLATOR: electrical, 1ph+n | 1
DTUR-10 | Dirty Utility, 10m2 | SE | HYTP-013 | OUTLET: hot water | 2 | for washer sanitisers
DTUR-10 | Dirty Utility, 10m2 | SE | HYTP-008 | OUTLET: cold water | 2 | for washer sanitisers
DTUR-10 | Dirty Utility, 10m2 | SE | HYDR-004 | DRAIN: floor waste | 1
DTUR-10 | Dirty Utility, 10m2 | SE | HYDR-006 | DRAIN: tundish | 1 | optional, dependent on machine requirements
DTUR-10 | Dirty Utility, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
DTUR-10 | Dirty Utility, 10m2 | SE | ELGP-106 | GPO: single, special, for equipment | 2 | for washer sanitisers
DTUR-10 | Dirty Utility, 10m2 | SE | ELGP-109.1 | GPO: single, wall mounted (V) | 1 | optional, for urinalysis machine
DTUR-10 | Dirty Utility, 10m2 | SE | HYGE-028 | DEVICE: backflow prevention | 1
DTUR-12 | Dirty Utility, 12m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
DTUR-12 | Dirty Utility, 12m2 | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1
DTUR-12 | Dirty Utility, 12m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
DTUR-12 | Dirty Utility, 12m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DTUR-12 | Dirty Utility, 12m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DTUR-12 | Dirty Utility, 12m2 | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
DTUR-12 | Dirty Utility, 12m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | pull handle/push plate, closer
DTUR-12 | Dirty Utility, 12m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 5
DTUR-12 | Dirty Utility, 12m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DTUR-12 | Dirty Utility, 12m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | with wet edge
DTUR-12 | Dirty Utility, 12m2 | FFE | FIHR-038 | RACK: bowl/basin | 1 | quantity to be based on service requirements
DTUR-12 | Dirty Utility, 12m2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
DTUR-12 | Dirty Utility, 12m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DTUR-12 | Dirty Utility, 12m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DTUR-12 | Dirty Utility, 12m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
DTUR-12 | Dirty Utility, 12m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DTUR-12 | Dirty Utility, 12m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DTUR-12 | Dirty Utility, 12m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DTUR-12 | Dirty Utility, 12m2 | FFE | FIHR-029 | HOOK: equipment | 2 | optional, for clean linen bags
DTUR-12 | Dirty Utility, 12m2 | FFE | FIHR-025 | HOOK: catheter bag | 2 | optional
DTUR-12 | Dirty Utility, 12m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 3
DTUR-12 | Dirty Utility, 12m2 | FFE | FIHR-053 | RACK: pan & urinal | 1 | quantity to be based on service requirements
DTUR-12 | Dirty Utility, 12m2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
DTUR-12 | Dirty Utility, 12m2 | FFE | FIJO-196 | SHELF: stainless steel, overbench, wall mount | 1
DTUR-12 | Dirty Utility, 12m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1
DTUR-12 | Dirty Utility, 12m2 | FFE | FQCW-017 | BIN: waste, general, 240L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
DTUR-12 | Dirty Utility, 12m2 | FFE | FQCW-006 | BIN: waste, clinical, 240L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
DTUR-12 | Dirty Utility, 12m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DTUR-12 | Dirty Utility, 12m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DTUR-12 | Dirty Utility, 12m2 | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
DTUR-12 | Dirty Utility, 12m2 | FFE | MMIC-095 | WASHER SANITISER: pan flusher | 1 | combined utensil/pan sanitiser may be provided
DTUR-12 | Dirty Utility, 12m2 | FFE | HYBA-052 | SINK: single, 450 x 350 x 250D | 1 | hot water
DTUR-12 | Dirty Utility, 12m2 | FFE | MMLA-013 | ANALYSER: urinalysis | 1 | optional
DTUR-12 | Dirty Utility, 12m2 | FFE | MMIC-097 | WASHER SANITISER: utensil washer | 1 | combined utensil/pan sanitiser may be provided
DTUR-12 | Dirty Utility, 12m2 | FFE | FIJO-100 | BENCH LEGS: stainless steel, for bench top | 1
DTUR-12 | Dirty Utility, 12m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional, for urinalysis machine
DTUR-12 | Dirty Utility, 12m2 | SE | ELPR-003 | ISOLATOR: electrical, 1ph+n | 1
DTUR-12 | Dirty Utility, 12m2 | SE | HYTP-013 | OUTLET: hot water | 2 | for washer sanitisers
DTUR-12 | Dirty Utility, 12m2 | SE | HYTP-008 | OUTLET: cold water | 2 | for washer sanitisers
DTUR-12 | Dirty Utility, 12m2 | SE | HYDR-004 | DRAIN: floor waste | 1
DTUR-12 | Dirty Utility, 12m2 | SE | HYDR-006 | DRAIN: tundish | 1 | optional, dependent on machine requirements
DTUR-12 | Dirty Utility, 12m2 | SE | ELBO-015 | SWITCH: light, single | 1
DTUR-12 | Dirty Utility, 12m2 | SE | ELGP-106 | GPO: single, special, for equipment | 2 | for washer sanitisers
DTUR-12 | Dirty Utility, 12m2 | SE | ELGP-109.1 | GPO: single, wall mounted (V) | 1 | optional, for urinalysis machine
DTUR-12 | Dirty Utility, 12m2 | SE | HYGE-028 | DEVICE: backflow prevention | 1
DTUR-14 | Dirty Utility, 14m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
DTUR-14 | Dirty Utility, 14m2 | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1
DTUR-14 | Dirty Utility, 14m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
DTUR-14 | Dirty Utility, 14m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
DTUR-14 | Dirty Utility, 14m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DTUR-14 | Dirty Utility, 14m2 | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
DTUR-14 | Dirty Utility, 14m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 2 | pull handle/push plate, closer
DTUR-14 | Dirty Utility, 14m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 6
DTUR-14 | Dirty Utility, 14m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DTUR-14 | Dirty Utility, 14m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | with wet edge
DTUR-14 | Dirty Utility, 14m2 | FFE | FIHR-038 | RACK: bowl/basin | 1 | quantity to be based on service requirements
DTUR-14 | Dirty Utility, 14m2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
DTUR-14 | Dirty Utility, 14m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2
DTUR-14 | Dirty Utility, 14m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DTUR-14 | Dirty Utility, 14m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
DTUR-14 | Dirty Utility, 14m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DTUR-14 | Dirty Utility, 14m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DTUR-14 | Dirty Utility, 14m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DTUR-14 | Dirty Utility, 14m2 | FFE | FIHR-029 | HOOK: equipment | 2 | optional, for clean linen bags
DTUR-14 | Dirty Utility, 14m2 | FFE | FIHR-025 | HOOK: catheter bag | 2 | optional
DTUR-14 | Dirty Utility, 14m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 3
DTUR-14 | Dirty Utility, 14m2 | FFE | FIHR-053 | RACK: pan & urinal | 1 | quantity to be based on service requirements
DTUR-14 | Dirty Utility, 14m2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
DTUR-14 | Dirty Utility, 14m2 | FFE | FIJO-196 | SHELF: stainless steel, overbench, wall mount | 1
DTUR-14 | Dirty Utility, 14m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1
DTUR-14 | Dirty Utility, 14m2 | FFE | FQCW-017 | BIN: waste, general, 240L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
DTUR-14 | Dirty Utility, 14m2 | FFE | FQCW-006 | BIN: waste, clinical, 240L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
DTUR-14 | Dirty Utility, 14m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DTUR-14 | Dirty Utility, 14m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DTUR-14 | Dirty Utility, 14m2 | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
DTUR-14 | Dirty Utility, 14m2 | FFE | MMIC-095 | WASHER SANITISER: pan flusher | 1 | combined utensil/pan sanitiser may be provided
DTUR-14 | Dirty Utility, 14m2 | FFE | HYBA-052 | SINK: single, 450 x 350 x 250D | 1 | hot water
DTUR-14 | Dirty Utility, 14m2 | FFE | MMLA-013 | ANALYSER: urinalysis | 1 | optional
DTUR-14 | Dirty Utility, 14m2 | FFE | MMIC-097 | WASHER SANITISER: utensil washer | 1 | combined utensil/pan sanitiser may be provided
DTUR-14 | Dirty Utility, 14m2 | FFE | FIJO-100 | BENCH LEGS: stainless steel, for bench top | 1
DTUR-14 | Dirty Utility, 14m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional, for urinalysis machine
DTUR-14 | Dirty Utility, 14m2 | SE | ELPR-003 | ISOLATOR: electrical, 1ph+n | 1
DTUR-14 | Dirty Utility, 14m2 | SE | HYTP-013 | OUTLET: hot water | 2 | for washer sanitisers
DTUR-14 | Dirty Utility, 14m2 | SE | HYTP-008 | OUTLET: cold water | 2 | for washer sanitisers
DTUR-14 | Dirty Utility, 14m2 | SE | HYDR-004 | DRAIN: floor waste | 1
DTUR-14 | Dirty Utility, 14m2 | SE | HYDR-006 | DRAIN: tundish | 1 | optional, dependent on machine requirements
DTUR-14 | Dirty Utility, 14m2 | SE | ELBO-010 | SWITCH: light, 2-way | 2
DTUR-14 | Dirty Utility, 14m2 | SE | ELGP-106 | GPO: single, special, for equipment | 2 | for washer sanitisers
DTUR-14 | Dirty Utility, 14m2 | SE | ELGP-109.1 | GPO: single, wall mounted (V) | 1 | optional, for urinalysis machine
DTUR-14 | Dirty Utility, 14m2 | SE | HYGE-028 | DEVICE: backflow prevention | 1
DTUR-S | Dirty Utility, Sub | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
DTUR-S | Dirty Utility, Sub | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1
DTUR-S | Dirty Utility, Sub | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
DTUR-S | Dirty Utility, Sub | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DTUR-S | Dirty Utility, Sub | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DTUR-S | Dirty Utility, Sub | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
DTUR-S | Dirty Utility, Sub | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | pull handle/push plate, closer
DTUR-S | Dirty Utility, Sub | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 5
DTUR-S | Dirty Utility, Sub | FA | CLCN-009 | CORNICE: wall trim fixing | 1
DTUR-S | Dirty Utility, Sub | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | with wet edge
DTUR-S | Dirty Utility, Sub | FFE | FIHR-038 | RACK: bowl/basin | 1 | quantity to be based on service requirements
DTUR-S | Dirty Utility, Sub | FFE | FIHR-016 | BRACKET: for sharps bin | 1
DTUR-S | Dirty Utility, Sub | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DTUR-S | Dirty Utility, Sub | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DTUR-S | Dirty Utility, Sub | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
DTUR-S | Dirty Utility, Sub | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DTUR-S | Dirty Utility, Sub | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DTUR-S | Dirty Utility, Sub | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DTUR-S | Dirty Utility, Sub | FFE | FIHR-029 | HOOK: equipment | 2 | optional, for clean linen bags
DTUR-S | Dirty Utility, Sub | FFE | FIHR-025 | HOOK: catheter bag | 2 | optional
DTUR-S | Dirty Utility, Sub | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 2
DTUR-S | Dirty Utility, Sub | FFE | FIHR-053 | RACK: pan & urinal | 1 | quantity to be based on service requirements
DTUR-S | Dirty Utility, Sub | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
DTUR-S | Dirty Utility, Sub | FFE | FIJO-196 | SHELF: stainless steel, overbench, wall mount | 1
DTUR-S | Dirty Utility, Sub | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1
DTUR-S | Dirty Utility, Sub | FFE | FQCW-006 | BIN: waste, clinical, 240L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
DTUR-S | Dirty Utility, Sub | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DTUR-S | Dirty Utility, Sub | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
DTUR-S | Dirty Utility, Sub | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
DTUR-S | Dirty Utility, Sub | FFE | MMIC-095 | WASHER SANITISER: pan flusher | 1 | combined utensil/pan sanitiser may be provided
DTUR-S | Dirty Utility, Sub | FFE | HYBA-052 | SINK: single, 450 x 350 x 250D | 1 | hot water
DTUR-S | Dirty Utility, Sub | FFE | MMLA-013 | ANALYSER: urinalysis | 1 | optional
DTUR-S | Dirty Utility, Sub | FFE | FIJO-100 | BENCH LEGS: stainless steel, for bench top | 1
DTUR-S | Dirty Utility, Sub | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional, for urinalysis machine
DTUR-S | Dirty Utility, Sub | SE | HYTP-013 | OUTLET: hot water | 1 | for washer sanitiser
DTUR-S | Dirty Utility, Sub | SE | HYTP-008 | OUTLET: cold water | 1 | for washer sanitiser
DTUR-S | Dirty Utility, Sub | SE | HYDR-004 | DRAIN: floor waste | 1
DTUR-S | Dirty Utility, Sub | SE | ELBO-015 | SWITCH: light, single | 1
DTUR-S | Dirty Utility, Sub | SE | ELGP-106 | GPO: single, special, for equipment | 1 | for washer sanitiser
DTUR-S | Dirty Utility, Sub | SE | ELGP-109.1 | GPO: single, wall mounted (V) | 1 | optional, for urinalysis machine
DTUR-S | Dirty Utility, Sub | SE | HYGE-028 | DEVICE: backflow prevention | 1
ECHO-TOE | Echocardiography - Transoesophageal | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
ECHO-TOE | Echocardiography - Transoesophageal | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
ECHO-TOE | Echocardiography - Transoesophageal | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
ECHO-TOE | Echocardiography - Transoesophageal | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
ECHO-TOE | Echocardiography - Transoesophageal | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
ECHO-TOE | Echocardiography - Transoesophageal | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
ECHO-TOE | Echocardiography - Transoesophageal | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | optional, to patient wc
ECHO-TOE | Echocardiography - Transoesophageal | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1
ECHO-TOE | Echocardiography - Transoesophageal | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
ECHO-TOE | Echocardiography - Transoesophageal | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 5
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MGFP-002 | FLOWMETER: medical air | 1 | patient services panel
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIHR-019 | BRACKET: for suction bottle | 2
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIHR-026 | HOOK: coat | 2 | optional
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIJO-131 | CUPBOARD: tall, single door, adj shelves, lockable | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIRT-025 | TRACK: curtain, bedscreen, U shaped | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | to entry door
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIDI-019 | DISPENSER: gown/apron | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIDI-014 | DISPENSER: disposable glove | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIDI-021 | DISPENSER: mask | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMGE-126 | TROLLEY: intubation | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | patient services panel
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMGE-088 | SAFE: dangerous drugs, small | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQBS-051 | STOOL: mobile, adj height | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQSN-006 | BIN: storage bins, for wall mounted panel | 3
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMGE-002 | ADAPTER: suction | 2 | patient services panel
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMGE-100 | SUCTION BOTTLE | 2
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQSM-016 | TROLLEY | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FQSM-003 | PEDESTAL: mobile, wire basket unit, 840H | 2
ECHO-TOE | Echocardiography - Transoesophageal | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
ECHO-TOE | Echocardiography - Transoesophageal | FFE | ITCL-014 | CLOCK: analogue, electric | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMHA-002 | DEFIBRILLATOR | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMGE-062 | MONITOR: patient, dynamap | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMMI-325 | DEVICE: imaging, echocardiography machine | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIJO-212 | BULKHEAD: joinery | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMGE-169 | SLING: arm, dynamic, ceiling mounted | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | MMBE-069 | TROLLEY: patient, echocardiography | 1
ECHO-TOE | Echocardiography - Transoesophageal | FFE | ITNE-071 | COMPUTER: desktop, with dual display screens | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 5
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | exam light
ECHO-TOE | Echocardiography - Transoesophageal | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELBO-015 | SWITCH: light, single | 2 | 1 optional, to optional clean up room
ECHO-TOE | Echocardiography - Transoesophageal | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
ECHO-TOE | Echocardiography - Transoesophageal | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
ECHO-TOE | Echocardiography - Transoesophageal | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
ECHO-TOE | Echocardiography - Transoesophageal | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELGP-109 | GPO: single, wall mounted | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELGP-208 | GPO: double, wall mounted | 5
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 3
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELBO-012 | SWITCH: light, dimmer | 1
ECHO-TOE | Echocardiography - Transoesophageal | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4 | on services panel
ECL-10 | Equipment Clean-Up, 10m2 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1
ECL-10 | Equipment Clean-Up, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
ECL-10 | Equipment Clean-Up, 10m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | Equal to Acrovyn
ECL-10 | Equipment Clean-Up, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
ECL-10 | Equipment Clean-Up, 10m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1
ECL-10 | Equipment Clean-Up, 10m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
ECL-10 | Equipment Clean-Up, 10m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | Equal to Acrovyn
ECL-10 | Equipment Clean-Up, 10m2 | FA | DWGL-020 | DOOR: 1100mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | 1100mm clear opening, lockable
ECL-10 | Equipment Clean-Up, 10m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | Wet edge, inset sink
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 2 | to basin and under bench
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1 | under bench
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1 | over bench
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FQSM-016 | TROLLEY | 2 | optional
ECL-10 | Equipment Clean-Up, 10m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | Refer to BHWS-B (similar)
ECL-10 | Equipment Clean-Up, 10m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 2 | to sinks
ECL-10 | Equipment Clean-Up, 10m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
ECL-10 | Equipment Clean-Up, 10m2 | FFE | HYTP-004 | HOSE: spray | 1 | optional
ECL-10 | Equipment Clean-Up, 10m2 | FFE | HYBA-055 | SINK: single, 500 x 450 x 250D | 2 | hot water
ECL-10 | Equipment Clean-Up, 10m2 | FFE | FIJO-124 | CUPBOARD: tall, triple door, adj shelves, lockable | 1 | optional
ECL-10 | Equipment Clean-Up, 10m2 | SE | HYDR-004 | DRAIN: floor waste | 1
ECL-10 | Equipment Clean-Up, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
ECL-10 | Equipment Clean-Up, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | Optional. To be decided at project level.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | DWGL-020 | DOOR: 1100mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | lockable
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | Wet edge and inset sink/s. Length to suit layout, refer RLS or project specific requirements.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 | Optional
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIDI-018 | DISPENSER: gown | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 1 | Optional. Number to suit layout.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIDI-008 | DISPENSER: detergent | 1 | To sink.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIDI-022 | DISPENSER: moisturiser | 0 | optional
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | MMGE-119 | TROLLEY: equipment | 1 | Optional.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIGE-009 | PLATFORM: equipment washing, wall mounted | 1 | Stainless steel, grated, height adjustable, wall mounted.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1 | At 350mm AFFL or underbench, stainless steel cupboard.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 2 | Over bench. Number to suit layout, refer RLS or project specific requirements.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | Refer to BHWS-B (similar). Elbow action levers or handsfree optional.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin. Elbow action levers or handsfree optional.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | HYTP-006 | HOSE: spray, pressure control | 1 | Above sink, optional. Depends on operational/infection control policies.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | HYTP-003 | HOSE REEL: retractable | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | FFE | HYBA-055 | SINK: single, 500 x 450 x 250D | 1 | hot water
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | SE | HYTP-013 | OUTLET: hot water | 2 | Hose reel and spray hose connection
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | SE | HYTP-008 | OUTLET: cold water | 2 | Hose reel and spray hose connection
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | SE | HYDR-004 | DRAIN: floor waste | 1 | To wash bay.
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | SE | ELBO-015 | SWITCH: light, single | 1
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
ECL-14 | Equipment Clean-Up, 14m2 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1
ECL-14 | Equipment Clean-Up, 14m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
ECL-14 | Equipment Clean-Up, 14m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | Equal to Acrovyn
ECL-14 | Equipment Clean-Up, 14m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
ECL-14 | Equipment Clean-Up, 14m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1
ECL-14 | Equipment Clean-Up, 14m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
ECL-14 | Equipment Clean-Up, 14m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | Equal to Acrovyn
ECL-14 | Equipment Clean-Up, 14m2 | FA | DWGL-020 | DOOR: 1100mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | lockable
ECL-14 | Equipment Clean-Up, 14m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | wet edge, inset sinks
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 2 | optional
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 2 | to basin and under bench
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1
ECL-14 | Equipment Clean-Up, 14m2 | FFE | MMGE-119 | TROLLEY: equipment | 3 | optional
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1 | under bench
ECL-14 | Equipment Clean-Up, 14m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1 | over bench
ECL-14 | Equipment Clean-Up, 14m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | Refer BHWS-B (similar). Elbow action levers or handsfree optional.
ECL-14 | Equipment Clean-Up, 14m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 2 | to sinks
ECL-14 | Equipment Clean-Up, 14m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin. Elbow action levers or handsfree optional.
ECL-14 | Equipment Clean-Up, 14m2 | FFE | HYTP-004 | HOSE: spray | 1 | optional
ECL-14 | Equipment Clean-Up, 14m2 | FFE | HYBA-055 | SINK: single, 500 x 450 x 250D | 2 | hot water
ECL-14 | Equipment Clean-Up, 14m2 | FFE | MMIC-097 | WASHER SANITISER: utensil washer | 1 | optional
ECL-14 | Equipment Clean-Up, 14m2 | SE | ELPR-003 | ISOLATOR: electrical, 1ph+n | 1
ECL-14 | Equipment Clean-Up, 14m2 | SE | HYTP-013 | OUTLET: hot water | 1
ECL-14 | Equipment Clean-Up, 14m2 | SE | HYTP-008 | OUTLET: cold water | 1
ECL-14 | Equipment Clean-Up, 14m2 | SE | HYDR-004 | DRAIN: floor waste | 1
ECL-14 | Equipment Clean-Up, 14m2 | SE | ELBO-015 | SWITCH: light, single | 1
ECL-14 | Equipment Clean-Up, 14m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
ECL-14 | Equipment Clean-Up, 14m2 | SE | ELGP-106 | GPO: single, special, for equipment | 1 | utensil washer, optional
ECL-8 | Equipment Clean-Up, 8m2 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1
ECL-8 | Equipment Clean-Up, 8m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
ECL-8 | Equipment Clean-Up, 8m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | Equal to Acrovyn
ECL-8 | Equipment Clean-Up, 8m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
ECL-8 | Equipment Clean-Up, 8m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1
ECL-8 | Equipment Clean-Up, 8m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
ECL-8 | Equipment Clean-Up, 8m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | Equal to Acrovyn
ECL-8 | Equipment Clean-Up, 8m2 | FA | DWGL-020 | DOOR: 1100mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | lockable
ECL-8 | Equipment Clean-Up, 8m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | wet edge, inset sink
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 1 | optional
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIJO-197 | SHELF: stainless steel, underbench | 1 | under bench
ECL-8 | Equipment Clean-Up, 8m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1 | over bench
ECL-8 | Equipment Clean-Up, 8m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | Refer BHWS-B (similar)
ECL-8 | Equipment Clean-Up, 8m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 2 | to sinks
ECL-8 | Equipment Clean-Up, 8m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
ECL-8 | Equipment Clean-Up, 8m2 | FFE | HYBA-055 | SINK: single, 500 x 450 x 250D | 2 | hot water
ECL-8 | Equipment Clean-Up, 8m2 | SE | HYDR-004 | DRAIN: floor waste | 1
ECL-8 | Equipment Clean-Up, 8m2 | SE | ELBO-015 | SWITCH: light, single | 1
ECL-8 | Equipment Clean-Up, 8m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
ENS-BR | Ensuite - Birthing | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | to be tiles for a tiled floor
ENS-BR | Ensuite - Birthing | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
ENS-BR | Ensuite - Birthing | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
ENS-BR | Ensuite - Birthing | FA | WLWA-019 | WALL FINISH: vinyl, full height | 1 | tiles are also acceptable
ENS-BR | Ensuite - Birthing | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1 | water resistant ceiling tiles are also acceptable
ENS-BR | Ensuite - Birthing | FA | DWSC-011 | DOOR: 1100mm c/o, single leaf, solid core, paint | 1 | outward opening, with privacy latch and emergency release
ENS-BR | Ensuite - Birthing | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
ENS-BR | Ensuite - Birthing | FFE | FIHR-026 | HOOK: coat | 2 | optional
ENS-BR | Ensuite - Birthing | FFE | FIRT-034 | TRACK: curtain, shower, 90 deg | 1 | optional, dependent on project/service requirements
ENS-BR | Ensuite - Birthing | FFE | FQGE-011 | CURTAIN: shower | 1 | optional, dependent on project/service requirements
ENS-BR | Ensuite - Birthing | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
ENS-BR | Ensuite - Birthing | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
ENS-BR | Ensuite - Birthing | FFE | FIDI-030 | DISPENSER: soap | 1
ENS-BR | Ensuite - Birthing | FFE | FIDI-036 | DISPENSER: toilet paper | 1
ENS-BR | Ensuite - Birthing | FFE | FIRT-006 | GRAB RAIL: WC | 1
ENS-BR | Ensuite - Birthing | FFE | FIRT-003 | GRAB RAIL: shower, horizontal | 2
ENS-BR | Ensuite - Birthing | FFE | FIBM-006 | MIRROR: fixed, glass | 1 | over basin
ENS-BR | Ensuite - Birthing | FFE | MGFP-004 | FLOWMETER: oxygen | 1
ENS-BR | Ensuite - Birthing | FFE | FQCW-037 | BIN: waste, sanitary | 1
ENS-BR | Ensuite - Birthing | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 1
ENS-BR | Ensuite - Birthing | FFE | FIRT-016 | RAIL: towel | 1 | strength and fixings, including in-wall support, to be equivalent to grab rails
ENS-BR | Ensuite - Birthing | FFE | HYWC-001 | TOILET SUITE: accessible | 1 | accessible suite to assist with mobility and support during labor
ENS-BR | Ensuite - Birthing | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1
ENS-BR | Ensuite - Birthing | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 2
ENS-BR | Ensuite - Birthing | FFE | HYTP-020 | SHOWER: hand-held | 2
ENS-BR | Ensuite - Birthing | FFE | HYBA-002 | BASIN: ensuite, with vanity unit | 1
ENS-BR | Ensuite - Birthing | FFE | FIGE-011 | BACKREST: accessible toilet | 1
ENS-BR | Ensuite - Birthing | FFE | MMMA-019 | SUPPORT FIXTURE: birthing, ladder, wall mounted, wet area | 1 | optional; strength and fixings, including in-wall support, to be equivalent to heavy duty grab rails
ENS-BR | Ensuite - Birthing | SE | ITCL-008 | BUTTON: nurse call, staff assist | 2 | with indicator button and light, IPX4 rating required
ENS-BR | Ensuite - Birthing | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1
ENS-BR | Ensuite - Birthing | SE | HYDR-004 | DRAIN: floor waste | 2
ENS-BR | Ensuite - Birthing | SE | ELBO-015 | SWITCH: light, single | 1 | IPX4 rating required, located outside room
ENS-BR | Ensuite - Birthing | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
ENS-BR | Ensuite - Birthing | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1
ENS-BR | Ensuite - Birthing | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | with indicator button and light, IPX4 rating required
ENS-BR | Ensuite - Birthing | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light, IPX4 rating required
ENS-BR | Ensuite - Birthing | SE | ELGP-208 | GPO: double, wall mounted | 1 | to be classified as a shaver outlet
ENS-BR | Ensuite - Birthing | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
GYAH-GP | Gymnasium, Group Therapy | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
GYAH-GP | Gymnasium, Group Therapy | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
GYAH-GP | Gymnasium, Group Therapy | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
GYAH-GP | Gymnasium, Group Therapy | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
GYAH-GP | Gymnasium, Group Therapy | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 2
GYAH-GP | Gymnasium, Group Therapy | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
GYAH-GP | Gymnasium, Group Therapy | FA | DWSC-027 | DOOR: 1600mm c/o, 1 1/2 leaf, solid core, paint | 1 | optional, to equipment store
GYAH-GP | Gymnasium, Group Therapy | FA | DWGL-041 | DOOR: 1500mm c/o, 1 1/2 leaf, half glazed, solid core/glass, paint | 1
GYAH-GP | Gymnasium, Group Therapy | FA | CLCN-009 | CORNICE: wall trim fixing | 1
GYAH-GP | Gymnasium, Group Therapy | FA | FLVY-010 | FLOOR FINISH: vinyl, seamless, inset floor markings | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | MMAH-016 | ERGOMETER: arm | 2
GYAH-GP | Gymnasium, Group Therapy | FFE | FQBS-030 | CHAIR: visitor, patient | 2 | to arm ergometer
GYAH-GP | Gymnasium, Group Therapy | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2
GYAH-GP | Gymnasium, Group Therapy | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | FQGE-010 | CURTAIN: privacy screen | 2 | to plinth
GYAH-GP | Gymnasium, Group Therapy | FFE | FQSN-007 | BIN: storage bins, wire mesh, stainless steel | 1 | to plinth
GYAH-GP | Gymnasium, Group Therapy | FFE | FIHR-029 | HOOK: equipment | 2 | to plinth
GYAH-GP | Gymnasium, Group Therapy | FFE | MMAH-002 | BICYCLE: exercise | 2
GYAH-GP | Gymnasium, Group Therapy | FFE | MMAH-003 | BICYCLE: exercise, recumbent | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1 | to plinth
GYAH-GP | Gymnasium, Group Therapy | FFE | FIRT-017 | STEEL MESH: framed, ceiling mount | 1 | optional, to plinth
GYAH-GP | Gymnasium, Group Therapy | FFE | MMAH-035 | PLINTH: single width | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | FIHR-061 | RACK: weights | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | MMAH-019 | MACHINE: rowing | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | MMAH-020 | MACHINE: stepping | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | MMAH-027 | TREADMILL: automatic | 2
GYAH-GP | Gymnasium, Group Therapy | FFE | FQDW-028 | WORKSTATION: straight, adj height, 1200W x 750D x 750H nom. | 2 | for arm ergometer
GYAH-GP | Gymnasium, Group Therapy | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
GYAH-GP | Gymnasium, Group Therapy | FFE | FIJO-225 | CUPBOARD: tall, pigeon hole shelf unit, gymnasium | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | FIJO-227 | CUPBOARD: tall, double door, gym equipment storage | 1
GYAH-GP | Gymnasium, Group Therapy | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 1
GYAH-GP | Gymnasium, Group Therapy | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
GYAH-GP | Gymnasium, Group Therapy | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to plinth
GYAH-GP | Gymnasium, Group Therapy | SE | ITCL-008 | BUTTON: nurse call, staff assist | 2 | 1 x to plinth, 1 x at entry door
GYAH-GP | Gymnasium, Group Therapy | SE | ELBO-015 | SWITCH: light, single | 1
GYAH-GP | Gymnasium, Group Therapy | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | optional, to plinth
GYAH-GP | Gymnasium, Group Therapy | SE | ITCL-004 | BUTTON: nurse call, emergency | 2 | 1 x to plinth, 1 x at entry door
GYAH-GP | Gymnasium, Group Therapy | SE | ELGP-109 | GPO: single, wall mounted | 1 | to plinth
GYAH-GP | Gymnasium, Group Therapy | SE | ELGP-208 | GPO: double, wall mounted | 9 | quantity to suit equipment requirements
GYAH-GP | Gymnasium, Group Therapy | SE | ELGP-208 | GPO: double, wall mounted | 3 | to plinth
GYAH-GP | Gymnasium, Group Therapy | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
GYAH-ID | Gymnasium, Individual Treatment | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
GYAH-ID | Gymnasium, Individual Treatment | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
GYAH-ID | Gymnasium, Individual Treatment | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
GYAH-ID | Gymnasium, Individual Treatment | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 2
GYAH-ID | Gymnasium, Individual Treatment | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1
GYAH-ID | Gymnasium, Individual Treatment | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
GYAH-ID | Gymnasium, Individual Treatment | FA | DWSC-027 | DOOR: 1600mm c/o, 1 1/2 leaf, solid core, paint | 1 | optional, to equipment store
GYAH-ID | Gymnasium, Individual Treatment | FA | DWGL-041 | DOOR: 1500mm c/o, 1 1/2 leaf, half glazed, solid core/glass, paint | 1
GYAH-ID | Gymnasium, Individual Treatment | FA | CLCN-009 | CORNICE: wall trim fixing | 1
GYAH-ID | Gymnasium, Individual Treatment | FA | FLVY-010 | FLOOR FINISH: vinyl, seamless, inset floor markings | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIBM-010 | MIRROR: fixed, safety glass, full length | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 4
GYAH-ID | Gymnasium, Individual Treatment | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 3
GYAH-ID | Gymnasium, Individual Treatment | FFE | FQGE-010 | CURTAIN: privacy screen | 3 | to plinths
GYAH-ID | Gymnasium, Individual Treatment | FFE | FQSN-007 | BIN: storage bins, wire mesh, stainless steel | 3 | to plinths
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIHR-029 | HOOK: equipment | 6 | to plinths
GYAH-ID | Gymnasium, Individual Treatment | FFE | MMAH-002 | BICYCLE: exercise | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 3 | to plinths
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIRT-017 | STEEL MESH: framed, ceiling mount | 3 | optional, to plinths
GYAH-ID | Gymnasium, Individual Treatment | FFE | MMAH-001 | BARS: parallel | 1 | bariatric capacity to be determine based on service requirements
GYAH-ID | Gymnasium, Individual Treatment | FFE | MMAH-033 | PLINTH: double width | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | MMAH-035 | PLINTH: single width | 2
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIHR-061 | RACK: weights | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 3
GYAH-ID | Gymnasium, Individual Treatment | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | MMAH-027 | TREADMILL: automatic | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 3
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
GYAH-ID | Gymnasium, Individual Treatment | FFE | FIJO-227 | CUPBOARD: tall, double door, gym equipment storage | 1
GYAH-ID | Gymnasium, Individual Treatment | FFE | MMAH-044 | STAIRS: exercise, 90 degree | 1
GYAH-ID | Gymnasium, Individual Treatment | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
GYAH-ID | Gymnasium, Individual Treatment | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | to plinths
GYAH-ID | Gymnasium, Individual Treatment | SE | ITCL-008 | BUTTON: nurse call, staff assist | 4 | 3 x to plinth, 1 x at entry door
GYAH-ID | Gymnasium, Individual Treatment | SE | ELBO-015 | SWITCH: light, single | 1
GYAH-ID | Gymnasium, Individual Treatment | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 3 | optional, to plinths
GYAH-ID | Gymnasium, Individual Treatment | SE | ITCL-004 | BUTTON: nurse call, emergency | 4 | 3 x to plinth, 1 x at entry door
GYAH-ID | Gymnasium, Individual Treatment | SE | ELGP-109 | GPO: single, wall mounted | 3 | to plinths
GYAH-ID | Gymnasium, Individual Treatment | SE | ELGP-208 | GPO: double, wall mounted | 3 | quantity to suit equipment requirements
GYAH-ID | Gymnasium, Individual Treatment | SE | ELGP-208 | GPO: double, wall mounted | 9 | to plinths
GYAH-ID | Gymnasium, Individual Treatment | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
HYDP | Hydrotherapy Pool | FA | FLGE-010 | FLOOR FINISH: tiles, ceramic, non-slip | 2
HYDP | Hydrotherapy Pool | FA | FLSK-004 | SKIRTING: tiles, coved, to match floor tiles | 12
HYDP | Hydrotherapy Pool | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 6
HYDP | Hydrotherapy Pool | FA | WLWA-011 | WALL FINISH: tiles, ceramic, glazed | 3 | to pool and all walls
HYDP | Hydrotherapy Pool | FA | CLFS-002 | CEILING: fibre cement, flush set, suspended, paint, washable | 1
HYDP | Hydrotherapy Pool | FA | DWWH-005 | WINDOW | 1 | desirable
HYDP | Hydrotherapy Pool | FA | DWOB-007 | OBSERVATION: internal window, alum/safety glass, powdercoat/paint | 4
HYDP | Hydrotherapy Pool | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 6 | to change rooms, cleaner's room, plant areas and alternate exit if required by BCA; lockable
HYDP | Hydrotherapy Pool | FA | DWGL-040 | DOOR: 1500mm c/o ,1 1/2 leaf, fully glazed, alum/safety glass, powdercoat/paint | 2 | to lobby/ waiting; lockable, restricted access
HYDP | Hydrotherapy Pool | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
HYDP | Hydrotherapy Pool | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
HYDP | Hydrotherapy Pool | FFE | MMAH-017 | HOIST: pool, hydraulic | 1 | weight limit requirement to be confirmed dependent on service requirements
HYDP | Hydrotherapy Pool | FFE | MMAH-032 | BLANKET: pool, auto-operated | 1 | optional, power requirements dependent on selected model
HYDP | Hydrotherapy Pool | FFE | ITBU-009 | CONTROL UNIT: public address system | 1 | for announcements and music
HYDP | Hydrotherapy Pool | SE | ITBU-020 | PUBLIC ADDRESS SYSTEM | 1
HYDP | Hydrotherapy Pool | SE | HYTP-008 | OUTLET: cold water | 1 | for washing down concourse
HYDP | Hydrotherapy Pool | SE | HYDR-004 | DRAIN: floor waste | 16 | every 3-4 metres on concourse
HYDP | Hydrotherapy Pool | SE | ELBO-015 | SWITCH: light, single | 1
HYDP | Hydrotherapy Pool | SE | ITCL-004 | BUTTON: nurse call, emergency | 4 | 1 each side of pool
HYDP | Hydrotherapy Pool | SE | ELGP-106 | GPO: single, special, for equipment | 1 | optional, for motorised pool blanket
HYDP | Hydrotherapy Pool | SE | ELGP-113 | GPO: single, wall mounted, waterproof, RCD protected | 4
INTV | Interview Room | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
INTV | Interview Room | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 6
INTV | Interview Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
INTV | Interview Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
INTV | Interview Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
INTV | Interview Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
INTV | Interview Room | FA | DWWH-005 | WINDOW | 1
INTV | Interview Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 2 | second door is optional for egress
INTV | Interview Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
INTV | Interview Room | FFE | FQGE-002 | ARTWORK | 1 | optional
INTV | Interview Room | FFE | DWWD-004 | BLIND: roller | 1
INTV | Interview Room | FFE | FQBS-010 | CHAIR: armchair | 1
INTV | Interview Room | FFE | FQBS-041 | LOUNGE: 2.5 seat | 1
INTV | Interview Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
INTV | Interview Room | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1
INTV | Interview Room | FFE | FQTA-075 | TABLE: square, 600W, low height | 2
INTV | Interview Room | FFE | FQBS-058 | CHAIR: visitor, patient, with arms | 4
INTV | Interview Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2
INTV | Interview Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | optional, + indicator button & light
INTV | Interview Room | SE | ELBO-015 | SWITCH: light, single | 1
INTV | Interview Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | optional, + indicator button & light
INTV | Interview Room | SE | ELGP-208 | GPO: double, wall mounted | 2
INTV | Interview Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
INTV | Interview Room | SE | ITSE-038 | READER: access control, general | 1 | optional to second egress door
LAUN-PT | Laundry - Patient, 6m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
LAUN-PT | Laundry - Patient, 6m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | To be decided at project level.
LAUN-PT | Laundry - Patient, 6m2 | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1
LAUN-PT | Laundry - Patient, 6m2 | FA | WLWA-007 | WALL FINISH: splashback | 1 | Selection of material to comply with Part D Infection Prevention and Control AusHFG.
LAUN-PT | Laundry - Patient, 6m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
LAUN-PT | Laundry - Patient, 6m2 | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
LAUN-PT | Laundry - Patient, 6m2 | FA | DWOB-002 | OBSERVATION: internal window, alum/glass, door height, powdercoat | 1 | Observation window, optional
LAUN-PT | Laundry - Patient, 6m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
LAUN-PT | Laundry - Patient, 6m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | lockable
LAUN-PT | Laundry - Patient, 6m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
LAUN-PT | Laundry - Patient, 6m2 | FFE | FIHR-007 | BRACKET: for dryer | 1 | for dryer
LAUN-PT | Laundry - Patient, 6m2 | FFE | FQSN-023 | CABINET: ironing board | 1 | lockable, with integral ironing board
LAUN-PT | Laundry - Patient, 6m2 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
LAUN-PT | Laundry - Patient, 6m2 | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 2 | Over bench. Number to suit layout, refer RLS or project specific requirements. Some lockable space may be needed.
LAUN-PT | Laundry - Patient, 6m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
LAUN-PT | Laundry - Patient, 6m2 | FFE | FQSM-030 | TROLLEY: laundry, basket | 1 | under bench
LAUN-PT | Laundry - Patient, 6m2 | FFE | FQGE-015 | DRYER: clothes, domestic | 1 | wall mounted @ 600mm AFFL on bracket
LAUN-PT | Laundry - Patient, 6m2 | FFE | FQGE-028 | IRON: clothes, electric | 1 | inside cupboard, hardwired.
LAUN-PT | Laundry - Patient, 6m2 | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
LAUN-PT | Laundry - Patient, 6m2 | FFE | HYBA-064 | SINK: single, laundry | 1 | with overflow drain, warm water, sinlge mixer to tub optional
LAUN-PT | Laundry - Patient, 6m2 | FFE | FQCW-056 | MACHINE: washing, domestic | 1 | heavy duty
LAUN-PT | Laundry - Patient, 6m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1 | With inset laundry tub, open under part. Length to suit layout, refer RLS or project specific requirements.
LAUN-PT | Laundry - Patient, 6m2 | SE | HYTP-013 | OUTLET: hot water | 1 | to washing machine
LAUN-PT | Laundry - Patient, 6m2 | SE | HYTP-008 | OUTLET: cold water | 1
LAUN-PT | Laundry - Patient, 6m2 | SE | HYDR-004 | DRAIN: floor waste | 1
LAUN-PT | Laundry - Patient, 6m2 | SE | MEEX-002 | EXHAUST: dryer | 1
LAUN-PT | Laundry - Patient, 6m2 | SE | ELBO-015 | SWITCH: light, single | 1
LAUN-PT | Laundry - Patient, 6m2 | SE | ELGP-209 | GPO: double, wall mounted, RCD protected | 1 | 30mA RCD protection, above bench
LAUN-PT | Laundry - Patient, 6m2 | SE | ELGP-111 | GPO: single, wall mounted, RCD protected | 3 | 30mA RCD protection, washer, dryer, ironing cabinet,
LINAC | Linear Accelerator Treatment Room | FA | FLVY-001 | FLOOR FINISH: vinyl, seamless, coved, anti-static | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
LINAC | Linear Accelerator Treatment Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 18
LINAC | Linear Accelerator Treatment Room | FA | WLPR-006 | WALL PROTECTION: corner guards | 9
LINAC | Linear Accelerator Treatment Room | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 14
LINAC | Linear Accelerator Treatment Room | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 16
LINAC | Linear Accelerator Treatment Room | FA | CLFS-005 | CEILING: plasterboard, flush set, suspended, paint, scrub resistant | 1
LINAC | Linear Accelerator Treatment Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | inset area above machine as requried
LINAC | Linear Accelerator Treatment Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
LINAC | Linear Accelerator Treatment Room | FFE | DWGE-013 | GATE: single leaf, steel, prefinished | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-016 | BRACKET: for sharps bin | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-006 | DISPENSER: cleaning wipes | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 3
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-140 | CUPBOARD: underbench, single door, adj shelves | 1
LINAC | Linear Accelerator Treatment Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-110 | TROLLEY: clean linen, small | 1
LINAC | Linear Accelerator Treatment Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1
LINAC | Linear Accelerator Treatment Room | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-002 | ADAPTER: suction | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
LINAC | Linear Accelerator Treatment Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
LINAC | Linear Accelerator Treatment Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITCL-012 | CAMERA: observation | 6 | 1 x optional, for total body irradiation or visualisation of additional monitoring equipment, pan, tilt and zoom (PTZ) capability is required, quantity to be based on equipment and service requirements
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | workstation on wheels is also acceptable
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-004 | LINEAR ACCELERATOR | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 2 | for patient's own music, individual to each bunker from control room, hearing loop required to speakers
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-058 | BENCH TOP: 600D, joinery under | 3
LINAC | Linear Accelerator Treatment Room | FFE | ITAV-010 | MICROPHONE: ceiling mounted | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-165 | CUPBOARD: tall, applicator storage, radiotherapy | 1 | custom joinery is required for vendor specific electron cones, including sliding rails, shelving to be reinforced for electron shields
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-166 | CUPBOARD: tall, shelving, horizontal patient device storage, radiotherapy | 3 | wide span shelving to be reinforced
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-167 | CUPBOARD: tall, shelving, vertical patient device storage, radiotherapy | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead is also acceptable
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-066 | COMPUTER: all-in-one, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | to wall mounted computer, may be wireless
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 6 | 1 x optional, to additional patient observation camera if provided
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-073 | DISPLAY SCREEN: patient treatment information | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-006 | ACCESS PANEL: medical physics port | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-013 | DISPLAY SCREEN: radiotherapy, ceiling mounted | 2
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-015 | LASER: radiotherapy, ceiling mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-016 | LASER: radiotherapy, lateral laser | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-231 | CUPBOARD: radiotherapy, laser enclosure | 3
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-017 | LASER: radiotherapy, sagittal laser | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-008 | DEVICE: surface guided radiation therapy system | 3 | as required for respiratory gating / motion management / surface guidance systems, to suit service requirements
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-011 | DEVICE: radiotherapy, daily QA | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-012 | GATE: light gate, radiotherapy bunker | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-093 | BRACKET: keyboard shelf, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-089 | BRACKET: surface guided radiation therapy system, ceiling mounted | 3
LINAC | Linear Accelerator Treatment Room | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 11 | 1 x optional, to additional patient observation camera if provided
LINAC | Linear Accelerator Treatment Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3
LINAC | Linear Accelerator Treatment Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | SE | ELPR-006 | RCD: residual current device | 5
LINAC | Linear Accelerator Treatment Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
LINAC | Linear Accelerator Treatment Room | SE | ELBO-001 | BUTTON: emergency stop | 4
LINAC | Linear Accelerator Treatment Room | SE | LIFX-004 | LIGHT: illuminated sign, "Beam On-Beam Off" | 1
LINAC | Linear Accelerator Treatment Room | SE | ELBO-010 | SWITCH: light, 2-way | 2
LINAC | Linear Accelerator Treatment Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
LINAC | Linear Accelerator Treatment Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1
LINAC | Linear Accelerator Treatment Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
LINAC | Linear Accelerator Treatment Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 4
LINAC | Linear Accelerator Treatment Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 2
LINAC | Linear Accelerator Treatment Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
LINAC | Linear Accelerator Treatment Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
LINAC | Linear Accelerator Treatment Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
LINAC | Linear Accelerator Treatment Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3
LINAC | Linear Accelerator Treatment Room | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 12 | 1 x to ceiling mounted laser, 1 x optional, to additional patient observation camera if provided
LINAC | Linear Accelerator Treatment Room | SE | ELGP-220.1 | GPO: emergency power, double (V) | 5 | to wall mounted lasers
LINAC | Linear Accelerator Treatment Room | SE | ELBO-019 | BUTTON: last person out | 1
LINAC | Linear Accelerator Treatment Room | SE | LIFX-038 | LIGHT: backlit feature panels | 2
LINAC-CR | Linear Accelerator Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
LINAC-CR | Linear Accelerator Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
LINAC-CR | Linear Accelerator Control Room | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1 | optional
LINAC-CR | Linear Accelerator Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
LINAC-CR | Linear Accelerator Control Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
LINAC-CR | Linear Accelerator Control Room | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | optional
LINAC-CR | Linear Accelerator Control Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
LINAC-CR | Linear Accelerator Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 5
LINAC-CR | Linear Accelerator Control Room | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 5
LINAC-CR | Linear Accelerator Control Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
LINAC-CR | Linear Accelerator Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1 | may be bench mounted, or integrated with microphone
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
LINAC-CR | Linear Accelerator Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 1
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 2
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-099 | CAC: cable access cap | 8
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead is also acceptable
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-213 | PANEL: joinery carcass end | 2
LINAC-CR | Linear Accelerator Control Room | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 2 | to patient observation monitors
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | for patient observation cameras to control pan, tilt and zoom; may be custom control depending on camera system
LINAC-CR | Linear Accelerator Control Room | FFE | FIHR-080 | BRACKET: for bench, 750D, wall mounted | 5 | legs are also acceptable, quantity to be confirmed to achieve required support, ensure location does not impact leg clearance at workstations
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-071 | COMPUTER: desktop, with dual display screens | 2 | for access to local network including patient records and images for planning and reporting
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-006 | ACCESS PANEL: medical physics port | 1
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-009 | COMPUTER: linear accelerator control system PC | 6 | quantity to be confirmed based on vendor and equipment requirements; current provision allows for 1 x linear accelerator control system, 1 x record and verify system, and 2-4 x imaging and patient set up systems (e.g. gating and patient positioning).
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-010 | CONSOLE: control, linear accelerator | 1
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-007 | CABINET: linear accelerator computer cabinet | 1 | size, configuration and location to be confirmed by vendor
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-074 | DISPLAY SCREEN: patient observation | 2
LINAC-CR | Linear Accelerator Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3
LINAC-CR | Linear Accelerator Control Room | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 8 | quantity to suit computer and display screen requirements
LINAC-CR | Linear Accelerator Control Room | SE | ELPR-006 | RCD: residual current device | 3
LINAC-CR | Linear Accelerator Control Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
LINAC-CR | Linear Accelerator Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1 | may be integral to linear accelerator control console
LINAC-CR | Linear Accelerator Control Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
LINAC-CR | Linear Accelerator Control Room | SE | ELGP-208 | GPO: double, wall mounted | 10
LINAC-CR | Linear Accelerator Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 14 | quantity to suit computer and display screen requirements
LINAC-CR | Linear Accelerator Control Room | SE | ELGP-106 | GPO: single, special, for equipment | 1 | to linear accelerator computer cabinet
LINAC-CR | Linear Accelerator Control Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
LINAC-CR | Linear Accelerator Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
LINAC-CR | Linear Accelerator Control Room | SE | ITIN-055 | OUTLET: data, USB, wall mounted | 1 | connection for music to bunker
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | carpet tiles are also acceptable
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | to be feather edge vinyl for carpet tiles
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | DWWH-005 | WINDOW | 1 | natural light and external outlook is essential
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | DWGL-002 | DOOR: 910mm c/o, single leaf, half glazed, solid core/glass, paint | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FFE | FQGE-002 | ARTWORK | 2
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FFE | MMGE-019 | CHAIR: patient | 2 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FFE | DWWD-004 | BLIND: roller | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FFE | FQBS-043 | LOUNGE: 2 seat, 1600 wide | 1 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | FFE | FQTA-075 | TABLE: square, 600W, low height | 2 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to television, MATV outlet may also be provided depending on local ICT policies
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | with indicator button and light
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | optional, dependent on patient cohort, provide in areas requiring high staff observation; with indicator button and light
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | optional, dependent on patient cohort, provide in areas requiring high staff observation; with indicator button and light
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 3 | 1 to television, 2 to room
LNGE-10 | Lounge - Patient / Visitor, 10m2 | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | carpet tiles are also acceptable
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 | to be feather edge vinyl for carpet tiles
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | DWGL-061 | DOOR: 1800mm c/o, double leaf, half glazed, solid core/glass, paint | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | DWWH-005 | WINDOW | 1 | natural light and external outlook is essential
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | FQGE-002 | ARTWORK | 2
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | FQBS-039 | LOUNGE: 1 seat | 2 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | MMGE-019 | CHAIR: patient | 2 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | DWWD-004 | BLIND: roller | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | FQBS-043 | LOUNGE: 2 seat, 1600 wide | 2 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | FIHR-045 | RACK: magazine/pamphlet | 1 | optional
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | FFE | FQTA-075 | TABLE: square, 600W, low height | 2 | furniture is indicative and selection will be dependent on patient/visitor cohort and service requirements
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to television, MATV outlet may also be provided depending on local ICT policies
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | with indicator button and light
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ELBO-015 | SWITCH: light, single | 1
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | optional, dependent on patient cohort, provide in areas requiring high staff observation; with indicator button and light
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | optional, dependent on patient cohort, provide in areas requiring high staff observation; with indicator button and light
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 3 | 1 to television, 2 to room
LNGE-20 | Lounge - Patient / Visitor, 20m2 | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
MAMMO | Mammography Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
MAMMO | Mammography Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
MAMMO | Mammography Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
MAMMO | Mammography Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
MAMMO | Mammography Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MAMMO | Mammography Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin
MAMMO | Mammography Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | tiled ceiling also acceptable if interventional procedures are not undertaken
MAMMO | Mammography Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
MAMMO | Mammography Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MAMMO | Mammography Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 6
MAMMO | Mammography Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
MAMMO | Mammography Room | FFE | FIHR-005 | BRACKET: for display screen, PACS, wall mounted | 1
MAMMO | Mammography Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 1 | to desk
MAMMO | Mammography Room | FFE | MMGE-019 | CHAIR: patient | 1
MAMMO | Mammography Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
MAMMO | Mammography Room | FFE | FIHR-026 | HOOK: coat | 2 | optional, for patient changing, 1 @1200H
MAMMO | Mammography Room | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | to entry door
MAMMO | Mammography Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
MAMMO | Mammography Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
MAMMO | Mammography Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
MAMMO | Mammography Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
MAMMO | Mammography Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
MAMMO | Mammography Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
MAMMO | Mammography Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1
MAMMO | Mammography Room | FFE | FIHR-055 | RACK: scanhead | 1
MAMMO | Mammography Room | FFE | MMGE-002 | ADAPTER: suction | 1
MAMMO | Mammography Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
MAMMO | Mammography Room | FFE | FQSM-003 | PEDESTAL: mobile, wire basket unit, 840H | 1
MAMMO | Mammography Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
MAMMO | Mammography Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
MAMMO | Mammography Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
MAMMO | Mammography Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | to desk
MAMMO | Mammography Room | FFE | MMMI-159 | DEVICE: imaging, mammography | 1
MAMMO | Mammography Room | FFE | FQDW-028 | WORKSTATION: straight, adj height, 1200W x 750D x 750H nom. | 1
MAMMO | Mammography Room | FFE | MMMI-162 | CONSOLE: control, mammography | 1
MAMMO | Mammography Room | FFE | FIJO-099 | CAC: cable access cap | 1
MAMMO | Mammography Room | FFE | FIRT-040 | TRACK: curtain, privacy, L shaped, angled | 1
MAMMO | Mammography Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
MAMMO | Mammography Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 1
MAMMO | Mammography Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
MAMMO | Mammography Room | FFE | ITNE-070 | DISPLAY SCREEN: PACS | 2
MAMMO | Mammography Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3
MAMMO | Mammography Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to desk, 1 outlet for telephone
MAMMO | Mammography Room | SE | ELPR-006 | RCD: residual current device | 2 | to services panel
MAMMO | Mammography Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on service panel
MAMMO | Mammography Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1 | auto on when modality is in use
MAMMO | Mammography Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
MAMMO | Mammography Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1
MAMMO | Mammography Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
MAMMO | Mammography Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | x-ray viewer
MAMMO | Mammography Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | clock
MAMMO | Mammography Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2
MAMMO | Mammography Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
MAMMO | Mammography Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 3 | imaging and control equipment
MAMMO | Mammography Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
MAMMO | Mammography Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
MAMMO | Mammography Room | SE | WLSH-004 | SHIELDING: lead lining | 1
MAMMO | Mammography Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
MAMMO | Mammography Room | SE | ELGP-105 | GPO: single, on services panel | 2
MAMMO | Mammography Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2
MED-14 | Medication Room, 14m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
MED-14 | Medication Room, 14m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
MED-14 | Medication Room, 14m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
MED-14 | Medication Room, 14m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
MED-14 | Medication Room, 14m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MED-14 | Medication Room, 14m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 4 | extent dependent on type and height storage solution
MED-14 | Medication Room, 14m2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
MED-14 | Medication Room, 14m2 | FA | DWGL-003 | DOOR: 910mm c/o, single leaf, half glazed, solid core/safety glass, paint | 1
MED-14 | Medication Room, 14m2 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MED-14 | Medication Room, 14m2 | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 1 | extent dependent on type and height storage solution
MED-14 | Medication Room, 14m2 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
MED-14 | Medication Room, 14m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
MED-14 | Medication Room, 14m2 | FFE | FIJO-157 | CUPBOARD: wall mount, single door, adj shelves, lockable | 1
MED-14 | Medication Room, 14m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
MED-14 | Medication Room, 14m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
MED-14 | Medication Room, 14m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
MED-14 | Medication Room, 14m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
MED-14 | Medication Room, 14m2 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
MED-14 | Medication Room, 14m2 | FFE | MMGE-087 | SAFE: dangerous drugs, medium | 1 | size and quantity to be based on service requirements
MED-14 | Medication Room, 14m2 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
MED-14 | Medication Room, 14m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
MED-14 | Medication Room, 14m2 | FFE | FIJO-108 | CUPBOARD: drawer unit, baskets, 3 x Wire (400mm wide) | 5
MED-14 | Medication Room, 14m2 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
MED-14 | Medication Room, 14m2 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
MED-14 | Medication Room, 14m2 | FFE | ITSE-012 | CAMERA: CCTV | 1 | optional, provision and placement dependent on local policies
MED-14 | Medication Room, 14m2 | FFE | MMGE-034 | FRIDGE: drugs, large | 1 | size and quantity to be based on service requirements
MED-14 | Medication Room, 14m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
MED-14 | Medication Room, 14m2 | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
MED-14 | Medication Room, 14m2 | FFE | FQSM-046 | SHELVING: storage solution, 600W x 400D x 2100H nom., wire, mobile | 3 | type and quantity of storage options to be determined based on service requirements
MED-14 | Medication Room, 14m2 | FFE | FQCW-061 | BIN: waste, pharmaceutical | 1
MED-14 | Medication Room, 14m2 | FFE | FQSM-048 | SHELVING: storage solution, shelving, 600W x 400D x 2100H nom., wire, mobile | 2 | type and quantity of storage options to be determined based on service requirements
MED-14 | Medication Room, 14m2 | FFE | FIDI-052 | DISPENSER: sticker roll | 1 | optional
MED-14 | Medication Room, 14m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to drug fridge, for temperature monitoring
MED-14 | Medication Room, 14m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | 2x optional, for future provision of automated dispensing cabinets
MED-14 | Medication Room, 14m2 | SE | LIFX-008 | LIGHT: task light, built in | 1 | under wall mounted joinery
MED-14 | Medication Room, 14m2 | SE | ELBO-015 | SWITCH: light, single | 2 | 1 x general room lighting, 1 x task light
MED-14 | Medication Room, 14m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
MED-14 | Medication Room, 14m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | optional, for future provision of automated dispensing cabinets
MED-14 | Medication Room, 14m2 | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to drug fridge
MED-14 | Medication Room, 14m2 | SE | ITSE-038 | READER: access control, general | 1
MEET-15 | Meeting Room, 15m2 | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
MEET-15 | Meeting Room, 15m2 | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
MEET-15 | Meeting Room, 15m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
MEET-15 | Meeting Room, 15m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MEET-15 | Meeting Room, 15m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
MEET-15 | Meeting Room, 15m2 | FA | DWWH-005 | WINDOW | 1 | desirable
MEET-15 | Meeting Room, 15m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional lockable
MEET-15 | Meeting Room, 15m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
MEET-15 | Meeting Room, 15m2 | FFE | DWWD-004 | BLIND: roller | 1
MEET-15 | Meeting Room, 15m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
MEET-15 | Meeting Room, 15m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
MEET-15 | Meeting Room, 15m2 | FFE | FIJO-194 | SHELF: laminate | 1
MEET-15 | Meeting Room, 15m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
MEET-15 | Meeting Room, 15m2 | FFE | ITAV-006 | CONTROL PANEL: AV system, wall mount | 1 | may be wall mounted or on table top
MEET-15 | Meeting Room, 15m2 | FFE | ITAV-051 | DISPLAY SCREEN: audio visual | 1
MEET-15 | Meeting Room, 15m2 | FFE | FQBS-059 | CHAIR: meeting rom, stackable | 8
MEET-15 | Meeting Room, 15m2 | FFE | ITAV-062 | CONTROL PANEL: room booking system | 1 | optional
MEET-15 | Meeting Room, 15m2 | FFE | ITAV-063 | COMPUTER: AV system | 1
MEET-15 | Meeting Room, 15m2 | FFE | ITAV-090 | SOUNDBAR: AV system, camera, speaker and microphone | 1 | functions may be provided by individual items dependent on local ICT policies and service requirements
MEET-15 | Meeting Room, 15m2 | FFE | FQTA-086 | TABLE: meeting room, folding | 2 | with services box if floor box provided
MEET-15 | Meeting Room, 15m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | optional, 1 to room booking panel and 1 to video conferencing control panel if provided
MEET-15 | Meeting Room, 15m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
MEET-15 | Meeting Room, 15m2 | SE | ITIN-046 | OUTLET: data, quad RJ45, wall mounted | 1 | optional
MEET-15 | Meeting Room, 15m2 | SE | ITIN-022 | OUTLET: data, double RJ45, in floorbox | 1 | optional
MEET-15 | Meeting Room, 15m2 | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 1 | optional; services may be wall mounted depending on layout and service requirements
MEET-15 | Meeting Room, 15m2 | SE | ELBO-015 | SWITCH: light, single | 1
MEET-15 | Meeting Room, 15m2 | SE | ELGP-208 | GPO: double, wall mounted | 4 | 3 are optional, to suit video conferencing and AV equipment as required
MEET-15 | Meeting Room, 15m2 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
MEET-15 | Meeting Room, 15m2 | SE | ELGP-203 | GPO: double, in floorbox | 1 | optional
MEET-15 | Meeting Room, 15m2 | SE | ITIN-145 | DESK BOX: for power, data, AV, flush mounted, softwired | 1 | optional; to table if floor box provided
MEET-20 | Meeting Room, 20m2 | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
MEET-20 | Meeting Room, 20m2 | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
MEET-20 | Meeting Room, 20m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
MEET-20 | Meeting Room, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MEET-20 | Meeting Room, 20m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
MEET-20 | Meeting Room, 20m2 | FA | DWWH-005 | WINDOW | 1 | desirable
MEET-20 | Meeting Room, 20m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional lockable
MEET-20 | Meeting Room, 20m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
MEET-20 | Meeting Room, 20m2 | FFE | DWWD-004 | BLIND: roller | 1
MEET-20 | Meeting Room, 20m2 | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1 | optional
MEET-20 | Meeting Room, 20m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
MEET-20 | Meeting Room, 20m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
MEET-20 | Meeting Room, 20m2 | FFE | FIJO-194 | SHELF: laminate | 1
MEET-20 | Meeting Room, 20m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
MEET-20 | Meeting Room, 20m2 | FFE | ITNE-004 | CAMERA: video conferencing | 1 | optional, multidirectional, may be provided to support camera in soundbar as required
MEET-20 | Meeting Room, 20m2 | FFE | ITAV-006 | CONTROL PANEL: AV system, wall mount | 1 | may be wall mounted or on table top
MEET-20 | Meeting Room, 20m2 | FFE | ITAV-051 | DISPLAY SCREEN: audio visual | 1
MEET-20 | Meeting Room, 20m2 | FFE | FQBS-059 | CHAIR: meeting rom, stackable | 12
MEET-20 | Meeting Room, 20m2 | FFE | ITAV-062 | CONTROL PANEL: room booking system | 1 | optional
MEET-20 | Meeting Room, 20m2 | FFE | ITAV-063 | COMPUTER: AV system | 1
MEET-20 | Meeting Room, 20m2 | FFE | ITAV-071 | MICROPHONE: AV system, ceiling mounted, array | 1 | optional, may be recessed in ceiling
MEET-20 | Meeting Room, 20m2 | FFE | ITAV-090 | SOUNDBAR: AV system, camera, speaker and microphone | 1 | functions may be provided by individual items dependent on local ICT policies and service requirements
MEET-20 | Meeting Room, 20m2 | FFE | FQTA-086 | TABLE: meeting room, folding | 2 | with services box if floor box provided
MEET-20 | Meeting Room, 20m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 is optional for booking system if provided
MEET-20 | Meeting Room, 20m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
MEET-20 | Meeting Room, 20m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 2
MEET-20 | Meeting Room, 20m2 | SE | ITIN-021 | OUTLET: data, double RJ45, ceiling mounted | 1 | to microphone array
MEET-20 | Meeting Room, 20m2 | SE | ITIN-022 | OUTLET: data, double RJ45, in floorbox | 1 | optional
MEET-20 | Meeting Room, 20m2 | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 1 | optional; services may be wall mounted depending on layout and service requirements
MEET-20 | Meeting Room, 20m2 | SE | ELBO-015 | SWITCH: light, single | 1
MEET-20 | Meeting Room, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 4
MEET-20 | Meeting Room, 20m2 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
MEET-20 | Meeting Room, 20m2 | SE | ELGP-203 | GPO: double, in floorbox | 1 | optional
MEET-20 | Meeting Room, 20m2 | SE | ITIN-145 | DESK BOX: for power, data, AV, flush mounted, softwired | 1 | optional; to table if floor box provided
MEET-30 | Meeting Room, 30m2 | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
MEET-30 | Meeting Room, 30m2 | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
MEET-30 | Meeting Room, 30m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
MEET-30 | Meeting Room, 30m2 | FA | WLWA-001 | WALL FINISH: acoustic panels, prefinished | 12 | optional; to suit acoustic requirements
MEET-30 | Meeting Room, 30m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MEET-30 | Meeting Room, 30m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
MEET-30 | Meeting Room, 30m2 | FA | DWWH-005 | WINDOW | 1 | desirable
MEET-30 | Meeting Room, 30m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional lockable
MEET-30 | Meeting Room, 30m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
MEET-30 | Meeting Room, 30m2 | FFE | DWWD-004 | BLIND: roller | 1
MEET-30 | Meeting Room, 30m2 | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
MEET-30 | Meeting Room, 30m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
MEET-30 | Meeting Room, 30m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
MEET-30 | Meeting Room, 30m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
MEET-30 | Meeting Room, 30m2 | FFE | ITNE-004 | CAMERA: video conferencing | 1 | multidirectional
MEET-30 | Meeting Room, 30m2 | FFE | ITAV-057 | SPEAKER: ceiling recessed | 4 | to replace or support audio from soundbar as required
MEET-30 | Meeting Room, 30m2 | FFE | FIJO-011 | BENCH: 450D, open under, laminate finish | 1
MEET-30 | Meeting Room, 30m2 | FFE | ITAV-006 | CONTROL PANEL: AV system, wall mount | 1 | may be wall mounted or on table top
MEET-30 | Meeting Room, 30m2 | FFE | ITAV-051 | DISPLAY SCREEN: audio visual | 1
MEET-30 | Meeting Room, 30m2 | FFE | FIJO-099 | CAC: cable access cap | 1
MEET-30 | Meeting Room, 30m2 | FFE | FIHR-065 | BRACKET: for bench, 450D, wall mounted | 2
MEET-30 | Meeting Room, 30m2 | FFE | FQBS-059 | CHAIR: meeting rom, stackable | 18
MEET-30 | Meeting Room, 30m2 | FFE | ITAV-062 | CONTROL PANEL: room booking system | 1 | optional
MEET-30 | Meeting Room, 30m2 | FFE | ITAV-063 | COMPUTER: AV system | 1
MEET-30 | Meeting Room, 30m2 | FFE | ITAV-071 | MICROPHONE: AV system, ceiling mounted, array | 1 | may be recessed in ceiling
MEET-30 | Meeting Room, 30m2 | FFE | FQTA-086 | TABLE: meeting room, folding | 3 | with services boxes to connect to floor boxes
MEET-30 | Meeting Room, 30m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 is optional for booking system if provided
MEET-30 | Meeting Room, 30m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
MEET-30 | Meeting Room, 30m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 2
MEET-30 | Meeting Room, 30m2 | SE | ITIN-021 | OUTLET: data, double RJ45, ceiling mounted | 1 | to microphone array
MEET-30 | Meeting Room, 30m2 | SE | ITIN-022 | OUTLET: data, double RJ45, in floorbox | 2
MEET-30 | Meeting Room, 30m2 | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 2
MEET-30 | Meeting Room, 30m2 | SE | ELBO-015 | SWITCH: light, single | 1
MEET-30 | Meeting Room, 30m2 | SE | ELGP-208 | GPO: double, wall mounted | 4
MEET-30 | Meeting Room, 30m2 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
MEET-30 | Meeting Room, 30m2 | SE | ELGP-203 | GPO: double, in floorbox | 2
MEET-30 | Meeting Room, 30m2 | SE | ITIN-145 | DESK BOX: for power, data, AV, flush mounted, softwired | 2 | to tables to connect to floor boxes
MEET-55 | Meeting Room, 55m2 | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
MEET-55 | Meeting Room, 55m2 | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
MEET-55 | Meeting Room, 55m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 2
MEET-55 | Meeting Room, 55m2 | FA | WLWA-001 | WALL FINISH: acoustic panels, prefinished | 12 | optional; to suit acoustic requirements
MEET-55 | Meeting Room, 55m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MEET-55 | Meeting Room, 55m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
MEET-55 | Meeting Room, 55m2 | FA | DWGL-052 | DOOR: 1600mm c/o, double leaf, observation panel, solid core/glass, paint | 1 | optional lockable
MEET-55 | Meeting Room, 55m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
MEET-55 | Meeting Room, 55m2 | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
MEET-55 | Meeting Room, 55m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
MEET-55 | Meeting Room, 55m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
MEET-55 | Meeting Room, 55m2 | FFE | FQDW-006 | LECTERN: mobile | 1
MEET-55 | Meeting Room, 55m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
MEET-55 | Meeting Room, 55m2 | FFE | ITNE-004 | CAMERA: video conferencing | 1
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-057 | SPEAKER: ceiling recessed | 6
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-006 | CONTROL PANEL: AV system, wall mount | 1
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 1
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-051 | DISPLAY SCREEN: audio visual | 1
MEET-55 | Meeting Room, 55m2 | FFE | FQBS-059 | CHAIR: meeting rom, stackable | 42
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-062 | CONTROL PANEL: room booking system | 1 | optional
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-063 | COMPUTER: AV system | 1
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-064 | CAMERA: AV system, ceiling mounted | 1 | optional, multidirectional, to rear of room for view of speaker
MEET-55 | Meeting Room, 55m2 | FFE | ITAV-071 | MICROPHONE: AV system, ceiling mounted, array | 2 | may be recessed in ceiling
MEET-55 | Meeting Room, 55m2 | FFE | FQTA-086 | TABLE: meeting room, folding | 4 | quantity and size/type to suit requirements
MEET-55 | Meeting Room, 55m2 | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 1 | optional, to speaker camera if provided
MEET-55 | Meeting Room, 55m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 is optional for booking system if provided
MEET-55 | Meeting Room, 55m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 2
MEET-55 | Meeting Room, 55m2 | SE | ITIN-021 | OUTLET: data, double RJ45, ceiling mounted | 2 | to microphone array
MEET-55 | Meeting Room, 55m2 | SE | ITIN-022 | OUTLET: data, double RJ45, in floorbox | 3 | 1 to each floorbox if provided
MEET-55 | Meeting Room, 55m2 | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 3 | 1 to lectern, 2 optional to support flexible layout, location to be confirmed to suit service requirements
MEET-55 | Meeting Room, 55m2 | SE | ELBO-015 | SWITCH: light, single | 1
MEET-55 | Meeting Room, 55m2 | SE | ELGP-208 | GPO: double, wall mounted | 4
MEET-55 | Meeting Room, 55m2 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 | as required
MEET-55 | Meeting Room, 55m2 | SE | ELGP-203 | GPO: double, in floorbox | 6 | 2 to each floorbox if provided
MEET-9 | Meeting Room, 9m2 | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
MEET-9 | Meeting Room, 9m2 | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
MEET-9 | Meeting Room, 9m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
MEET-9 | Meeting Room, 9m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
MEET-9 | Meeting Room, 9m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
MEET-9 | Meeting Room, 9m2 | FA | DWWH-005 | WINDOW | 1 | desirable
MEET-9 | Meeting Room, 9m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional lockable
MEET-9 | Meeting Room, 9m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
MEET-9 | Meeting Room, 9m2 | FFE | DWWD-004 | BLIND: roller | 1
MEET-9 | Meeting Room, 9m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
MEET-9 | Meeting Room, 9m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1 | optional
MEET-9 | Meeting Room, 9m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
MEET-9 | Meeting Room, 9m2 | FFE | ITAV-006 | CONTROL PANEL: AV system, wall mount | 1 | optional; may be wall mounted or on table top
MEET-9 | Meeting Room, 9m2 | FFE | ITAV-051 | DISPLAY SCREEN: audio visual | 1 | optional
MEET-9 | Meeting Room, 9m2 | FFE | FQTA-023 | TABLE: rectangular, 1500L x 900W, work height | 1
MEET-9 | Meeting Room, 9m2 | FFE | FQBS-059 | CHAIR: meeting rom, stackable | 4
MEET-9 | Meeting Room, 9m2 | FFE | ITAV-062 | CONTROL PANEL: room booking system | 1 | optional
MEET-9 | Meeting Room, 9m2 | FFE | ITAV-063 | COMPUTER: AV system | 1 | optional
MEET-9 | Meeting Room, 9m2 | FFE | ITAV-090 | SOUNDBAR: AV system, camera, speaker and microphone | 1 | optional; functions may be provided by individual items dependent on local ICT policies and service requirements
MEET-9 | Meeting Room, 9m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | optional, 1 to room booking panel and 1 to video conferencing control panel if provided
MEET-9 | Meeting Room, 9m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
MEET-9 | Meeting Room, 9m2 | SE | ITIN-046 | OUTLET: data, quad RJ45, wall mounted | 1 | optional, to suit video conferencing and AV equipment as required
MEET-9 | Meeting Room, 9m2 | SE | ELBO-015 | SWITCH: light, single | 1
MEET-9 | Meeting Room, 9m2 | SE | ELGP-208 | GPO: double, wall mounted | 4 | 3 are optional, to suit video conferencing and AV equipment as required
MOR-AU | Mortuary - Autopsy Room | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1 | junction with evisceration table to be considered to prevent water ingress
MOR-AU | Mortuary - Autopsy Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
MOR-AU | Mortuary - Autopsy Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
MOR-AU | Mortuary - Autopsy Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
MOR-AU | Mortuary - Autopsy Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MOR-AU | Mortuary - Autopsy Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
MOR-AU | Mortuary - Autopsy Room | FA | DWOB-005 | OBSERVATION: internal window, alum/glass, powdercoat | 1 | to observation area if provided
MOR-AU | Mortuary - Autopsy Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | door seals required
MOR-AU | Mortuary - Autopsy Room | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | door seals required
MOR-AU | Mortuary - Autopsy Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MOR-AU | Mortuary - Autopsy Room | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 8
MOR-AU | Mortuary - Autopsy Room | FFE | FIJO-021 | BENCH: 600D, open under, stainless steel lipped | 1 | with wet edge and integral sink
MOR-AU | Mortuary - Autopsy Room | FFE | FIHR-016 | BRACKET: for sharps bin | 2
MOR-AU | Mortuary - Autopsy Room | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
MOR-AU | Mortuary - Autopsy Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
MOR-AU | Mortuary - Autopsy Room | FFE | FQCW-004 | BIN: waste, clinical, 120L | 1
MOR-AU | Mortuary - Autopsy Room | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
MOR-AU | Mortuary - Autopsy Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
MOR-AU | Mortuary - Autopsy Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
MOR-AU | Mortuary - Autopsy Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
MOR-AU | Mortuary - Autopsy Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 2 | to basin
MOR-AU | Mortuary - Autopsy Room | FFE | FIHR-029 | HOOK: equipment | 1 | optional, for lifter straps/slings
MOR-AU | Mortuary - Autopsy Room | FFE | MMGE-119 | TROLLEY: equipment | 2
MOR-AU | Mortuary - Autopsy Room | FFE | FIRT-005 | GRAB RAIL: stainless steel | 2 | to boot wash
MOR-AU | Mortuary - Autopsy Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
MOR-AU | Mortuary - Autopsy Room | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
MOR-AU | Mortuary - Autopsy Room | FFE | FQBS-051 | STOOL: mobile, adj height | 1 | optional, dependent on project requirements
MOR-AU | Mortuary - Autopsy Room | FFE | FQCW-015 | BIN: waste, general, 120L | 1
MOR-AU | Mortuary - Autopsy Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
MOR-AU | Mortuary - Autopsy Room | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
MOR-AU | Mortuary - Autopsy Room | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 2
MOR-AU | Mortuary - Autopsy Room | FFE | HYTP-023 | SHOWER: safety, eyewash combination | 1
MOR-AU | Mortuary - Autopsy Room | FFE | LIFX-014 | LIGHT: procedure, ceiling mounted | 1 | hardwired, a second light may be provided dependent on service requirements
MOR-AU | Mortuary - Autopsy Room | FFE | HYTP-030 | TAPWARE: cold | 1 | to boot wash
MOR-AU | Mortuary - Autopsy Room | FFE | HYTP-034 | TAPWARE: laboratory | 1 | to dissection table
MOR-AU | Mortuary - Autopsy Room | FFE | HYTP-003 | HOSE REEL: retractable | 1
MOR-AU | Mortuary - Autopsy Room | FFE | HYBA-055 | SINK: single, 500 x 450 x 250D | 1
MOR-AU | Mortuary - Autopsy Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1 | suitable for use in wet zone
MOR-AU | Mortuary - Autopsy Room | FFE | FIJO-100 | BENCH LEGS: stainless steel, for bench top | 1
MOR-AU | Mortuary - Autopsy Room | FFE | HYTP-028 | SPOUT: bath, wall mounted | 1 | to boot wash
MOR-AU | Mortuary - Autopsy Room | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1
MOR-AU | Mortuary - Autopsy Room | FFE | ITNE-066 | COMPUTER: all-in-one, wall mounted | 1 | suitable for use in wet zone
MOR-AU | Mortuary - Autopsy Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | suitable for use in wet zone
MOR-AU | Mortuary - Autopsy Room | FFE | MMBE-070 | LIFTER: patient, ceiling mounted, motor and sling component, mortuary | 1 | optional, mobile lifter may be used, bariatric capacity to be confirmed based on service requirements; control may be suspended from motor or wall mounted depending on model
MOR-AU | Mortuary - Autopsy Room | FFE | MMLA-086 | TABLE: evisceration, height adjustable | 1 | bariatric capacity to be determined based on service requirements, services on table to be considered: suction capabilities (built-in or provided), power to table, power outlets (if not on table), down draft (optional, built-in to table), water connection (built-in to table)
MOR-AU | Mortuary - Autopsy Room | FFE | MMLA-087 | TABLE: dissection, height adjustable | 1 | with integral sink
MOR-AU | Mortuary - Autopsy Room | FFE | MMGE-185 | SCALES: digital, benchtop | 1
MOR-AU | Mortuary - Autopsy Room | FFE | MMBE-071 | LIFTER: patient, ceiling mounted, track, linear, mortuary | 1 | optional, mobile lifter may be used, bariatric capacity to be confirmed based on service requirements
MOR-AU | Mortuary - Autopsy Room | FFE | LIFX-037 | LIGHT: LED strip, ceiling mounted | 1 | task lighting over dissection table, articulated task light is also acceptable dependent on service requirements
MOR-AU | Mortuary - Autopsy Room | FFE | FIRT-046 | GRAB RAIL: stainless steel, freestanding | 1 | to boot wash
MOR-AU | Mortuary - Autopsy Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3 | IP rated to suit AS/NZS wet zone requirements
MOR-AU | Mortuary - Autopsy Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | IP rated to suit AS/NZS wet zone requirements
MOR-AU | Mortuary - Autopsy Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | suitable for use in wet zone
MOR-AU | Mortuary - Autopsy Room | SE | HYTP-013 | OUTLET: hot water | 2 | connection only, to evisceration and dissection tables, dependent on service requirements and vendor specifications
MOR-AU | Mortuary - Autopsy Room | SE | HYTP-008 | OUTLET: cold water | 2 | connection only, to evisceration and dissection tables, dependent on service requirements and vendor specifications
MOR-AU | Mortuary - Autopsy Room | SE | HYDR-001 | DRAIN: special requirements | 1 | to evisceration table, confirm requirements at project level
MOR-AU | Mortuary - Autopsy Room | SE | HYDR-004 | DRAIN: floor waste | 1 | to boot wash
MOR-AU | Mortuary - Autopsy Room | SE | HYDR-003 | DRAIN: floor trough | 2
MOR-AU | Mortuary - Autopsy Room | SE | HYTP-012 | OUTLET: hosecock, cold water | 1 | to hose reel
MOR-AU | Mortuary - Autopsy Room | SE | ELBO-015 | SWITCH: light, single | 2 | 1x general room lights, 1x task light over dissection bench. IP rated to suit AS/NZS wet zone requirements
MOR-AU | Mortuary - Autopsy Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | IP rated to suit AS/NZS wet zone requirements
MOR-AU | Mortuary - Autopsy Room | SE | ELGP-106 | GPO: single, special, for equipment | 2 | power to evisceration and dissection tables (if not part of table configuration). Required to be emergency power/generator backed. IP rated to suit AS/NZS wet zone requirements
MOR-AU | Mortuary - Autopsy Room | SE | ELGP-106 | GPO: single, special, for equipment | 8 | 4x outlets on evisceration and dissection tables, dependent on service requirements and vendor specifications. Required to be emergency power/generator backed. IP rated to suit AS/NZS wet zone requirements
MOR-AU | Mortuary - Autopsy Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to computer. IP rated to suit AS/NZS wet zone requirements
MOR-AU | Mortuary - Autopsy Room | SE | MGAS-060 | OUTLET: suction | 2 | on evisceration and dissection tables, dependent on service requirements and vendor specifications
MOR-AU | Mortuary - Autopsy Room | SE | HYGE-026 | TROUGH: boot wash | 1
MOR-AU | Mortuary - Autopsy Room | SE | MEEX-017 | EXHAUST: downdraft, mortuary | 1 | optional, on evisceration and dissection tables, dependent on service requirements and vendor specifications
MOR-BH | Mortuary - Body Holding | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1
MOR-BH | Mortuary - Body Holding | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
MOR-BH | Mortuary - Body Holding | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
MOR-BH | Mortuary - Body Holding | FA | WLPR-021 | WALL PROTECTION: crash rail, at 900 AFFL | 2
MOR-BH | Mortuary - Body Holding | FA | WLPR-006 | WALL PROTECTION: corner guards | 1
MOR-BH | Mortuary - Body Holding | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
MOR-BH | Mortuary - Body Holding | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MOR-BH | Mortuary - Body Holding | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
MOR-BH | Mortuary - Body Holding | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
MOR-BH | Mortuary - Body Holding | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MOR-BH | Mortuary - Body Holding | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 5
MOR-BH | Mortuary - Body Holding | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
MOR-BH | Mortuary - Body Holding | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
MOR-BH | Mortuary - Body Holding | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
MOR-BH | Mortuary - Body Holding | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
MOR-BH | Mortuary - Body Holding | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
MOR-BH | Mortuary - Body Holding | FFE | FIHR-029 | HOOK: equipment | 1 | for lifter straps/slings
MOR-BH | Mortuary - Body Holding | FFE | MMGE-133 | TROLLEY: mortuary, hydraulic | 1 | bariatric capacity to be confirmed based on service requirements
MOR-BH | Mortuary - Body Holding | FFE | HYBA-010 | BASIN: handwash, Type B | 1
MOR-BH | Mortuary - Body Holding | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
MOR-BH | Mortuary - Body Holding | FFE | HYTP-003 | HOSE REEL: retractable | 1
MOR-BH | Mortuary - Body Holding | FFE | MMGE-183 | FRIDGE: body holding, triple tier | 3 | bariatric capacity to be confirmed based on service requirements; power and BMS requirements to be confirmed by project
MOR-BH | Mortuary - Body Holding | FFE | MMBE-070 | LIFTER: patient, ceiling mounted, motor and sling component, mortuary | 1 | mobile lifter may be used, bariatric capacity to be confirmed based on service requirements; control may be suspended from motor or wall mounted depending on model
MOR-BH | Mortuary - Body Holding | FFE | MMBE-071 | LIFTER: patient, ceiling mounted, track, linear, mortuary | 1 | mobile lifter may be used, bariatric capacity to be confirmed based on service requirements
MOR-BH | Mortuary - Body Holding | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3 | temperature monitoring, IP rated to suit AS/NZS wet zone requirements, 1 to each refrigerator, BMS requirements to be confirmed by project
MOR-BH | Mortuary - Body Holding | SE | ITBU-001 | ALARM: equipment monitoring, temperature | 1 | temperature monitoring, alarm and BMS requirements to be confirmed by project
MOR-BH | Mortuary - Body Holding | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to lifter
MOR-BH | Mortuary - Body Holding | SE | HYDR-004 | DRAIN: floor waste | 1
MOR-BH | Mortuary - Body Holding | SE | HYTP-012 | OUTLET: hosecock, cold water | 1
MOR-BH | Mortuary - Body Holding | SE | ELBO-015 | SWITCH: light, single | 1 | IP rated to suit AS/NZS wet zone requirements
MOR-BH | Mortuary - Body Holding | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 | IP rated to suit AS/NZS wet zone requirements
MOR-BH | Mortuary - Body Holding | SE | ELGP-106 | GPO: single, special, for equipment | 3 | 1 to each refrigerator; required to be emergency power/generator backed. IP rated to suit AS/NZS wet zone requirements
MOR-BR | Mortuary - Body Reception | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
MOR-BR | Mortuary - Body Reception | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 2
MOR-BR | Mortuary - Body Reception | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 2
MOR-BR | Mortuary - Body Reception | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
MOR-BR | Mortuary - Body Reception | FA | CLCN-009 | CORNICE: wall trim fixing | 1
MOR-BR | Mortuary - Body Reception | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
MOR-BR | Mortuary - Body Reception | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
MOR-BR | Mortuary - Body Reception | FFE | FIDI-018 | DISPENSER: gown | 1
MOR-BR | Mortuary - Body Reception | FFE | FIDI-021 | DISPENSER: mask | 1
MOR-BR | Mortuary - Body Reception | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 2
MOR-BR | Mortuary - Body Reception | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
MOR-BR | Mortuary - Body Reception | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
MOR-BR | Mortuary - Body Reception | FFE | ITNE-023 | PRINTER: networked | 1
MOR-BR | Mortuary - Body Reception | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
MOR-BR | Mortuary - Body Reception | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
MOR-BR | Mortuary - Body Reception | FFE | FQDW-045 | DESK: adj height, 2400W x 750D x 750H nom. | 1 | length to accommodate A3 landscape mortuary register
MOR-BR | Mortuary - Body Reception | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | 1 outlet for telephone
MOR-BR | Mortuary - Body Reception | SE | LIFX-008 | LIGHT: task light, built in | 2 | to underside of overhead cupboards
MOR-BR | Mortuary - Body Reception | SE | ELBO-015 | SWITCH: light, single | 1 | for task lighting
MOR-BR | Mortuary - Body Reception | SE | ELGP-208 | GPO: double, wall mounted | 3 | emergency power may be provided depending on operational policies and service requirements
MOR-VR | Mortuary - Viewing Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
MOR-VR | Mortuary - Viewing Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
MOR-VR | Mortuary - Viewing Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
MOR-VR | Mortuary - Viewing Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MOR-VR | Mortuary - Viewing Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
MOR-VR | Mortuary - Viewing Room | FA | DWOB-005 | OBSERVATION: internal window, alum/glass, powdercoat | 1
MOR-VR | Mortuary - Viewing Room | FA | DWPR-003 | DOOR PROTECTION: kickplate, to 300 AFFL, vinyl, prefinished | 1
MOR-VR | Mortuary - Viewing Room | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1
MOR-VR | Mortuary - Viewing Room | FA | DWGL-052 | DOOR: 1600mm c/o, double leaf, observation panel, solid core/glass, paint | 1
MOR-VR | Mortuary - Viewing Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MOR-VR | Mortuary - Viewing Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
MOR-VR | Mortuary - Viewing Room | FFE | FIRT-031 | TRACK: curtain, privacy | 1
MOR-VR | Mortuary - Viewing Room | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | heavy/domestic fabric is acceptable
MOR-VR | Mortuary - Viewing Room | FFE | MMGE-186 | TROLLEY: mortuary, transport | 1
MOR-VR | Mortuary - Viewing Room | SE | ELBO-015 | SWITCH: light, single | 1
MOR-W | Mortuary - Waiting | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
MOR-W | Mortuary - Waiting | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 6
MOR-W | Mortuary - Waiting | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
MOR-W | Mortuary - Waiting | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MOR-W | Mortuary - Waiting | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
MOR-W | Mortuary - Waiting | FA | DWOB-005 | OBSERVATION: internal window, alum/glass, powdercoat | 1
MOR-W | Mortuary - Waiting | FA | DWPR-003 | DOOR PROTECTION: kickplate, to 300 AFFL, vinyl, prefinished | 2
MOR-W | Mortuary - Waiting | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | lockable
MOR-W | Mortuary - Waiting | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | lockable
MOR-W | Mortuary - Waiting | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MOR-W | Mortuary - Waiting | FFE | FQGE-002 | ARTWORK | 1
MOR-W | Mortuary - Waiting | FFE | FQBS-032 | CHAIR: waiting | 3
MOR-W | Mortuary - Waiting | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
MOR-W | Mortuary - Waiting | FFE | FQCW-016 | BIN: waste, general, 20L | 1
MOR-W | Mortuary - Waiting | FFE | FQBS-043 | LOUNGE: 2 seat, 1600 wide | 1
MOR-W | Mortuary - Waiting | FFE | FQTA-034 | TABLE: rectangular, 900L x 600W, low height | 1
MOR-W | Mortuary - Waiting | SE | ELBO-015 | SWITCH: light, single | 1
MOR-W | Mortuary - Waiting | SE | ELGP-109 | GPO: single, wall mounted | 2
MOR-W | Mortuary - Waiting | SE | ITSE-006 | BUTTON: duress, fixed | 1
MRICR | MRI Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
MRICR | MRI Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 4
MRICR | MRI Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
MRICR | MRI Control Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
MRICR | MRI Control Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
MRICR | MRI Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 3
MRICR | MRI Control Room | FFE | DWWD-004 | BLIND: roller | 1
MRICR | MRI Control Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
MRICR | MRI Control Room | FFE | FIHR-026 | HOOK: coat | 2 | optional
MRICR | MRI Control Room | FFE | FQSN-031 | CABINET: short, single door, tambour, adj shelves, lockable | 1
MRICR | MRI Control Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
MRICR | MRI Control Room | FFE | FQSN-050 | LOCKER: 4 tier (quarter size) | 2 | optional
MRICR | MRI Control Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
MRICR | MRI Control Room | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 2 | for RIS/PMS and patient observation
MRICR | MRI Control Room | FFE | MMGE-032 | DISPLAY SCREEN: slave, for patient monitor | 1 | wall mounted, on bracket, cardiac/ anaesthetic monitor
MRICR | MRI Control Room | FFE | ITNE-023 | PRINTER: networked | 1 | on benchtop
MRICR | MRI Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1
MRICR | MRI Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
MRICR | MRI Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 1 | on bench
MRICR | MRI Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 1 | high bench at 900-950 AFFL optional
MRICR | MRI Control Room | FFE | MMMI-171 | BOX: alarm, MRI | 1
MRICR | MRI Control Room | FFE | MMMI-193 | COMPUTER: tower, acquisition, MRI | 1
MRICR | MRI Control Room | FFE | MMMI-195 | WINDOW: radiofrequency, MRI | 1
MRICR | MRI Control Room | FFE | MMMI-197 | BUTTON: quench, MRI | 1
MRICR | MRI Control Room | FFE | MMMI-199 | PANEL: waveguide, MRI | 2
MRICR | MRI Control Room | FFE | MMMI-202 | CONSOLE: control, MRI | 1
MRICR | MRI Control Room | FFE | FIJO-099 | CAC: cable access cap | 2
MRICR | MRI Control Room | FFE | FIHR-066 | BRACKET: for display screen, wall mounted, articulated, adj height | 1
MRICR | MRI Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 2
MRICR | MRI Control Room | FFE | MMMI-369 | DISPLAY SCREEN: contrast injector | 1
MRICR | MRI Control Room | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 2
MRICR | MRI Control Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 2
MRICR | MRI Control Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
MRICR | MRI Control Room | FFE | ITNE-070 | DISPLAY SCREEN: PACS | 2
MRICR | MRI Control Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3 | 1 outlet for patient slave monitor, one outlet for telephone
MRICR | MRI Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 4
MRICR | MRI Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1
MRICR | MRI Control Room | SE | ELGP-135 | GPO: UPS, single, wall mounted | 4
MRICR | MRI Control Room | SE | MGFP-001 | ALARM PANEL: medical gases | 1
MRICR | MRI Control Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
MRICR | MRI Control Room | SE | ELGP-208 | GPO: double, wall mounted | 2 | 1 for printer
MRICR | MRI Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 5
MRICR | MRI Control Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
MRICR | MRI Control Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
MRICR | MRI Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
MRICR | MRI Control Room | SE | MEGE-014 | ALARM: oxygen depletion | 1
MRICR | MRI Control Room | SE | HYGE-025 | CONTROL PANEL: process cooling water pump | 2
MRICR | MRI Control Room | SE | MEGE-017 | CONTROL PANEL: HVAC | 1
MRIR | MRI Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
MRIR | MRI Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
MRIR | MRI Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
MRIR | MRI Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
MRIR | MRI Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
MRIR | MRI Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | tiled ceiling also acceptable; ceiling treatment to enhance patient experience to be considered
MRIR | MRI Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MRIR | MRI Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 9
MRIR | MRI Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | MRI compatible
MRIR | MRI Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | MRI compatible
MRIR | MRI Room | FFE | FIJO-116 | CUPBOARD: tall, double door, adj shelves | 3 | for coils
MRIR | MRI Room | FFE | FQBS-037 | FOOTSTOOL | 1 | MRI compatible
MRIR | MRI Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | MRI compatible
MRIR | MRI Room | FFE | MMGE-002 | ADAPTER: suction | 2 | MRI compatible
MRIR | MRI Room | FFE | MMGE-100 | SUCTION BOTTLE | 2 | MRI compatible
MRIR | MRI Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
MRIR | MRI Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1
MRIR | MRI Room | FFE | MMMI-175 | CABIN: radio frequency, faraday cage, MRI | 1 | 3 dimensions
MRIR | MRI Room | FFE | MMMI-180 | DEVICE: imaging, magnet, MRI | 1 | refer to manufacturer's requirements for quench pipes and other associated equipment
MRIR | MRI Room | FFE | MMMI-181 | DOOR: radio frequency, MRI | 1 | 1200mm clear opening; with door protection to 900mm AFFL
MRIR | MRI Room | FFE | MMMI-185 | PIPE: quench, MRI | 1 | installed to manufacturer's specifications with minimal bends
MRIR | MRI Room | FFE | MMMI-183 | INJECTOR: contrast, MRI | 1
MRIR | MRI Room | FFE | MMMI-188 | CAMERA: observation, MRI | 2 | with direct view of patient
MRIR | MRI Room | FFE | MMMI-191 | TABLE: patient, MRI | 1 | detachable table also acceptable; use of detachable table may also negate need for MRI compatible patient transport trolley
MRIR | MRI Room | FFE | MMMI-195 | WINDOW: radiofrequency, MRI | 1 | 1 to control, 1 to anaesthetic preparation
MRIR | MRI Room | FFE | MMMI-198 | MACHINE: anaesthetic, MRI Safe | 1
MRIR | MRI Room | FFE | MMMI-199 | PANEL: waveguide, MRI | 4
MRIR | MRI Room | FFE | MMMI-200 | TROLLEY: patient, transport, MRI Safe | 1
MRIR | MRI Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | to cupboards
MRIR | MRI Room | FFE | MMMI-368 | LASER: positioning lights, MRI | 1 | optional, for radiotherapy planning
MRIR | MRI Room | FFE | FIHR-082 | BRACKET: for observation camera, MRI Safe, wall mounted | 2
MRIR | MRI Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
MRIR | MRI Room | SE | ELPR-006 | RCD: residual current device | 2 | 1 for each UPS & emergency power circuits
MRIR | MRI Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | 1-laser positioning light if required
MRIR | MRI Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
MRIR | MRI Room | SE | ELBO-001 | BUTTON: emergency stop | 1
MRIR | MRI Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1
MRIR | MRI Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2
MRIR | MRI Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1
MRIR | MRI Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
MRIR | MRI Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2
MRIR | MRI Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
MRIR | MRI Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | 1-clock
MRIR | MRI Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 2
MRIR | MRI Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
MRIR | MRI Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
MRIR | MRI Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | MRI compatible
MRIR | MRI Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
MRIR | MRI Room | SE | MMMI-201 | PANEL: penetration, MRI | 1
MRIR | MRI Room | SE | ELGP-133 | GPO: UPS, single, on services panel | 2
MRIR | MRI Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4
MRIR | MRI Room | SE | ELPR-007 | UPS: status indicator | 1
MRIR | MRI Room | SE | MEGE-016 | SENSOR: oxygen depletion | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 2
OFF-1P-12 | Office - 1 Person, 12m2 | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
OFF-1P-12 | Office - 1 Person, 12m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
OFF-1P-12 | Office - 1 Person, 12m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FA | DWWH-005 | WINDOW | 1 | desirable
OFF-1P-12 | Office - 1 Person, 12m2 | FA | DWOB-002 | OBSERVATION: internal window, alum/glass, door height, powdercoat | 1 | with privacy film
OFF-1P-12 | Office - 1 Person, 12m2 | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | lockable
OFF-1P-12 | Office - 1 Person, 12m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FQBS-029 | CHAIR: visitor, office areas | 3
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | DWWD-004 | BLIND: roller | 1 | to window if provided
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FIHR-026 | HOOK: coat | 2 | optional, prioritize lower hook for wheelchair users if only one hook provided
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1 | lockable
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | dual screens may be provided to suit service requirements
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FQTA-057 | TABLE: round, 900 diam, work height | 1
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FQSN-135 | CABINET: single door, tambour, adj shelves, lockable, 1200H | 1 | lockable, size to suit service requirements
OFF-1P-12 | Office - 1 Person, 12m2 | FFE | FQDW-054 | WORKSTATION: straight, adj height, 1600W x 750D x 750H nom. | 1
OFF-1P-12 | Office - 1 Person, 12m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1
OFF-1P-12 | Office - 1 Person, 12m2 | SE | ELBO-015 | SWITCH: light, single | 1
OFF-1P-12 | Office - 1 Person, 12m2 | SE | ELGP-208 | GPO: double, wall mounted | 4 | extent of emergency power to be determined based on service requirements
OFF-1P-12 | Office - 1 Person, 12m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1 | inclusion dependent on location, not required if room located in staff only area
OFF-1P-9 | Office - 1 Person, 9m2 | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
OFF-1P-9 | Office - 1 Person, 9m2 | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
OFF-1P-9 | Office - 1 Person, 9m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
OFF-1P-9 | Office - 1 Person, 9m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
OFF-1P-9 | Office - 1 Person, 9m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
OFF-1P-9 | Office - 1 Person, 9m2 | FA | DWWH-005 | WINDOW | 1 | desirable
OFF-1P-9 | Office - 1 Person, 9m2 | FA | DWOB-002 | OBSERVATION: internal window, alum/glass, door height, powdercoat | 1 | with privacy film
OFF-1P-9 | Office - 1 Person, 9m2 | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | lockable
OFF-1P-9 | Office - 1 Person, 9m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FQBS-029 | CHAIR: visitor, office areas | 2
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | DWWD-004 | BLIND: roller | 1 | to window if provided
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FIHR-026 | HOOK: coat | 2 | optional, prioritize lower hook for wheelchair users if only one hook provided
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1 | lockable
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | dual screens may be provided to suit service requirements
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FQSN-135 | CABINET: single door, tambour, adj shelves, lockable, 1200H | 1 | lockable, size to suit service requirements
OFF-1P-9 | Office - 1 Person, 9m2 | FFE | FQDW-054 | WORKSTATION: straight, adj height, 1600W x 750D x 750H nom. | 1
OFF-1P-9 | Office - 1 Person, 9m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1
OFF-1P-9 | Office - 1 Person, 9m2 | SE | ELBO-015 | SWITCH: light, single | 1
OFF-1P-9 | Office - 1 Person, 9m2 | SE | ELGP-208 | GPO: double, wall mounted | 3 | extent of emergency power to be determined based on service requirements
OFF-1P-9 | Office - 1 Person, 9m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1 | inclusion dependent on location, not required if room located in staff only area
OFF-2P | Office - 2 Person | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
OFF-2P | Office - 2 Person | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5
OFF-2P | Office - 2 Person | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
OFF-2P | Office - 2 Person | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
OFF-2P | Office - 2 Person | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
OFF-2P | Office - 2 Person | FA | DWWH-005 | WINDOW | 1 | desirable
OFF-2P | Office - 2 Person | FA | DWOB-002 | OBSERVATION: internal window, alum/glass, door height, powdercoat | 1 | with privacy film
OFF-2P | Office - 2 Person | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | lockable
OFF-2P | Office - 2 Person | FA | CLCN-009 | CORNICE: wall trim fixing | 1
OFF-2P | Office - 2 Person | FFE | FQBS-018 | CHAIR: office, ergonomic | 2
OFF-2P | Office - 2 Person | FFE | FQBS-029 | CHAIR: visitor, office areas | 2
OFF-2P | Office - 2 Person | FFE | DWWD-004 | BLIND: roller | 1 | to window if provided
OFF-2P | Office - 2 Person | FFE | FIHR-026 | HOOK: coat | 2 | optional, prioritize lower hook for wheelchair users if only one hook provided
OFF-2P | Office - 2 Person | FFE | FQCW-016 | BIN: waste, general, 20L | 1
OFF-2P | Office - 2 Person | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 2 | lockable
OFF-2P | Office - 2 Person | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 2 | optional
OFF-2P | Office - 2 Person | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 2 | dual screens may be provided to suit service requirements
OFF-2P | Office - 2 Person | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2
OFF-2P | Office - 2 Person | FFE | FQDW-054 | WORKSTATION: straight, adj height, 1600W x 750D x 750H nom. | 2
OFF-2P | Office - 2 Person | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 2
OFF-2P | Office - 2 Person | SE | ELBO-015 | SWITCH: light, single | 1
OFF-2P | Office - 2 Person | SE | ELGP-208 | GPO: double, wall mounted | 6 | extent of emergency power to be determined based on service requirements
OFF-2P | Office - 2 Person | SE | ITSE-006 | BUTTON: duress, fixed | 1 | inclusion dependent on location, not required if room located in staff only area
OFF-CLN | Office - Clinical Workroom | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | carpet tiles are also acceptable
OFF-CLN | Office - Clinical Workroom | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
OFF-CLN | Office - Clinical Workroom | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
OFF-CLN | Office - Clinical Workroom | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
OFF-CLN | Office - Clinical Workroom | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
OFF-CLN | Office - Clinical Workroom | FA | DWOB-002 | OBSERVATION: internal window, alum/glass, door height, powdercoat | 1 | optional
OFF-CLN | Office - Clinical Workroom | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
OFF-CLN | Office - Clinical Workroom | FA | CLCN-009 | CORNICE: wall trim fixing | 1
OFF-CLN | Office - Clinical Workroom | FFE | FQBS-018 | CHAIR: office, ergonomic | 4
OFF-CLN | Office - Clinical Workroom | FFE | FQCW-016 | BIN: waste, general, 20L | 1
OFF-CLN | Office - Clinical Workroom | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1 | to patient flow monitor
OFF-CLN | Office - Clinical Workroom | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 6 | overhead joinery may be configured to include pigeon holes for record storage as required
OFF-CLN | Office - Clinical Workroom | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
OFF-CLN | Office - Clinical Workroom | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional, above bench
OFF-CLN | Office - Clinical Workroom | FFE | ITBU-013 | DISPLAY SCREEN: patient flow system | 1
OFF-CLN | Office - Clinical Workroom | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 4
OFF-CLN | Office - Clinical Workroom | FFE | ITNE-023 | PRINTER: networked | 1
OFF-CLN | Office - Clinical Workroom | FFE | ITNE-049 | TELEPHONE: handset, desktop | 4
OFF-CLN | Office - Clinical Workroom | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 1
OFF-CLN | Office - Clinical Workroom | FFE | FIJO-099 | CAC: cable access cap | 4
OFF-CLN | Office - Clinical Workroom | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead is also acceptable
OFF-CLN | Office - Clinical Workroom | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | to patient flow monitor
OFF-CLN | Office - Clinical Workroom | FFE | FIHR-080 | BRACKET: for bench, 750D, wall mounted | 3 | legs are also acceptable, quantity to be confirmed to achieve required support, ensure location does not impact leg clearance at workstations
OFF-CLN | Office - Clinical Workroom | FFE | ITCL-037 | CHARGER: wireless duress devices | 2 | optional
OFF-CLN | Office - Clinical Workroom | FFE | FIHR-093 | BRACKET: keyboard shelf, wall mounted | 1 | to patient flow monitor
OFF-CLN | Office - Clinical Workroom | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 x to patient flow monitor, 1 x to printer
OFF-CLN | Office - Clinical Workroom | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 4 | location of outlets to meet local operational/WHS requirements
OFF-CLN | Office - Clinical Workroom | SE | LIFX-008 | LIGHT: task light, built in | 6 | under wall mounted joinery
OFF-CLN | Office - Clinical Workroom | SE | ELBO-015 | SWITCH: light, single | 2 | 1 x general room lighting, 1 x task light
OFF-CLN | Office - Clinical Workroom | SE | ELGP-208 | GPO: double, wall mounted | 7 | location of outlets to meet local operational/WHS requirements
OFF-CLN | Office - Clinical Workroom | SE | ELGP-208 | GPO: double, wall mounted | 3 | for BYO devices, to be located above desk for easy access
OFF-CLN | Office - Clinical Workroom | SE | ELGP-208 | GPO: double, wall mounted | 1 | optional, for duress charging station if provided
OFF-CLN | Office - Clinical Workroom | SE | ITSE-006 | BUTTON: duress, fixed | 1 | below desk, wall mounted is also acceptable
OFF-CLN | Office - Clinical Workroom | SE | ELGP-130 | GPO: UPS, single | 1 | to printer; UPS to be provided if downtime computer required
OFF-CLN | Office - Clinical Workroom | SE | ELGP-235 | GPO: UPS, double, wall mounted | 2 | UPS to be provided if downtime computer required. location of outlets to meet local operational/WHS requirements
OFF-WI-5 | Office - Write-up, 5m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
OFF-WI-5 | Office - Write-up, 5m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
OFF-WI-5 | Office - Write-up, 5m2 | FA | WLPR-010 | WALL PROTECTION: corner guards, to 900 AFFL | 1 | Align heights of corner guards with other corridor elements
OFF-WI-5 | Office - Write-up, 5m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
OFF-WI-5 | Office - Write-up, 5m2 | FA | CLGE-001 | CEILING: plasterboard, bulkhead, paint | 1 | Between above bench joinery and ceiling
OFF-WI-5 | Office - Write-up, 5m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
OFF-WI-5 | Office - Write-up, 5m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
OFF-WI-5 | Office - Write-up, 5m2 | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1 | Standing or seating height, refer RLS or project specific requirements
OFF-WI-5 | Office - Write-up, 5m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 2 | For standing height bench, use ergonomic high chair for standring height
OFF-WI-5 | Office - Write-up, 5m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
OFF-WI-5 | Office - Write-up, 5m2 | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 4 | Above bench, number to suit layout. Refer to RLS or project specific requirements
OFF-WI-5 | Office - Write-up, 5m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | Between bench and overbench joinery
OFF-WI-5 | Office - Write-up, 5m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 2 | Consider location of harddrive, if below bench add bracket
OFF-WI-5 | Office - Write-up, 5m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2
OFF-WI-5 | Office - Write-up, 5m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | Number of outlets per faceplate as per ICT policy. 1 per workspace
OFF-WI-5 | Office - Write-up, 5m2 | SE | LIFX-008 | LIGHT: task light, built in | 2 | For above bench joinery
OFF-WI-5 | Office - Write-up, 5m2 | SE | ELBO-015 | SWITCH: light, single | 2 | Pelmet lighting
OFF-WI-5 | Office - Write-up, 5m2 | SE | ITCL-001 | ANNUNCIATOR: nurse call | 1 | Located in corridor adjacent, number to suit design
OFF-WI-5 | Office - Write-up, 5m2 | SE | ELGP-208 | GPO: double, wall mounted | 4 | 2 per workspace
OFF-WS | Office - Workstation | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 2
OFF-WS | Office - Workstation | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
OFF-WS | Office - Workstation | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
OFF-WS | Office - Workstation | FA | CLCN-009 | CORNICE: wall trim fixing | 1
OFF-WS | Office - Workstation | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
OFF-WS | Office - Workstation | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | optional; to be rationalised between multiple workstations
OFF-WS | Office - Workstation | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1
OFF-WS | Office - Workstation | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | dual screens may be provided to suit service requirements
OFF-WS | Office - Workstation | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | optional
OFF-WS | Office - Workstation | FFE | FQDW-043 | PARTITION: workstation system, acoustic | 1 | a modesty panel may been to be required depending on layout, this may be an extension of the acoustic panel or a separate panel depending on selected system.
OFF-WS | Office - Workstation | FFE | FQDW-054 | WORKSTATION: straight, adj height, 1600W x 750D x 750H nom. | 1
OFF-WS | Office - Workstation | FFE | FQDW-075 | WORKSTATION SYSTEM: cable tray, soft wired power and data | 1 | separation of power and data cables may be required
OFF-WS | Office - Workstation | SE | ITIN-029 | OUTLET: data, double RJ45, desk box | 1
OFF-WS | Office - Workstation | SE | ELGP-212 | GPO: double, desk box | 3
OFF-WS | Office - Workstation | SE | ITIN-145 | DESK BOX: for power, data, AV, flush mounted, softwired | 1
OVBR | Overnight Stay - Bedroom | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
OVBR | Overnight Stay - Bedroom | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 8
OVBR | Overnight Stay - Bedroom | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1
OVBR | Overnight Stay - Bedroom | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1 | optional
OVBR | Overnight Stay - Bedroom | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
OVBR | Overnight Stay - Bedroom | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
OVBR | Overnight Stay - Bedroom | FA | DWWH-005 | WINDOW | 1 | essential
OVBR | Overnight Stay - Bedroom | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | lockable
OVBR | Overnight Stay - Bedroom | FA | CLCN-009 | CORNICE: wall trim fixing | 1
OVBR | Overnight Stay - Bedroom | FFE | FQBS-018 | CHAIR: office, ergonomic | 1 | optional
OVBR | Overnight Stay - Bedroom | FFE | DWWD-004 | BLIND: roller | 1 | blackout fabric
OVBR | Overnight Stay - Bedroom | FFE | FIJO-130 | CUPBOARD: tall, single door, adj shelves and clothes rail | 1
OVBR | Overnight Stay - Bedroom | FFE | FQCW-016 | BIN: waste, general, 20L | 1
OVBR | Overnight Stay - Bedroom | FFE | FQBS-004 | BED: domestic, single with mattress | 1
OVBR | Overnight Stay - Bedroom | FFE | FQTA-001 | TABLE: bedside | 1
OVBR | Overnight Stay - Bedroom | FFE | FQCA-030 | FRIDGE: domestic, 120L | 1 | optional, dependent on project requirements
OVBR | Overnight Stay - Bedroom | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | may be wall mounted
OVBR | Overnight Stay - Bedroom | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead is also acceptable
OVBR | Overnight Stay - Bedroom | FFE | FQDW-053 | DESK: fixed height, 1000W x 600D x 720H | 1 | optional
OVBR | Overnight Stay - Bedroom | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | 1 x optional, to desk if provided
OVBR | Overnight Stay - Bedroom | SE | LIFX-011 | LIGHT: reading light | 1
OVBR | Overnight Stay - Bedroom | SE | ELBO-015 | SWITCH: light, single | 2 | 1 x general room lighting, 1 x reading light
OVBR | Overnight Stay - Bedroom | SE | ELGP-109 | GPO: single, wall mounted | 1 | optional, to fridge if provided
OVBR | Overnight Stay - Bedroom | SE | ELGP-208 | GPO: double, wall mounted | 1
OVBR | Overnight Stay - Bedroom | SE | ITSE-038 | READER: access control, general | 1 | optional
OVBR | Overnight Stay - Bedroom | SE | ELGP-208.1 | GPO: double, wall mounted (V) | 1
OVES | Overnight Stay - Ensuite | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
OVES | Overnight Stay - Ensuite | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
OVES | Overnight Stay - Ensuite | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
OVES | Overnight Stay - Ensuite | FA | WLWA-019 | WALL FINISH: vinyl, full height | 6 | tiles are also acceptable
OVES | Overnight Stay - Ensuite | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
OVES | Overnight Stay - Ensuite | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | with privacy latch, inward opening door to have emergency release
OVES | Overnight Stay - Ensuite | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
OVES | Overnight Stay - Ensuite | FFE | FIHR-026 | HOOK: coat | 2
OVES | Overnight Stay - Ensuite | FFE | FIRT-034 | TRACK: curtain, shower, 90 deg | 1
OVES | Overnight Stay - Ensuite | FFE | FQGE-011 | CURTAIN: shower | 1
OVES | Overnight Stay - Ensuite | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
OVES | Overnight Stay - Ensuite | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
OVES | Overnight Stay - Ensuite | FFE | FIDI-030 | DISPENSER: soap | 2 | 1 - to basin, 1 - to shower
OVES | Overnight Stay - Ensuite | FFE | FIDI-036 | DISPENSER: toilet paper | 1
OVES | Overnight Stay - Ensuite | FFE | FIBM-006 | MIRROR: fixed, glass | 1
OVES | Overnight Stay - Ensuite | FFE | FQCW-037 | BIN: waste, sanitary | 1
OVES | Overnight Stay - Ensuite | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 1 | optional
OVES | Overnight Stay - Ensuite | FFE | FIRT-016 | RAIL: towel | 1
OVES | Overnight Stay - Ensuite | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1
OVES | Overnight Stay - Ensuite | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 1
OVES | Overnight Stay - Ensuite | FFE | HYTP-018 | SHOWER: fixed rose | 1
OVES | Overnight Stay - Ensuite | FFE | HYWC-008 | TOILET SUITE: standard | 1 | with 100mm clearout access point
OVES | Overnight Stay - Ensuite | FFE | HYBA-002 | BASIN: ensuite, with vanity unit | 1
OVES | Overnight Stay - Ensuite | SE | HYDR-004 | DRAIN: floor waste | 1
OVES | Overnight Stay - Ensuite | SE | ELBO-015 | SWITCH: light, single | 1
OVES | Overnight Stay - Ensuite | SE | ELGP-208 | GPO: double, wall mounted | 1
PAR | Parenting Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
PAR | Parenting Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
PAR | Parenting Room | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | tiles are also acceptable
PAR | Parenting Room | FA | WLWA-007 | WALL FINISH: splashback | 1
PAR | Parenting Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
PAR | Parenting Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 4
PAR | Parenting Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PAR | Parenting Room | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | with privacy latch, emergency release to be provided on inward opening door
PAR | Parenting Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PAR | Parenting Room | FFE | FQBS-035 | CHANGE TABLE: baby | 1
PAR | Parenting Room | FFE | FIHR-026 | HOOK: coat | 2
PAR | Parenting Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
PAR | Parenting Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to sink
PAR | Parenting Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to sink
PAR | Parenting Room | FFE | FQBS-010 | CHAIR: armchair | 1 | to be suitable for breastfeeding
PAR | Parenting Room | FFE | FQCW-019 | BIN: waste, general, 50L | 1 | for nappies, sanitary waste stream may be required dependent on operational policies
PAR | Parenting Room | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
PAR | Parenting Room | FFE | HYBA-051 | SINK: single | 1 | depth to meet accessible requirements; integrated solid surface sink into solid surface bench is also acceptable
PAR | Parenting Room | FFE | FQTA-049 | TABLE: round, 500 diam, low height | 1
PAR | Parenting Room | FFE | FIJO-213 | PANEL: joinery carcass end | 1
PAR | Parenting Room | FFE | FIJO-228 | JOINERY UNIT: sink shroud, accessible | 1 | to sink
PAR | Parenting Room | FFE | FIJO-232 | BENCH TOP: 650D, joinery under, laminate finish | 1 | solid surface is also acceptable, dependent on requirements
PAR | Parenting Room | SE | ELBO-015 | SWITCH: light, single | 1
PAR | Parenting Room | SE | ELGP-208 | GPO: double, wall mounted | 1
PAR | Parenting Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
PBSC | Patient Bay - Specimen Collection | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PBSC | Patient Bay - Specimen Collection | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
PBSC | Patient Bay - Specimen Collection | FA | WLPR-006 | WALL PROTECTION: corner guards | 2
PBSC | Patient Bay - Specimen Collection | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
PBSC | Patient Bay - Specimen Collection | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 3
PBSC | Patient Bay - Specimen Collection | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PBSC | Patient Bay - Specimen Collection | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PBSC | Patient Bay - Specimen Collection | FFE | FIHR-016 | BRACKET: for sharps bin | 1 | on trolley
PBSC | Patient Bay - Specimen Collection | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 1 | optional
PBSC | Patient Bay - Specimen Collection | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PBSC | Patient Bay - Specimen Collection | FFE | FIHR-026 | HOOK: coat | 2 | optional, prioritize lower hook for wheelchair users if only one hook provided
PBSC | Patient Bay - Specimen Collection | FFE | FIRT-031 | TRACK: curtain, privacy | 1
PBSC | Patient Bay - Specimen Collection | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PBSC | Patient Bay - Specimen Collection | FFE | FQSN-007 | BIN: storage bins, wire mesh, stainless steel | 1 | extent and type of storage to be based on service requirements and local infection control policies
PBSC | Patient Bay - Specimen Collection | FFE | FQCW-016 | BIN: waste, general, 20L | 1
PBSC | Patient Bay - Specimen Collection | FFE | FIDI-014 | DISPENSER: disposable glove | 1
PBSC | Patient Bay - Specimen Collection | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1
PBSC | Patient Bay - Specimen Collection | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1 | on trolley; may be wall mounted or mobile
PBSC | Patient Bay - Specimen Collection | FFE | FQBS-051 | STOOL: mobile, adj height | 1
PBSC | Patient Bay - Specimen Collection | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1
PBSC | Patient Bay - Specimen Collection | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | optional
PBSC | Patient Bay - Specimen Collection | FFE | FQGE-056 | PRINTER: desktop, labels | 1 | optional, on WOW
PBSC | Patient Bay - Specimen Collection | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | a desk may be provided depending on operational model
PBSC | Patient Bay - Specimen Collection | FFE | MMGE-189 | TROLLEY: phlebotomy 600W x 490D x 900H nominal | 1 | with drop down shelf
PBSC | Patient Bay - Specimen Collection | FFE | MMBE-094 | CHAIR: recliner, clinical, phlebotomy | 1 | electric, height adjustable and able to lie flat, electric examination couch may also be provided
PBSC | Patient Bay - Specimen Collection | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
PBSC | Patient Bay - Specimen Collection | SE | ELPR-006 | RCD: residual current device | 1
PBSC | Patient Bay - Specimen Collection | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | optional, to exam light
PBSC | Patient Bay - Specimen Collection | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
PBSC | Patient Bay - Specimen Collection | SE | ELBO-015 | SWITCH: light, single | 1 | optional, to exam light
PBSC | Patient Bay - Specimen Collection | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
PBSC | Patient Bay - Specimen Collection | SE | ELGP-109 | GPO: single, wall mounted | 1 | to chair
PBSC | Patient Bay - Specimen Collection | SE | ELGP-208 | GPO: double, wall mounted | 1
PBSC | Patient Bay - Specimen Collection | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
PBSC | Patient Bay - Specimen Collection | SE | ITSE-006 | BUTTON: duress, fixed | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 2
PBTR-MD | Patient Bay - Medical Day Treatment | FA | WLPR-002 | WALL PROTECTION: bed protection wall panel, composite, laminate finish | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FA | DWWH-005 | WINDOW | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | MGFP-002 | FLOWMETER: medical air | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FIHR-019 | BRACKET: for suction bottle | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FIHR-016 | BRACKET: for sharps bin | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FQBS-030 | CHAIR: visitor, patient | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | DWWD-004 | BLIND: roller | 1 | dependent on local infection control policies and service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FQGE-010 | CURTAIN: privacy screen | 2
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | MMGE-119 | TROLLEY: equipment | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FQSN-051 | LOCKER: 2 drawers, open under, lockable, mobile | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FQTA-013 | TABLE: overbed | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | MGFP-004 | FLOWMETER: oxygen | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | FQBS-051 | STOOL: mobile, adj height | 1 | quantity to be rationalised across the unit to suit number of staff
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | MMGE-002 | ADAPTER: suction | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | MMGE-100 | SUCTION BOTTLE | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | MMGE-023 | CHAIR: recliner, clinical, electric | 1 | or bed
PBTR-MD | Patient Bay - Medical Day Treatment | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | optional, dependent on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ELPR-006 | RCD: residual current device | 2
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment screen, may be replaced with data outlet in accordance with ICT policy
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ELGP-101 | GPO: single, ceiling mounted | 2 | to patient entertainment screen
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
PBTR-MD | Patient Bay - Medical Day Treatment | SE | LIFX-011 | LIGHT: reading light | 1 | hardwired
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ELBO-015 | SWITCH: light, single | 1 | to exam light
PBTR-MD | Patient Bay - Medical Day Treatment | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | + indicator button & light
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | call button on the handset, reading light control
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ELGP-109 | GPO: single, wall mounted | 2 | low level, for bed/recliner
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ELGP-206 | GPO: double, on services panel, RCD protected | 1
PBTR-MD | Patient Bay - Medical Day Treatment | SE | MMSP-026 | SERVICES PANEL: wall mounted | 2 | 1 panel for medical gases may be shared between two bays depending on service requirements
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ELGP-105 | GPO: single, on services panel | 2
PBTR-MD | Patient Bay - Medical Day Treatment | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent; a higher slip rating may be required depending on local WHS policies
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | optional, to side walls if provided
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 2 | optional, to side walls if provided
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | DWWH-005 | WINDOW | 1 | access to natural light and external outlook essential
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FIHR-019 | BRACKET: for suction bottle | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FIHR-016 | BRACKET: for sharps bin | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FQBS-030 | CHAIR: visitor, patient | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | DWWD-004 | BLIND: roller | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FQGE-010 | CURTAIN: privacy screen | 2
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FQTA-013 | TABLE: overbed | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | MGFP-004 | FLOWMETER: oxygen | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | MMGE-002 | ADAPTER: suction | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | MMGE-100 | SUCTION BOTTLE | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | MMGE-023 | CHAIR: recliner, clinical, electric | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | MMGE-049 | MACHINE: dialysis | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | MMGE-096 | STATION: combination dialysis connection point and waste water tundish | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | optional
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | FFE | HYDR-019 | ACCESS PANEL: waterless trap | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ELPR-006 | RCD: residual current device | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment system, may be replaced with data outlet depending on ICT policy
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to patient entertainment system
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | HYDR-004 | DRAIN: floor waste | 1 | in case of spills, to be consolidated between multiple bays
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | LIFX-011 | LIGHT: reading light | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ELBO-015 | SWITCH: light, single | 1 | optional
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | MGAS-062 | OUTLET: suction, on services panel | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | call button and reading light switch on handset
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | low level, to recliner
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | SE | ELGP-222 | GPO: emergency power, double, on services panel | 3
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent; a higher slip rating may be required depending on local WHS policies
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | optional, to side walls if provided
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 2 | optional, to side walls if provided
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | DWWH-005 | WINDOW | 1 | access to natural light and external outlook essential
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FIHR-019 | BRACKET: for suction bottle | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FIHR-016 | BRACKET: for sharps bin | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FQBS-030 | CHAIR: visitor, patient | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | DWWD-004 | BLIND: roller | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FQGE-010 | CURTAIN: privacy screen | 2
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FQTA-013 | TABLE: overbed | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | MGFP-004 | FLOWMETER: oxygen | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | MMGE-002 | ADAPTER: suction | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | MMGE-100 | SUCTION BOTTLE | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | MMBE-009 | BED: patient, electric | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | MMGE-049 | MACHINE: dialysis | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | MMGE-096 | STATION: combination dialysis connection point and waste water tundish | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | optional
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | FFE | HYDR-019 | ACCESS PANEL: waterless trap | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ELPR-006 | RCD: residual current device | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment system, may be replaced with data outlet depending on ICT policy
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to patient entertainment system
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | HYDR-004 | DRAIN: floor waste | 1 | in case of spills, to be consolidated between multiple bays
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | LIFX-011 | LIGHT: reading light | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ELBO-015 | SWITCH: light, single | 1 | optional
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | MGAS-062 | OUTLET: suction, on services panel | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | call button and reading light switch on handset
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | low level, to bed
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | SE | ELGP-222 | GPO: emergency power, double, on services panel | 3
PET-CT | PET-CT Imaging Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
PET-CT | PET-CT Imaging Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
PET-CT | PET-CT Imaging Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
PET-CT | PET-CT Imaging Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
PET-CT | PET-CT Imaging Room | FA | WLWA-007 | WALL FINISH: splashback | 1
PET-CT | PET-CT Imaging Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
PET-CT | PET-CT Imaging Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 6
PET-CT | PET-CT Imaging Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PET-CT | PET-CT Imaging Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1 | automated
PET-CT | PET-CT Imaging Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1 | automated
PET-CT | PET-CT Imaging Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PET-CT | PET-CT Imaging Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 1
PET-CT | PET-CT Imaging Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | on services panel
PET-CT | PET-CT Imaging Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2
PET-CT | PET-CT Imaging Room | FFE | FQBS-030 | CHAIR: visitor, patient | 1
PET-CT | PET-CT Imaging Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-141 | CUPBOARD: underbench, single door, adj shelves, lockable | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 3
PET-CT | PET-CT Imaging Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door
PET-CT | PET-CT Imaging Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-045 | IV STAND: mobile | 1 | may be ceiling mounted
PET-CT | PET-CT Imaging Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
PET-CT | PET-CT Imaging Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | on services panel
PET-CT | PET-CT Imaging Room | FFE | MMBE-042 | SLIDE: patient | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-142 | TROLLEY: sharps bin | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-091 | SPHYGMOMANOMETER: portable | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-002 | ADAPTER: suction | 2 | on services panel
PET-CT | PET-CT Imaging Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
PET-CT | PET-CT Imaging Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
PET-CT | PET-CT Imaging Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
PET-CT | PET-CT Imaging Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
PET-CT | PET-CT Imaging Room | FFE | ITCL-012 | CAMERA: observation | 3
PET-CT | PET-CT Imaging Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
PET-CT | PET-CT Imaging Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-207 | DEVICE: imaging, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-209 | PHANTOM: PET/CT | 1 | stored within shielded custom joinery
PET-CT | PET-CT Imaging Room | FFE | MMMI-217 | TABLE: patient, fixed, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-107 | INJECTOR: contrast, CT, ceiling mounted | 1 | also may be mobile or wall mounted
PET-CT | PET-CT Imaging Room | FFE | MMMI-218 | LASER: positioning lights, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-377 | WARMER: contrast | 1 | optional
PET-CT | PET-CT Imaging Room | FFE | FIJO-218 | CUPBOARD: tall, double door, adj shelves, lockable, heavy duty | 1
PET-CT | PET-CT Imaging Room | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 2
PET-CT | PET-CT Imaging Room | FFE | FIHR-078 | BRACKET: for observation camera, wall mounted | 1
PET-CT | PET-CT Imaging Room | FFE | FIHR-092 | HOOK: anchor, patient ladder | 1
PET-CT | PET-CT Imaging Room | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 2
PET-CT | PET-CT Imaging Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3
PET-CT | PET-CT Imaging Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | on wall near PET-CT unit
PET-CT | PET-CT Imaging Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room
PET-CT | PET-CT Imaging Room | SE | ELPR-006 | RCD: residual current device | 1 | on services panel
PET-CT | PET-CT Imaging Room | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 2
PET-CT | PET-CT Imaging Room | SE | ELGP-101 | GPO: single, ceiling mounted | 3 | 1-laser positioning light
PET-CT | PET-CT Imaging Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
PET-CT | PET-CT Imaging Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional, on services panel
PET-CT | PET-CT Imaging Room | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | alarm for temperature range
PET-CT | PET-CT Imaging Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 2 | to outside of all doors to room, auto on when modality is in use, vendor to provide signal wiring to light
PET-CT | PET-CT Imaging Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional, on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
PET-CT | PET-CT Imaging Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel + indicator button & light
PET-CT | PET-CT Imaging Room | SE | ELGP-109 | GPO: single, wall mounted | 2 | 1-clock
PET-CT | PET-CT Imaging Room | SE | ELGP-208 | GPO: double, wall mounted | 3
PET-CT | PET-CT Imaging Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | to walls
PET-CT | PET-CT Imaging Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 1
PET-CT | PET-CT Imaging Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
PET-CT | PET-CT Imaging Room | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
PET-CT | PET-CT Imaging Room | SE | ELGP-203 | GPO: double, in floorbox | 4
PET-CT | PET-CT Imaging Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
PET-CT | PET-CT Imaging Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | patient services
PET-CT | PET-CT Imaging Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to room doors, walls, (ceiling) as required by radiation consultant design and specification
PET-CT | PET-CT Imaging Room | SE | ELBO-012 | SWITCH: light, dimmer | 1 | number determined by number of doors
PET-CT | PET-CT Imaging Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PET-CT | PET-CT Imaging Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4 | on service panel
PET-CT | PET-CT Imaging Room | SE | ELPR-007 | UPS: status indicator | 1 | on service panel
PET-CT | PET-CT Imaging Room | SE | LIFX-038 | LIGHT: backlit feature panels | 2
PHA-CO | Pharmacy - Counter | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1 | vinyl is also acceptable
PHA-CO | Pharmacy - Counter | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 12 | seamless coved vinyl skirting if vinyl flooring provided
PHA-CO | Pharmacy - Counter | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 7
PHA-CO | Pharmacy - Counter | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PHA-CO | Pharmacy - Counter | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PHA-CO | Pharmacy - Counter | FFE | FQSN-028 | CABINET: pass through | 1 | lockable, under bench, for transfer of large items
PHA-CO | Pharmacy - Counter | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
PHA-CO | Pharmacy - Counter | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
PHA-CO | Pharmacy - Counter | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 2
PHA-CO | Pharmacy - Counter | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | quantity and type of bins to be based on local waste management policies
PHA-CO | Pharmacy - Counter | FFE | FIJO-114 | CUPBOARD: drawer unit, lockable | 1
PHA-CO | Pharmacy - Counter | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 2
PHA-CO | Pharmacy - Counter | FFE | ITNE-023 | PRINTER: networked | 1 | optional
PHA-CO | Pharmacy - Counter | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2
PHA-CO | Pharmacy - Counter | FFE | FIJO-058 | BENCH TOP: 600D, joinery under | 2
PHA-CO | Pharmacy - Counter | FFE | FIJO-051 | BENCH TOP: 450D, joinery under | 1
PHA-CO | Pharmacy - Counter | FFE | FIJO-099 | CAC: cable access cap | 2
PHA-CO | Pharmacy - Counter | FFE | FIJO-213 | PANEL: joinery carcass end | 1
PHA-CO | Pharmacy - Counter | FFE | FQCW-064 | BIN: waste, paper, secure, 20L | 1 | quantity and type of bins to be based on local waste management policies
PHA-CO | Pharmacy - Counter | FFE | FIJO-229 | JOINERY UNIT: reception counter, pharmacy | 1
PHA-CO | Pharmacy - Counter | FFE | FIJO-230 | JOINERY UNIT: reception counter, pharmacy, accessible | 1
PHA-CO | Pharmacy - Counter | FFE | WLPT-015 | SCREEN: security, safety glass, with communication slot and pass through | 2
PHA-CO | Pharmacy - Counter | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3
PHA-CO | Pharmacy - Counter | SE | ELBO-015 | SWITCH: light, single | 1
PHA-CO | Pharmacy - Counter | SE | ELGP-208 | GPO: double, wall mounted | 3
PHA-CO | Pharmacy - Counter | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | extent of emergency power to be determined based on service requirements
PHA-CO | Pharmacy - Counter | SE | ITSE-006 | BUTTON: duress, fixed | 1 | shared, below desk
PHA-DB | Pharmacy - Distribution Workstation | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PHA-DB | Pharmacy - Distribution Workstation | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
PHA-DB | Pharmacy - Distribution Workstation | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PHA-DB | Pharmacy - Distribution Workstation | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PHA-DB | Pharmacy - Distribution Workstation | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PHA-DB | Pharmacy - Distribution Workstation | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
PHA-DB | Pharmacy - Distribution Workstation | FFE | FQDW-002 | DESK: adj height, 1800W x 750D x 750H nom. | 1
PHA-DB | Pharmacy - Distribution Workstation | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
PHA-DB | Pharmacy - Distribution Workstation | FFE | FIJO-108 | CUPBOARD: drawer unit, baskets, 3 x Wire (400mm wide) | 1
PHA-DB | Pharmacy - Distribution Workstation | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | may be on wall or desk mounted bracket, an additional screen may be provided depending on project ICT requirements
PHA-DB | Pharmacy - Distribution Workstation | FFE | ITNE-003 | BAR CODING: reader | 1
PHA-DB | Pharmacy - Distribution Workstation | FFE | FIJO-069 | BENCH TOP: 750D, joinery under, laminate finish | 1
PHA-DB | Pharmacy - Distribution Workstation | FFE | ITNE-072 | BAR CODING: PDE dock | 1 | may be shared between multiple workstations
PHA-DB | Pharmacy - Distribution Workstation | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
PHA-DB | Pharmacy - Distribution Workstation | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | extent of emergency power to be determined based on service requirements
PHA-DS | Pharmacy - Dispensing Workstation | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PHA-DS | Pharmacy - Dispensing Workstation | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PHA-DS | Pharmacy - Dispensing Workstation | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PHA-DS | Pharmacy - Dispensing Workstation | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
PHA-DS | Pharmacy - Dispensing Workstation | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | quantity and type of bins to be based on local waste management policies
PHA-DS | Pharmacy - Dispensing Workstation | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 1 | optional, depending on project ICT requirements; may be on desk mounted bracket
PHA-DS | Pharmacy - Dispensing Workstation | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | may be on desk mounted bracket
PHA-DS | Pharmacy - Dispensing Workstation | FFE | ITNE-003 | BAR CODING: reader | 1
PHA-DS | Pharmacy - Dispensing Workstation | FFE | FQDW-046 | WORKSTATION: pharmacy,straight, adj height, 1800W x 750D x 900H nom. | 1
PHA-DS | Pharmacy - Dispensing Workstation | FFE | FQDW-047 | WORKSTATION: pharmacy, services spine bench | 1 | services spine may be shared between multiple workstations
PHA-DS | Pharmacy - Dispensing Workstation | FFE | FQDW-048 | WORKSTATION: pharmacy, services spine shelves | 1 | services spine may be shared between multiple workstations
PHA-DS | Pharmacy - Dispensing Workstation | FFE | FQCW-064 | BIN: waste, paper, secure, 20L | 1 | quantity and type of bins to be based on local waste management policies
PHA-DS | Pharmacy - Dispensing Workstation | FFE | FQGE-093 | PRINTER: desktop, labels, pharmacy | 1
PHA-DS | Pharmacy - Dispensing Workstation | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2
PHA-DS | Pharmacy - Dispensing Workstation | SE | ELGP-208 | GPO: double, wall mounted | 1
PHA-DS | Pharmacy - Dispensing Workstation | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | extent of emergency power to be determined based on service requirements
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent to extend under basin
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 11
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 6
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 4
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | DWGL-003 | DOOR: 910mm c/o, single leaf, half glazed, solid core/safety glass, paint | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1 | may be height adjustable item depending on service requirements
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIHR-016 | BRACKET: for sharps bin | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIJO-141 | CUPBOARD: underbench, single door, adj shelves, lockable | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 3
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIJO-148 | CUPBOARD: wall mount, double door, glazed, adj shelves, lockable | 3
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIJO-159 | CUPBOARD: wall mount, single door, glazed, adj shelves, lockable | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FQSM-016 | TROLLEY | 1 | optional
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FQCW-019 | BIN: waste, general, 50L | 1 | quantity and type of bins to be based on local waste management policies
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | HYBA-010 | BASIN: handwash, Type B | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | HYBA-069 | SINK: single, with drainer, 930mmL | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | MEEX-013 | FUME CUPBOARD | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 2
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FIJO-212 | BULKHEAD: joinery | 2
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | MMGE-185 | SCALES: digital, benchtop | 2
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | FFE | FQCW-065 | BIN: waste, pharmaceutical 50L | 1 | quantity and type of bins to be based on local waste management policies
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | SE | LIFX-008 | LIGHT: task light, built in | 2 | under wall mounted joinery
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 2 x task lights
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | SE | ELGP-208 | GPO: double, wall mounted | 4 | provision of emergency power to be determined based on service requirements
PHA-RE | Pharmacy - Returns Workstation | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PHA-RE | Pharmacy - Returns Workstation | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
PHA-RE | Pharmacy - Returns Workstation | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PHA-RE | Pharmacy - Returns Workstation | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PHA-RE | Pharmacy - Returns Workstation | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PHA-RE | Pharmacy - Returns Workstation | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
PHA-RE | Pharmacy - Returns Workstation | FFE | FQDW-002 | DESK: adj height, 1800W x 750D x 750H nom. | 1
PHA-RE | Pharmacy - Returns Workstation | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
PHA-RE | Pharmacy - Returns Workstation | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | may be on wall or desk mounted bracket, an additional screen may be provided depending on project ICT requirements
PHA-RE | Pharmacy - Returns Workstation | FFE | ITNE-003 | BAR CODING: reader | 1
PHA-RE | Pharmacy - Returns Workstation | FFE | FQSM-045 | SHELVING: storage solution, 400W x 600D x 2100H nom., wire, mobile | 2
PHA-RE | Pharmacy - Returns Workstation | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
PHA-RE | Pharmacy - Returns Workstation | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | extent of emergency power to be determined based on service requirements
PLAY | Play Area - Paediatric | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PLAY | Play Area - Paediatric | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
PLAY | Play Area - Paediatric | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
PLAY | Play Area - Paediatric | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PLAY | Play Area - Paediatric | FA | DWWH-005 | WINDOW | 1 | desirable
PLAY | Play Area - Paediatric | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PLAY | Play Area - Paediatric | FFE | FQGE-002 | ARTWORK | 1
PLAY | Play Area - Paediatric | FFE | FQBS-012 | CHAIR: child | 2
PLAY | Play Area - Paediatric | FFE | FQBS-032 | CHAIR: waiting | 2
PLAY | Play Area - Paediatric | FFE | DWWD-004 | BLIND: roller | 1
PLAY | Play Area - Paediatric | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PLAY | Play Area - Paediatric | FFE | FQTA-003 | TABLE: children's play | 1
PLAY | Play Area - Paediatric | FFE | FQTA-075 | TABLE: square, 600W, low height | 1
PLAY | Play Area - Paediatric | SE | ELGP-208 | GPO: double, wall mounted | 1
PLNT-WT | Plant - Water Treatment | FA | FLGE-003 | FLOOR FINISH: concrete, sealed, trowel | 1
PLNT-WT | Plant - Water Treatment | FA | FLSK-002 | SKIRTING: concrete, coved to 150 AFFL, sealed | 6
PLNT-WT | Plant - Water Treatment | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
PLNT-WT | Plant - Water Treatment | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 4
PLNT-WT | Plant - Water Treatment | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
PLNT-WT | Plant - Water Treatment | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1
PLNT-WT | Plant - Water Treatment | FA | DWSC-027 | DOOR: 1600mm c/o, 1 1/2 leaf, solid core, paint | 1 | lockable
PLNT-WT | Plant - Water Treatment | FA | DWSC-027 | DOOR: 1600mm c/o, 1 1/2 leaf, solid core, paint | 1 | optional secondary access door, lockable
PLNT-WT | Plant - Water Treatment | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PLNT-WT | Plant - Water Treatment | FFE | FQBS-051 | STOOL: mobile, adj height | 1 | optional
PLNT-WT | Plant - Water Treatment | FFE | MEGE-010 | HEAT EXCHANGER: supplied by mechanical chilled water | 1 | as required; may be necessary to cool feed water where high temperatures are expected, refer to technical expert for required outlet temperature (generally <25°C)
PLNT-WT | Plant - Water Treatment | FFE | HYGE-016 | PUMP: pressure booster | 2 | to suit water pressure drop through pre-filtration train where building water pressure and flow is inadequate
PLNT-WT | Plant - Water Treatment | FFE | HYGE-012 | WATER TREATMENT: reverse osmosis plant | 1 | details according to manufacturer's specifications
PLNT-WT | Plant - Water Treatment | FFE | FQSN-138 | PALLET: half size | 1
PLNT-WT | Plant - Water Treatment | FFE | HYGE-027 | DEVICE: UV irradiator | 1
PLNT-WT | Plant - Water Treatment | FFE | FQDW-051 | DESK: adj height, 1500W x 600D x 900H nom. | 1 | optional
PLNT-WT | Plant - Water Treatment | SE | ITBU-002 | ALARM: equipment monitoring, water treatment malfunction | 1 | connected to BMS and staff station in unit
PLNT-WT | Plant - Water Treatment | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
PLNT-WT | Plant - Water Treatment | SE | HYTP-008 | OUTLET: cold water | 1
PLNT-WT | Plant - Water Treatment | SE | HYDR-004 | DRAIN: floor waste | 1 | floor waste must be capable of handling high flow rates and high temperature water (up to 90°C) from the RO water plant
PLNT-WT | Plant - Water Treatment | SE | HYDR-006 | DRAIN: tundish | 2
PLNT-WT | Plant - Water Treatment | SE | ELBO-015 | SWITCH: light, single | 1
PLNT-WT | Plant - Water Treatment | SE | ELBO-010 | SWITCH: light, 2-way | 1 | optional
PLNT-WT | Plant - Water Treatment | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 8 | for water treatment equipment, high level (power tools to be on a separate circuit)
PLNT-WT | Plant - Water Treatment | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | optional, dependent on equipment
PLNT-WT | Plant - Water Treatment | SE | HYGE-020 | WATER TREATMENT: carbon filter | 3 | 1 is optional; size/quantity based on feed water quality and confirmed by a dialysis water pre-treatment expert
PLNT-WT | Plant - Water Treatment | SE | HYGE-019 | WATER TREATMENT: brine tank | 1 | size/quantity based on feed water quality and confirmed by a dialysis water pre-treatment expert
PLNT-WT | Plant - Water Treatment | SE | HYGE-024 | WATER TREATMENT: softener tank | 1 | size/quantity based on feed water quality and confirmed by a dialysis water pre-treatment expert
PLNT-WT | Plant - Water Treatment | SE | HYGE-021 | WATER TREATMENT: multimedia tank | 1 | size/quantity based on feed water quality and confirmed by a dialysis water pre-treatment expert
PLNT-WT | Plant - Water Treatment | SE | HYGE-023 | WATER TREATMENT: particle filter, 5 microns | 1 | size/quantity based on feed water quality and confirmed by a dialysis water pre-treatment expert
PLNT-WT | Plant - Water Treatment | SE | HYGE-022 | WATER TREATMENT: particle filter, 1 micron | 1 | size/quantity based on feed water quality and confirmed by a dialysis water pre-treatment expert
PLNT-WT | Plant - Water Treatment | SE | HYDR-017 | DRAIN: floor waste, 300 x 300 | 1 | floor waste must be capable of handling high flow rates and high temperature water (up to 90°C) from the RO water plant
PLNT-WT | Plant - Water Treatment | SE | ITBU-024 | ALARM: chlorine meter | 1 | optional, dependent on local operational policies
PLNT-WT | Plant - Water Treatment | SE | HYGE-028 | DEVICE: backflow prevention | 1
PLNT-WT | Plant - Water Treatment | SE | HYDR-018 | DRAIN: tundish, floor drain | 1 | tundish must be capable of handling high flow rates and high temperature water (up to 90°C) from the RO water plant
PLST | Plaster Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
PLST | Plaster Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
PLST | Plaster Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
PLST | Plaster Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
PLST | Plaster Room | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
PLST | Plaster Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PLST | Plaster Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1
PLST | Plaster Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PLST | Plaster Room | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
PLST | Plaster Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PLST | Plaster Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
PLST | Plaster Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | quantity to suit number of suction outlets
PLST | Plaster Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PLST | Plaster Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 3
PLST | Plaster Room | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 3
PLST | Plaster Room | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PLST | Plaster Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PLST | Plaster Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
PLST | Plaster Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PLST | Plaster Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PLST | Plaster Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PLST | Plaster Room | FFE | MMBE-043 | TROLLEY: patient, examination | 1 | or plinth/examination couch
PLST | Plaster Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | quantity to suit number of oxygen outlets
PLST | Plaster Room | FFE | MMMA-011 | MIXER: oxygen and nitrous oxide | 1 | optional, may be mobile, on trolley or wall mounted
PLST | Plaster Room | FFE | MMAH-028 | TROLLEY: plaster | 1
PLST | Plaster Room | FFE | MMGE-002 | ADAPTER: suction | 2 | quantity to suit number of suction outlets
PLST | Plaster Room | FFE | MMGE-100 | SUCTION BOTTLE | 2 | quantity to suit number of suction outlets
PLST | Plaster Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
PLST | Plaster Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
PLST | Plaster Room | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
PLST | Plaster Room | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
PLST | Plaster Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
PLST | Plaster Room | FFE | MMGE-031 | DIAGNOSTIC SET: wall mounted | 1 | optional, may be portable/mobile on stand
PLST | Plaster Room | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
PLST | Plaster Room | FFE | HYDR-008 | TRAP: plaster | 1 | to be easily accessible for cleaning and maintenance
PLST | Plaster Room | FFE | MMAH-023 | SAW: plaster | 1 | integral/localised dust extraction required
PLST | Plaster Room | FFE | HYBA-053 | SINK: single, 450 x 350 x 300D | 1 | hot water
PLST | Plaster Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
PLST | Plaster Room | FFE | FIJO-060 | BENCH TOP: 600D, joinery under, stainless steel lipped | 1 | with sink and wet edge
PLST | Plaster Room | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1 | optional; type, size and quantity of bins to be determined based on service requirements and waste management policies
PLST | Plaster Room | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
PLST | Plaster Room | FFE | MMGE-001 | ADAPTER: scavenge | 1
PLST | Plaster Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
PLST | Plaster Room | FFE | FIRT-039 | TRACK: curtain, privacy, Z shaped, angled | 1
PLST | Plaster Room | FFE | MMAH-045 | VACUUM: plaster saw | 1 | or independent dust extraction system
PLST | Plaster Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
PLST | Plaster Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
PLST | Plaster Room | SE | ELPR-006 | RCD: residual current device | 4
PLST | Plaster Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to exam light
PLST | Plaster Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
PLST | Plaster Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1
PLST | Plaster Room | SE | LIFX-008 | LIGHT: task light, built in | 3
PLST | Plaster Room | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 1 x task light, 1 x examination light
PLST | Plaster Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | 1 is optional, quantity dependent on service requirements
PLST | Plaster Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1
PLST | Plaster Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
PLST | Plaster Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | 1 is optional, quantity dependent on service requirements
PLST | Plaster Room | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | + indicator button & light
PLST | Plaster Room | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | call button on the handset
PLST | Plaster Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
PLST | Plaster Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level, for patient trolley
PLST | Plaster Room | SE | ELGP-208 | GPO: double, wall mounted | 4
PLST | Plaster Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
PLST | Plaster Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
PLST | Plaster Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PLST | Plaster Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PLST | Plaster Room | SE | ELGP-105 | GPO: single, on services panel | 3
PLST | Plaster Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3
PROC | Procedure Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
PROC | Procedure Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
PROC | Procedure Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
PROC | Procedure Room | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
PROC | Procedure Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
PROC | Procedure Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 4
PROC | Procedure Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PROC | Procedure Room | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 2
PROC | Procedure Room | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
PROC | Procedure Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PROC | Procedure Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
PROC | Procedure Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | quantity to suit number of suction outlets
PROC | Procedure Room | FFE | FIHR-016 | BRACKET: for sharps bin | 1
PROC | Procedure Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PROC | Procedure Room | FFE | FIHR-026 | HOOK: coat | 2 | optional
PROC | Procedure Room | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
PROC | Procedure Room | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1
PROC | Procedure Room | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PROC | Procedure Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PROC | Procedure Room | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
PROC | Procedure Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
PROC | Procedure Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PROC | Procedure Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PROC | Procedure Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PROC | Procedure Room | FFE | MMGE-119 | TROLLEY: equipment | 1
PROC | Procedure Room | FFE | MMBE-043 | TROLLEY: patient, examination | 1
PROC | Procedure Room | FFE | FIHR-011 | BRACKET: for monitor, wall mounted | 1 | optional, for cardiac/vital signs monitor
PROC | Procedure Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | quantity to suit number of oxygen outlets
PROC | Procedure Room | FFE | MMMA-011 | MIXER: oxygen and nitrous oxide | 1 | optional, may be mobile, on trolley or wall mounted
PROC | Procedure Room | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
PROC | Procedure Room | FFE | FQBS-052 | STOOL: mobile, high, adj height | 1
PROC | Procedure Room | FFE | MMGE-002 | ADAPTER: suction | 2 | quantity to suit number of suction outlets
PROC | Procedure Room | FFE | MMGE-100 | SUCTION BOTTLE | 2 | quantity to suit number of suction outlets
PROC | Procedure Room | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
PROC | Procedure Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
PROC | Procedure Room | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
PROC | Procedure Room | FFE | LIFX-013 | LIGHT: procedure, ceiling mounted, small | 1
PROC | Procedure Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
PROC | Procedure Room | FFE | MMGE-031 | DIAGNOSTIC SET: wall mounted | 1 | optional, may be portable/mobile on stand
PROC | Procedure Room | FFE | MMHA-008 | MONITOR: patient, cardiac with vital signs | 1 | optional
PROC | Procedure Room | FFE | ITNE-023 | PRINTER: networked | 1 | optional
PROC | Procedure Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
PROC | Procedure Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
PROC | Procedure Room | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1 | optional; type, size and quantity of bins to be determined based on service requirements and waste management policies
PROC | Procedure Room | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
PROC | Procedure Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
PROC | Procedure Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 to telephone, 1 - optional, for cardiac/vital signs monitor
PROC | Procedure Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2
PROC | Procedure Room | SE | ELPR-006 | RCD: residual current device | 4
PROC | Procedure Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to procedure light
PROC | Procedure Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
PROC | Procedure Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional, dependent on service requirements, provide if nitrous oxide required
PROC | Procedure Room | SE | LIFX-008 | LIGHT: task light, built in | 1
PROC | Procedure Room | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 1 x task light, 1 x examination light
PROC | Procedure Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | 1 is optional, quantity dependent on service requirements
PROC | Procedure Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional, dependent on service requirements
PROC | Procedure Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
PROC | Procedure Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | 1 is optional, quantity dependent on service requirements
PROC | Procedure Room | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | + indicator button & light
PROC | Procedure Room | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | call button on the handset
PROC | Procedure Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
PROC | Procedure Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level, for patient trolley
PROC | Procedure Room | SE | ELGP-208 | GPO: double, wall mounted | 2
PROC | Procedure Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
PROC | Procedure Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | optional, for cardiac/vital signs monitor
PROC | Procedure Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
PROC | Procedure Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PROC | Procedure Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PROC | Procedure Room | SE | ELGP-105 | GPO: single, on services panel | 3
PROC | Procedure Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3
PT-HOLD-A | Patient Bay - Holding, Chair | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | acoustic tiles are also acceptable
PT-HOLD-A | Patient Bay - Holding, Chair | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | MGFP-002 | FLOWMETER: medical air | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FIHR-019 | BRACKET: for suction bottle | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FQBS-030 | CHAIR: visitor, patient | 1 | optional
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FIDI-014 | DISPENSER: disposable glove | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1 | type and extent of storage dependant on operational policies for stock centralisation/decanting and infection prevention and control policies
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | MGFP-004 | FLOWMETER: oxygen | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FQSN-006 | BIN: storage bins, for wall mounted panel | 4 | type and size of bins dependant on items to be stored
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | MMGE-002 | ADAPTER: suction | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | MMGE-100 | SUCTION BOTTLE | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | shared between multiple bays
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | shared between multiple bays
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 1
PT-HOLD-A | Patient Bay - Holding, Chair | FFE | MMBE-091 | CHAIR: recliner, clinical, electric | 1
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | optional, for patient monitoring
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
PT-HOLD-A | Patient Bay - Holding, Chair | SE | LIFX-032 | LIGHT: downlight, examination, ceiling mounted | 1 | optional, provision dependant on service requirements; articulated examination light is also acceptable
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel, with indicator button and light
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ELBO-015 | SWITCH: light, single | 1 | optional, to services panel, for examination downlight if provided
PT-HOLD-A | Patient Bay - Holding, Chair | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
PT-HOLD-A | Patient Bay - Holding, Chair | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
PT-HOLD-A | Patient Bay - Holding, Chair | SE | MGAS-062 | OUTLET: suction, on services panel | 1
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on services panel, with indicator button and light
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on services panel
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ELGP-109 | GPO: single, wall mounted | 1 | low, for patient recliner
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ELGP-208 | GPO: double, wall mounted | 1 | for patient's own devices
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | optional, for patient monitoring
PT-HOLD-A | Patient Bay - Holding, Chair | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | quantity to suit service requirements
PT-HOLD-A | Patient Bay - Holding, Chair | SE | ELGP-122 | GPO: emergency power, single, on services panel | 5 | type and quantity to suit service requirements
PT-HOLD-B | Patient Bay - Holding, Bed | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | acoustic tiles are also acceptable
PT-HOLD-B | Patient Bay - Holding, Bed | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | MGFP-002 | FLOWMETER: medical air | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FIHR-019 | BRACKET: for suction bottle | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FQBS-030 | CHAIR: visitor, patient | 1 | optional
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FIDI-014 | DISPENSER: disposable glove | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1 | type and extent of storage dependant on operational policies for stock centralisation/decanting and infection prevention and control policies
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | MGFP-004 | FLOWMETER: oxygen | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FQSN-006 | BIN: storage bins, for wall mounted panel | 4 | type and size of bins dependant on items to be stored
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | MMGE-002 | ADAPTER: suction | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | MMGE-100 | SUCTION BOTTLE | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | MMBE-007 | BED: patient, day procedure, electric | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | shared between multiple bays
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | shared between multiple bays
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FIHR-068 | BRACKET: for alcohol based hand rub, attached to bed | 1
PT-HOLD-B | Patient Bay - Holding, Bed | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 1
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | optional, for patient monitoring
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ELPR-006 | RCD: residual current device | 2 | to services panel
PT-HOLD-B | Patient Bay - Holding, Bed | SE | LIFX-032 | LIGHT: downlight, examination, ceiling mounted | 1 | optional, provision dependant on service requirements; articulated examination light is also acceptable
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel, with indicator button and light
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ELBO-015 | SWITCH: light, single | 1 | optional, to services panel, for examination downlight if provided
PT-HOLD-B | Patient Bay - Holding, Bed | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
PT-HOLD-B | Patient Bay - Holding, Bed | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
PT-HOLD-B | Patient Bay - Holding, Bed | SE | MGAS-062 | OUTLET: suction, on services panel | 1
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on services panel, with indicator button and light
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on services panel
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ELGP-109 | GPO: single, wall mounted | 1 | low, for patient bed
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ELGP-208 | GPO: double, wall mounted | 1 | for patient's own devices
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | optional, for patient monitoring
PT-HOLD-B | Patient Bay - Holding, Bed | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | quantity to suit service requirements
PT-HOLD-B | Patient Bay - Holding, Bed | SE | ELGP-122 | GPO: emergency power, single, on services panel | 5 | type and quantity to suit service requirements
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MGFP-002 | FLOWMETER: medical air | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-019 | BRACKET: for suction bottle | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-029 | HOOK: equipment | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-119 | TROLLEY: equipment | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-045 | IV STAND: mobile | 1 | bed mounted IV pole or ceiling mounted IV track are also acceptable
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1 | type and extent of storage dependant on operational policies for stock centralisation/decanting and infection prevention and control policies
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MGFP-004 | FLOWMETER: oxygen | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-142 | TROLLEY: sharps bin | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FQSN-006 | BIN: storage bins, for wall mounted panel | 4 | type and size of bins dependant on items to be stored
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-002 | ADAPTER: suction | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-100 | SUCTION BOTTLE | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMBE-007 | BED: patient, day procedure, electric | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMHA-008 | MONITOR: patient, cardiac with vital signs | 1 | may be located on left or right side of bedhead depending on monitor functionality and staff workflows; appropriate wall support required for mounting bracket
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1 | may be located on left or right side of bedhead depending on monitor functionality and staff workflows; appropriate wall support required
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-068 | BRACKET: for alcohol based hand rub, attached to bed | 1 | may be wall mounted depending on bedhead arrangment
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to patient monitor
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | LIFX-032 | LIGHT: downlight, examination, ceiling mounted | 1 | articulated examination light is also acceptable
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel, with indicator button and light
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELBO-015 | SWITCH: light, single | 1 | on services panel, for examination downlight
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MGAS-062 | OUTLET: suction, on services panel | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on services panel, with indicator button and light
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on services panel
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELGP-208 | GPO: double, wall mounted | 1 | low, for patient bed and pressure mattress
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | to patient monitor
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | quantity to suit service requirements
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELGP-122 | GPO: emergency power, single, on services panel | 7 | type and quantity to suit service requirements
PTRY | Pantry | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
PTRY | Pantry | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
PTRY | Pantry | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | floor finish must be minimum slip rating R10/pendulum P3 or agreed equivalent
PTRY | Pantry | FA | WLWA-007 | WALL FINISH: splashback | 4 | to bench with sink, to basin
PTRY | Pantry | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PTRY | Pantry | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
PTRY | Pantry | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | with hold open device to support trolley movement as required
PTRY | Pantry | FA | CLCN-009 | CORNICE: wall trim fixing | 1
PTRY | Pantry | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
PTRY | Pantry | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 1
PTRY | Pantry | FFE | FIJO-156 | CUPBOARD: wall mount, single door, adj shelves | 1
PTRY | Pantry | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PTRY | Pantry | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
PTRY | Pantry | FFE | FIDI-025 | DISPENSER: paper/hand towel | 2 | 1 to basin, 1 to sink
PTRY | Pantry | FFE | FIDI-030 | DISPENSER: soap | 2 | 1 to basin, 1 to sink
PTRY | Pantry | FFE | FIJO-104 | CUPBOARD: drawer unit, 3 equal | 1
PTRY | Pantry | FFE | FQCW-019 | BIN: waste, general, 50L | 1 | size, type and quantity of bins dependent on local waste management policies
PTRY | Pantry | FFE | HYBA-011 | BASIN: handwash, Type C | 1
PTRY | Pantry | FFE | HYGE-003 | BOILING/CHILLED WATER UNIT: Combination | 1 | ensure adequate ventilation in joinery
PTRY | Pantry | FFE | HYTP-029 | TAPWARE: boiling/chilled water unit (comb) | 1 | instant hot/cold water tap to sink; high/gooseneck tap may be required where large kettles are to be filled for tea service
PTRY | Pantry | FFE | HYBA-069 | SINK: single, with drainer, 930mmL | 1
PTRY | Pantry | FFE | FQCA-015 | DISPENSER: ice, potable | 1 | optional; item to comply with infection control requirements
PTRY | Pantry | FFE | FQCA-067 | OVEN: microwave, domestic | 1
PTRY | Pantry | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
PTRY | Pantry | FFE | FQCA-027 | FRIDGE: commercial, floor standing, single door | 1 | Clear area for fridge to allow for adequate ventilation and full opening of door. May be connected to BMS for temperature monitoring depending on service requirements,
PTRY | Pantry | FFE | FIJO-069 | BENCH TOP: 750D, joinery under, laminate finish | 1 | minimum depth 600mm; where an ice machine is provided the depth may need to increase to accommodate appliance
PTRY | Pantry | FFE | FIJO-163 | CUPBOARD: tall, single microwave | 1
PTRY | Pantry | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead is also acceptable
PTRY | Pantry | FFE | FQCW-062 | BIN: waste, recycling, comingled, 50L | 1 | size, type and quantity of bins dependent on local waste management policies
PTRY | Pantry | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | optional, to fridge, for temperature monitoring as required
PTRY | Pantry | SE | HYTP-008 | OUTLET: cold water | 1 | to boiling/ chilled water unit
PTRY | Pantry | SE | ELBO-015 | SWITCH: light, single | 1
PTRY | Pantry | SE | ELGP-109 | GPO: single, wall mounted | 3 | 1 to fridge, 1 to microwave, 1 to boiling/ chilled water unit
PTRY | Pantry | SE | ELGP-208 | GPO: double, wall mounted | 2 | over bench
PTRY | Pantry | SE | ITSE-038 | READER: access control, general | 1 | optional, dependent on security requirements and location of room
RECP-10 | Reception, 10m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | carpet tiles are also acceptable
RECP-10 | Reception, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 18
RECP-10 | Reception, 10m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
RECP-10 | Reception, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
RECP-10 | Reception, 10m2 | FA | CLGE-001 | CEILING: plasterboard, bulkhead, paint | 3
RECP-10 | Reception, 10m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
RECP-10 | Reception, 10m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
RECP-10 | Reception, 10m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
RECP-10 | Reception, 10m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
RECP-10 | Reception, 10m2 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
RECP-10 | Reception, 10m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
RECP-10 | Reception, 10m2 | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 1
RECP-10 | Reception, 10m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1
RECP-10 | Reception, 10m2 | FFE | FIJO-182 | JOINERY UNIT: reception desk | 1
RECP-10 | Reception, 10m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
RECP-10 | Reception, 10m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | dual screens may be provided to suit service requirements
RECP-10 | Reception, 10m2 | FFE | ITNE-023 | PRINTER: networked | 1 | optional
RECP-10 | Reception, 10m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
RECP-10 | Reception, 10m2 | FFE | FIJO-051 | BENCH TOP: 450D, joinery under | 1
RECP-10 | Reception, 10m2 | FFE | FIJO-099 | CAC: cable access cap | 1
RECP-10 | Reception, 10m2 | FFE | WLPT-014 | SCREEN: acoustic, privacy, safety glass | 3 | optional
RECP-10 | Reception, 10m2 | FFE | WLPT-015 | SCREEN: security, safety glass, with communication slot and pass through | 2 | optional
RECP-10 | Reception, 10m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
RECP-10 | Reception, 10m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | location of outlets to meet local operational/WHS requirements
RECP-10 | Reception, 10m2 | SE | ELBO-003 | CONTROL PANEL: lighting | 1 | optional
RECP-10 | Reception, 10m2 | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | optional
RECP-10 | Reception, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
RECP-10 | Reception, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 3 | 1 x for BYO devices, to be located above desk for easy access, location of outlets to meet local operational/WHS requirements
RECP-10 | Reception, 10m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | extent of emergency power to be determined based on service requirements, location of outlets to meet local operational/WHS requirements
RECP-10 | Reception, 10m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1
RECP-10 | Reception, 10m2 | SE | ITSE-038 | READER: access control, general | 1 | optional
RECP-10 | Reception, 10m2 | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1 | optional, to external entry and/or unit entry as required
RECP-15 | Reception, 15m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | carpet tiles are also acceptable
RECP-15 | Reception, 15m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 18
RECP-15 | Reception, 15m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
RECP-15 | Reception, 15m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
RECP-15 | Reception, 15m2 | FA | CLGE-001 | CEILING: plasterboard, bulkhead, paint | 3
RECP-15 | Reception, 15m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
RECP-15 | Reception, 15m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
RECP-15 | Reception, 15m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
RECP-15 | Reception, 15m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 2
RECP-15 | Reception, 15m2 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 2
RECP-15 | Reception, 15m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
RECP-15 | Reception, 15m2 | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 2
RECP-15 | Reception, 15m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 2
RECP-15 | Reception, 15m2 | FFE | FIJO-182 | JOINERY UNIT: reception desk | 1
RECP-15 | Reception, 15m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
RECP-15 | Reception, 15m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 2 | dual screens may be provided to suit service requirements
RECP-15 | Reception, 15m2 | FFE | ITNE-023 | PRINTER: networked | 1 | optional
RECP-15 | Reception, 15m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2
RECP-15 | Reception, 15m2 | FFE | FIJO-051 | BENCH TOP: 450D, joinery under | 1
RECP-15 | Reception, 15m2 | FFE | FIJO-099 | CAC: cable access cap | 2
RECP-15 | Reception, 15m2 | FFE | WLPT-014 | SCREEN: acoustic, privacy, safety glass | 3 | optional
RECP-15 | Reception, 15m2 | FFE | WLPT-015 | SCREEN: security, safety glass, with communication slot and pass through | 3 | optional
RECP-15 | Reception, 15m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
RECP-15 | Reception, 15m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 2 | location of outlets to meet local operational/WHS requirements
RECP-15 | Reception, 15m2 | SE | ELBO-003 | CONTROL PANEL: lighting | 1 | optional
RECP-15 | Reception, 15m2 | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | optional
RECP-15 | Reception, 15m2 | SE | ELBO-015 | SWITCH: light, single | 1
RECP-15 | Reception, 15m2 | SE | ELGP-208 | GPO: double, wall mounted | 5 | 2 x for BYO devices, to be located above desk for easy access, location of outlets to meet local operational/WHS requirements
RECP-15 | Reception, 15m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | extent of emergency power to be determined based on service requirements, location of outlets to meet local operational/WHS requirements
RECP-15 | Reception, 15m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1
RECP-15 | Reception, 15m2 | SE | ITSE-038 | READER: access control, general | 1 | optional
RECP-15 | Reception, 15m2 | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1 | optional, to external entry and/or unit entry as required
RECP-20 | Reception, 20m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | carpet tiles are also acceptable
RECP-20 | Reception, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 18
RECP-20 | Reception, 20m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
RECP-20 | Reception, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
RECP-20 | Reception, 20m2 | FA | CLGE-001 | CEILING: plasterboard, bulkhead, paint | 3
RECP-20 | Reception, 20m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
RECP-20 | Reception, 20m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
RECP-20 | Reception, 20m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
RECP-20 | Reception, 20m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 3
RECP-20 | Reception, 20m2 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 2
RECP-20 | Reception, 20m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 2
RECP-20 | Reception, 20m2 | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 2
RECP-20 | Reception, 20m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 2
RECP-20 | Reception, 20m2 | FFE | FIJO-182 | JOINERY UNIT: reception desk | 1
RECP-20 | Reception, 20m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
RECP-20 | Reception, 20m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 3 | dual screens may be provided to suit service requirements
RECP-20 | Reception, 20m2 | FFE | ITNE-023 | PRINTER: networked | 1 | optional
RECP-20 | Reception, 20m2 | FFE | FQGE-056 | PRINTER: desktop, labels | 1 | optional
RECP-20 | Reception, 20m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 3
RECP-20 | Reception, 20m2 | FFE | FIJO-051 | BENCH TOP: 450D, joinery under | 1
RECP-20 | Reception, 20m2 | FFE | FIJO-099 | CAC: cable access cap | 3
RECP-20 | Reception, 20m2 | FFE | WLPT-014 | SCREEN: acoustic, privacy, safety glass | 3 | optional
RECP-20 | Reception, 20m2 | FFE | WLPT-015 | SCREEN: security, safety glass, with communication slot and pass through | 4 | optional
RECP-20 | Reception, 20m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional for label printer
RECP-20 | Reception, 20m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
RECP-20 | Reception, 20m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 3 | location of outlets to meet local operational/WHS requirements
RECP-20 | Reception, 20m2 | SE | ELBO-003 | CONTROL PANEL: lighting | 1 | optional
RECP-20 | Reception, 20m2 | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | optional
RECP-20 | Reception, 20m2 | SE | ELBO-015 | SWITCH: light, single | 1
RECP-20 | Reception, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 8 | 3 x for BYO devices, to be located above desk for easy access, location of outlets to meet local operational/WHS requirements, 1 x optional for label printer
RECP-20 | Reception, 20m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 3 | extent of emergency power to be determined based on service requirements, location of outlets to meet local operational/WHS requirements
RECP-20 | Reception, 20m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1
RECP-20 | Reception, 20m2 | SE | ITSE-038 | READER: access control, general | 1 | optional
RECP-20 | Reception, 20m2 | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1 | optional, to external entry and/or unit entry as required
REPR | Reporting Room | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
REPR | Reporting Room | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 5 | or floor vinyl coved skirting
REPR | Reporting Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
REPR | Reporting Room | FA | WLWA-001 | WALL FINISH: acoustic panels, prefinished | 3
REPR | Reporting Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1 | both sides
REPR | Reporting Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
REPR | Reporting Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
REPR | Reporting Room | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1
REPR | Reporting Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
REPR | Reporting Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 3 | 1 per workstation
REPR | Reporting Room | FFE | FIHR-026 | HOOK: coat | 2 | optional
REPR | Reporting Room | FFE | FQCW-016 | BIN: waste, general, 20L | 2
REPR | Reporting Room | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
REPR | Reporting Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
REPR | Reporting Room | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 4 | RIS/PMS monitor per workstation
REPR | Reporting Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2
REPR | Reporting Room | FFE | MMGE-152 | VIEWER: X-ray, 2 screen | 2 | 1 per workstation
REPR | Reporting Room | FFE | FIJO-099 | CAC: cable access cap | 2
REPR | Reporting Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
REPR | Reporting Room | FFE | FIHR-066 | BRACKET: for display screen, wall mounted, articulated, adj height | 2 | optional, for wall mounted display screen, 1 per workstation
REPR | Reporting Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 2
REPR | Reporting Room | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 4
REPR | Reporting Room | FFE | FQDW-043 | PARTITION: workstation system, acoustic | 1
REPR | Reporting Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 2
REPR | Reporting Room | FFE | FQDW-044 | WORKSTATION: straight, adj height, 1800W x 800D x 750H nom. | 2 | adjustable sit to standing height workstation optional
REPR | Reporting Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 2
REPR | Reporting Room | FFE | ITNE-069 | DISPLAY SCREEN: PACS, portrait | 4
REPR | Reporting Room | FFE | ITNE-070 | DISPLAY SCREEN: PACS | 2 | optional, wall mounted for viewing by teams, 1 per workstation
REPR | Reporting Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 4 | 2 to each workstation for deskbox soft wired deskbox connection
REPR | Reporting Room | SE | LIFX-011 | LIGHT: reading light | 2 | 1 per w/s - task light
REPR | Reporting Room | SE | ELGP-109 | GPO: single, wall mounted | 2 | 1 to each workstation for deskbox soft wired deskbox connection
REPR | Reporting Room | SE | ELGP-109 | GPO: single, wall mounted | 2 | optional, for use with wall mounted display screen, 1 per workstation
REPR | Reporting Room | SE | ELGP-208 | GPO: double, wall mounted | 2 | to bench, x-ray viewer
REPR | Reporting Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
REPR | Reporting Room | SE | ELBO-012 | SWITCH: light, dimmer | 3 | to room
REPR | Reporting Room | SE | ITIN-066 | OUTLET: HDMI, wall mounted | 2 | 1 to each workstation for deskbox soft wired deskbox connection
REPR | Reporting Room | SE | ITIN-066 | OUTLET: HDMI, wall mounted | 2 | optional, for use with wall mounted display screen, 1 per workstation
REPR | Reporting Room | SE | ITIN-029 | OUTLET: data, double RJ45, desk box | 4
REPR | Reporting Room | SE | ITIN-067 | OUTLET: HDMI, desk box | 2 | optional, 1 per workstation
REPR | Reporting Room | SE | ELGP-212 | GPO: double, desk box | 8
REPR | Reporting Room | SE | ITIN-145 | DESK BOX: for power, data, AV, flush mounted, softwired | 2
REPW | Reporting Workstation | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
REPW | Reporting Workstation | FA | FLSK-009 | SKIRTING: vinyl, feather edged, 150H, prefinished | 1 | or floor vinyl coved skirting
REPW | Reporting Workstation | FA | WLWA-001 | WALL FINISH: acoustic panels, prefinished | 1 | to long edge of w/s where it abuts a wall
REPW | Reporting Workstation | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
REPW | Reporting Workstation | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
REPW | Reporting Workstation | FA | CLCN-009 | CORNICE: wall trim fixing | 1
REPW | Reporting Workstation | FFE | FQBS-018 | CHAIR: office, ergonomic | 2 | optional, 1 per w/s with a 2nd per w/s
REPW | Reporting Workstation | FFE | FQCW-016 | BIN: waste, general, 20L | 1
REPW | Reporting Workstation | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1 | optional, for ceiling mounted display screen
REPW | Reporting Workstation | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 2 | RIS/PMS monitor per w/s
REPW | Reporting Workstation | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
REPW | Reporting Workstation | FFE | FIJO-099 | CAC: cable access cap | 1
REPW | Reporting Workstation | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
REPW | Reporting Workstation | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 2
REPW | Reporting Workstation | FFE | FQDW-043 | PARTITION: workstation system, acoustic | 2
REPW | Reporting Workstation | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 1
REPW | Reporting Workstation | FFE | FQDW-044 | WORKSTATION: straight, adj height, 1800W x 800D x 750H nom. | 1 | adjustable sit to standing height workstation optional
REPW | Reporting Workstation | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
REPW | Reporting Workstation | FFE | ITNE-069 | DISPLAY SCREEN: PACS, portrait | 2
REPW | Reporting Workstation | FFE | ITNE-070 | DISPLAY SCREEN: PACS | 1 | optional, ceiling mounted, for viewing by teams
REPW | Reporting Workstation | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | for deskbox soft wired deskbox connection
REPW | Reporting Workstation | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | optional, for use with optional large format ceiling mounted display screen, 1 per w/s
REPW | Reporting Workstation | SE | LIFX-011 | LIGHT: reading light | 1
REPW | Reporting Workstation | SE | ELGP-109 | GPO: single, wall mounted | 1 | for deskbox soft wired deskbox connection
REPW | Reporting Workstation | SE | ELBO-012 | SWITCH: light, dimmer | 1
REPW | Reporting Workstation | SE | ITIN-061 | OUTLET: HDMI, ceiling mounted | 1 | optional, for use with optional large format ceiling mounted display screen, 1 per w/s
REPW | Reporting Workstation | SE | ITIN-066 | OUTLET: HDMI, wall mounted | 1 | optional, for deskbox soft wired deskbox connection
REPW | Reporting Workstation | SE | ITIN-029 | OUTLET: data, double RJ45, desk box | 2
REPW | Reporting Workstation | SE | ITIN-067 | OUTLET: HDMI, desk box | 1 | optional
REPW | Reporting Workstation | SE | ELGP-212 | GPO: double, desk box | 4
REPW | Reporting Workstation | SE | ITIN-145 | DESK BOX: for power, data, AV, flush mounted, softwired | 1
SECR | Security Observation Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
SECR | Security Observation Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
SECR | Security Observation Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
SECR | Security Observation Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
SECR | Security Observation Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
SECR | Security Observation Room | FA | DWWH-005 | WINDOW | 1 | optional, provide as required for surveillance of external areas
SECR | Security Observation Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | lockable, with secure staff-only access
SECR | Security Observation Room | FA | DWOB-016 | OBSERVATION: internal window, alum/safety glass, one way, powdercoat | 1
SECR | Security Observation Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SECR | Security Observation Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 2
SECR | Security Observation Room | FFE | DWWD-004 | BLIND: roller | 1 | optional, to external window if provided
SECR | Security Observation Room | FFE | FQSN-061 | SAFE: medium | 1 | for securing lost items of value as required
SECR | Security Observation Room | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 1
SECR | Security Observation Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 3
SECR | Security Observation Room | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
SECR | Security Observation Room | FFE | FQSN-060 | SAFE: guns | 1 | optional, for securing contraband and weapons, provision in alignment with operational policies and jurisdictional requirements; may be floor mounted.
SECR | Security Observation Room | FFE | DWWH-002 | HATCH: pass-through | 1 | sliding window for reception
SECR | Security Observation Room | FFE | ITSE-019 | CAMERA: identification photo | 1 | optional, for provision of identification cards, access cards and visitor passes, this may occur at a central security unit reception
SECR | Security Observation Room | FFE | FQGE-003 | CHARGER: battery | 1 | for radios, cameras and other devices as required
SECR | Security Observation Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 2
SECR | Security Observation Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | optional, for provision of identification cards, access cards and visitor passes, this may occur at a central security unit reception
SECR | Security Observation Room | FFE | ITSE-048 | DISPLAY SCREEN: CCTV system | 1 | type and quantity to meet security requirements for CCTV camera feed observation
SECR | Security Observation Room | FFE | ITNE-023 | PRINTER: networked | 1
SECR | Security Observation Room | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 2 | 1 to each computer workstation
SECR | Security Observation Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
SECR | Security Observation Room | FFE | ITSE-010 | CABINET: key management, electronic | 1 | optional
SECR | Security Observation Room | FFE | ITNE-022 | PRINTER: identification / access control card | 1 | optional, for provision of identification cards, access cards and visitor passes, this may occur at a central security unit reception
SECR | Security Observation Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 1
SECR | Security Observation Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead is also acceptable
SECR | Security Observation Room | FFE | FIHR-080 | BRACKET: for bench, 750D, wall mounted | 2
SECR | Security Observation Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | optional, to PC for provision of identification cards, access cards and visitor passes, this may occur at a central security unit reception
SECR | Security Observation Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | 1 to each workstation, 1 to bench with printers
SECR | Security Observation Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to CCTV display screen; quantity to meet security requirements
SECR | Security Observation Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 6 | 2 to each workstation, 2 to bench with printers
SECR | Security Observation Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | optional, to PC for provision of identification cards, access cards and visitor passes, this may occur at a central security unit reception
SECR | Security Observation Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | to CCTV display screen; quantity to meet security requirements
SECR | Security Observation Room | SE | ITSE-006 | BUTTON: duress, fixed | 1 | mounted under bench; may be wall mounted
SECR | Security Observation Room | SE | ITSE-038 | READER: access control, general | 1
SECR | Security Observation Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
SHAC | Shower - Accessible | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
SHAC | Shower - Accessible | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
SHAC | Shower - Accessible | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
SHAC | Shower - Accessible | FA | WLWA-019 | WALL FINISH: vinyl, full height | 6 | tiles are also acceptable
SHAC | Shower - Accessible | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1 | water resistant ceiling tiles are also acceptable
SHAC | Shower - Accessible | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | with privacy latch and emergency release
SHAC | Shower - Accessible | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
SHAC | Shower - Accessible | FFE | MMGE-024 | CHAIR: shower | 1
SHAC | Shower - Accessible | FFE | FIHR-026 | HOOK: coat | 2
SHAC | Shower - Accessible | FFE | FIRT-033 | TRACK: curtain, shower | 1 | optional, refer to local infection control policy
SHAC | Shower - Accessible | FFE | FQGE-011 | CURTAIN: shower | 1 | optional, refer to local infection control policy
SHAC | Shower - Accessible | FFE | FIJO-194 | SHELF: laminate | 1
SHAC | Shower - Accessible | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 1 | or soap dispenser
SHAC | Shower - Accessible | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 1
SHAC | Shower - Accessible | FFE | HYTP-045 | SHOWER: hand-held, accessible | 1
SHAC | Shower - Accessible | FFE | FIRT-045 | GRAB RAIL: shower, horizontal and vertical | 1
SHAC | Shower - Accessible | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | optional, IPX4 rating required
SHAC | Shower - Accessible | SE | HYDR-004 | DRAIN: floor waste | 1
SHAC | Shower - Accessible | SE | ELBO-015 | SWITCH: light, single | 1
SHAC | Shower - Accessible | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | optional, IPX4 rating required
SHAC | Shower - Accessible | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | optional, IPX4 rating required
SHPT | Shower - Patient | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
SHPT | Shower - Patient | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
SHPT | Shower - Patient | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
SHPT | Shower - Patient | FA | WLWA-019 | WALL FINISH: vinyl, full height | 6 | tiles are also acceptable
SHPT | Shower - Patient | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1 | water resistant ceiling tiles are also acceptable
SHPT | Shower - Patient | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | with privacy latch and emergency release
SHPT | Shower - Patient | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
SHPT | Shower - Patient | FFE | MMGE-024 | CHAIR: shower | 1 | or commode chair, optional
SHPT | Shower - Patient | FFE | FIHR-026 | HOOK: coat | 2
SHPT | Shower - Patient | FFE | FIRT-034 | TRACK: curtain, shower, 90 deg | 1 | optional, refer to local infection control policy
SHPT | Shower - Patient | FFE | FQGE-011 | CURTAIN: shower | 1 | optional, refer to local infection control policy
SHPT | Shower - Patient | FFE | FQCW-016 | BIN: waste, general, 20L | 1
SHPT | Shower - Patient | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
SHPT | Shower - Patient | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
SHPT | Shower - Patient | FFE | FIBM-006 | MIRROR: fixed, glass | 1
SHPT | Shower - Patient | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 1 | or soap dispenser
SHPT | Shower - Patient | FFE | FIRT-016 | RAIL: towel | 1 | with grab rail strength/fixings
SHPT | Shower - Patient | FFE | HYBA-011 | BASIN: handwash, Type C | 1
SHPT | Shower - Patient | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1
SHPT | Shower - Patient | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 1
SHPT | Shower - Patient | FFE | HYTP-020 | SHOWER: hand-held | 1
SHPT | Shower - Patient | FFE | FIRT-045 | GRAB RAIL: shower, horizontal and vertical | 1
SHPT | Shower - Patient | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | IPX4 rating required
SHPT | Shower - Patient | SE | HYDR-004 | DRAIN: floor waste | 1
SHPT | Shower - Patient | SE | ELBO-015 | SWITCH: light, single | 1 | IPX4 rating required
SHPT | Shower - Patient | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | IPX4 rating required
SHPT | Shower - Patient | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | IPX4 rating required
SHPT | Shower - Patient | SE | ELGP-208 | GPO: double, wall mounted | 1 | optional, to be classified as a shaver outlet or installed within a cupboard if provided
SHST | Shower - Staff | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
SHST | Shower - Staff | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to be suitable for bare feet; tiles are also acceptable
SHST | Shower - Staff | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
SHST | Shower - Staff | FA | WLWA-019 | WALL FINISH: vinyl, full height | 6 | tiles are also acceptable
SHST | Shower - Staff | FA | CLFS-009 | CEILING: plasterboard, water resistant, flush set, suspended, paint, washable | 1 | water resistant ceiling tiles are also acceptable
SHST | Shower - Staff | FA | DWSC-001 | DOOR: 850mm c/o, single leaf, solid core, paint | 1 | with privacy latch
SHST | Shower - Staff | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
SHST | Shower - Staff | FFE | FIGE-003 | BENCH SEAT: fixed, indoor | 1 | material to be suitable for wet areas
SHST | Shower - Staff | FFE | FIHR-026 | HOOK: coat | 1
SHST | Shower - Staff | FFE | FIRT-033 | TRACK: curtain, shower | 1 | optional, refer to local infection control policy
SHST | Shower - Staff | FFE | FQGE-011 | CURTAIN: shower | 1 | optional, refer to local infection control policy
SHST | Shower - Staff | FFE | FIDI-030 | DISPENSER: soap | 1
SHST | Shower - Staff | FFE | FIBM-006 | MIRROR: fixed, glass | 1
SHST | Shower - Staff | FFE | FIDI-049 | HOLDER: soap/shampoo, non-recessed | 1 | optional
SHST | Shower - Staff | FFE | FIRT-016 | RAIL: towel | 1
SHST | Shower - Staff | FFE | HYTP-039 | TAPWARE: mixer, shower, lever action | 1
SHST | Shower - Staff | FFE | HYTP-018 | SHOWER: fixed rose | 1
SHST | Shower - Staff | SE | HYDR-004 | DRAIN: floor waste | 1
SHST | Shower - Staff | SE | ELBO-015 | SWITCH: light, single | 1
SHST | Shower - Staff | SE | ITSE-038 | READER: access control, general | 1 | optional
SPECT-CT | SPECT-CT Imaging Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
SPECT-CT | SPECT-CT Imaging Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 11
SPECT-CT | SPECT-CT Imaging Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-007 | WALL FINISH: splashback | 2
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 7
SPECT-CT | SPECT-CT Imaging Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
SPECT-CT | SPECT-CT Imaging Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | optional, patient observation by camera may be considered, glazing to be AS1428 compliant safety glass
SPECT-CT | SPECT-CT Imaging Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-045 | IV STAND: mobile | 1 | optional
SPECT-CT | SPECT-CT Imaging Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMBE-042 | SLIDE: patient | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-139 | TROLLEY: procedure | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-002 | ADAPTER: suction | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-262 | DEVICE: nuclear medicine, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | ITCL-012 | CAMERA: observation | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-264 | CART: collimator, SPECT/CT | 4 | required quantity dependent on patient mix
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-265 | BOARD: distribution, mains, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-263 | BOX: line connection, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-100 | INJECTOR: contrast, CT | 1 | optional, floor or ceiling mounted to be confirmed
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-272 | BOARD: disconnect, mains, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-273 | PHANTOM: SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-377 | WARMER: contrast | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-078 | BRACKET: for observation camera, wall mounted | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-179 | MACHINE: electrocardiography (ECG), c/w trolley and printer shelf | 1 | procured with SPECT/CT system vendor
SPECT-CT | SPECT-CT Imaging Room | FFE | MMHA-032 | MONITOR: patient, cardiac, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FQCW-066 | BIN: waste, sharps, on benchtop, shielded | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | on wall near SPECT/CT unit
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-006 | RCD: residual current device | 1 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional, on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1 | auto on when modality is in use
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional, on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel + indicator button & light
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-109 | GPO: single, wall mounted | 2 | clock
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 5 | to walls
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 2
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ELGT-007 | UPS: uninterruptable power supply unit | 1 | optional, dependent on vendor requirements
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
SPECT-CT | SPECT-CT Imaging Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | patient services
SPECT-CT | SPECT-CT Imaging Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors extent as specified by radiation consultant
SPECT-CT | SPECT-CT Imaging Room | SE | ELBO-012 | SWITCH: light, dimmer | 1 | number determined by number of doors
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-007 | UPS: status indicator | 1
SPECT-CT | SPECT-CT Imaging Room | SE | LIFX-038 | LIGHT: backlit feature panels | 2
SPECT-CTCR | SPECT-CT Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
SPECT-CTCR | SPECT-CT Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
SPECT-CTCR | SPECT-CT Control Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
SPECT-CTCR | SPECT-CT Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
SPECT-CTCR | SPECT-CT Control Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
SPECT-CTCR | SPECT-CT Control Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | safety glass
SPECT-CTCR | SPECT-CT Control Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
SPECT-CTCR | SPECT-CT Control Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
SPECT-CTCR | SPECT-CT Control Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 2
SPECT-CTCR | SPECT-CT Control Room | FFE | DWWD-004 | BLIND: roller | 1 | to observation window for patient privacy
SPECT-CTCR | SPECT-CT Control Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-044 | RACK: lead apron | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | MMGE-032 | DISPLAY SCREEN: slave, for patient monitor | 1 | optional
SPECT-CTCR | SPECT-CT Control Room | FFE | FQGE-056 | PRINTER: desktop, labels | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1 | communication with patient, on bench required if integrated intercom system not provided by equipment manufacturer
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 1 | communication with patient, on bench required if integrated intercom system not provided by equipment manufacturer
SPECT-CTCR | SPECT-CT Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 2
SPECT-CTCR | SPECT-CT Control Room | FFE | MMMI-269 | COMPUTER: tower, acquisition, SPECT/CT | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | MMMI-274 | CONSOLE: control, SPECT/CT | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-066 | BRACKET: for display screen, wall mounted, articulated, adj height | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | MMMI-369 | DISPLAY SCREEN: contrast injector | 1 | optional
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 2
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-070 | DISPLAY SCREEN: PACS | 2
SPECT-CTCR | SPECT-CT Control Room | FFE | FQDW-054 | WORKSTATION: straight, adj height, 1600W x 750D x 750H nom. | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FQGE-094 | SCANNER: desktop, document, upright | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3 | 1 outlet for display screen, patient monitor slave; one outlet for telephone
SPECT-CTCR | SPECT-CT Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 4
SPECT-CTCR | SPECT-CT Control Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1 | mains power, consult manufacturer for safety and compatibility with equipment.
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-135 | GPO: UPS, single, wall mounted | 5
SPECT-CTCR | SPECT-CT Control Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1 | auto on when modality is in use
SPECT-CTCR | SPECT-CT Control Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-208 | GPO: double, wall mounted | 1 | one for printer
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 4
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 2
SRM-15 | Staff Room, 15m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent around sink
SRM-15 | Staff Room, 15m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
SRM-15 | Staff Room, 15m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
SRM-15 | Staff Room, 15m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
SRM-15 | Staff Room, 15m2 | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 2
SRM-15 | Staff Room, 15m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
SRM-15 | Staff Room, 15m2 | FA | DWWH-005 | WINDOW | 1 | desirable
SRM-15 | Staff Room, 15m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
SRM-15 | Staff Room, 15m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SRM-15 | Staff Room, 15m2 | FFE | FSPV-004 | FIRE EXTINGUISHER | 1
SRM-15 | Staff Room, 15m2 | FFE | FQBS-015 | CHAIR: dining | 4
SRM-15 | Staff Room, 15m2 | FFE | FQBS-039 | LOUNGE: 1 seat | 2
SRM-15 | Staff Room, 15m2 | FFE | DWWD-004 | BLIND: roller | 1
SRM-15 | Staff Room, 15m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
SRM-15 | Staff Room, 15m2 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
SRM-15 | Staff Room, 15m2 | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 1
SRM-15 | Staff Room, 15m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to sink
SRM-15 | Staff Room, 15m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to sink
SRM-15 | Staff Room, 15m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
SRM-15 | Staff Room, 15m2 | FFE | FSPV-005 | FIRE EXTINGUISHER: blanket | 1
SRM-15 | Staff Room, 15m2 | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 1
SRM-15 | Staff Room, 15m2 | FFE | FQCW-019 | BIN: waste, general, 50L | 1 | quantity and type of bins to be based on local waste management policies
SRM-15 | Staff Room, 15m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
SRM-15 | Staff Room, 15m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1
SRM-15 | Staff Room, 15m2 | FFE | HYGE-003 | BOILING/CHILLED WATER UNIT: Combination | 1 | provide grille in joinery for ventilation
SRM-15 | Staff Room, 15m2 | FFE | HYTP-029 | TAPWARE: boiling/chilled water unit (comb) | 1
SRM-15 | Staff Room, 15m2 | FFE | FQCA-067 | OVEN: microwave, domestic | 1
SRM-15 | Staff Room, 15m2 | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
SRM-15 | Staff Room, 15m2 | FFE | FQCA-032 | FRIDGE: domestic, 300L | 1
SRM-15 | Staff Room, 15m2 | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
SRM-15 | Staff Room, 15m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
SRM-15 | Staff Room, 15m2 | FFE | FQTA-055 | TABLE: round, 900 diam, dining height | 1
SRM-15 | Staff Room, 15m2 | FFE | FQTA-075 | TABLE: square, 600W, low height | 1
SRM-15 | Staff Room, 15m2 | FFE | FQBS-042 | LOUNGE: 2 seat, 1200 wide | 1
SRM-15 | Staff Room, 15m2 | FFE | FIJO-163 | CUPBOARD: tall, single microwave | 1
SRM-15 | Staff Room, 15m2 | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead is also acceptable
SRM-15 | Staff Room, 15m2 | FFE | HYBA-080 | SINK: single, with drainer, accessible | 1
SRM-15 | Staff Room, 15m2 | FFE | FIJO-226 | CUPBOARD: drawer unit, bin enclosure | 1 | inclusion will depend on waste management and infection control policies
SRM-15 | Staff Room, 15m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 x to television, 1 x to telephone
SRM-15 | Staff Room, 15m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
SRM-15 | Staff Room, 15m2 | SE | ITIN-070 | OUTLET: MATV | 1 | to television, may be replaced with data outlet in accordance with ICT policy
SRM-15 | Staff Room, 15m2 | SE | ELBO-015 | SWITCH: light, single | 1
SRM-15 | Staff Room, 15m2 | SE | ITCL-001 | ANNUNCIATOR: nurse call | 1 | optional, as required in patient care areas, audible
SRM-15 | Staff Room, 15m2 | SE | ELGP-109 | GPO: single, wall mounted | 4
SRM-15 | Staff Room, 15m2 | SE | ELGP-208 | GPO: double, wall mounted | 4
SRM-15 | Staff Room, 15m2 | SE | ITSE-038 | READER: access control, general | 1 | optional
SRM-35 | Staff Room, 35m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent around sink
SRM-35 | Staff Room, 35m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 13
SRM-35 | Staff Room, 35m2 | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
SRM-35 | Staff Room, 35m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
SRM-35 | Staff Room, 35m2 | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 3
SRM-35 | Staff Room, 35m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
SRM-35 | Staff Room, 35m2 | FA | DWWH-005 | WINDOW | 1 | desirable
SRM-35 | Staff Room, 35m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
SRM-35 | Staff Room, 35m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SRM-35 | Staff Room, 35m2 | FFE | FSPV-004 | FIRE EXTINGUISHER | 1
SRM-35 | Staff Room, 35m2 | FFE | FQBS-015 | CHAIR: dining | 16
SRM-35 | Staff Room, 35m2 | FFE | FQBS-039 | LOUNGE: 1 seat | 4
SRM-35 | Staff Room, 35m2 | FFE | DWWD-004 | BLIND: roller | 1
SRM-35 | Staff Room, 35m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
SRM-35 | Staff Room, 35m2 | FFE | FIJO-116 | CUPBOARD: tall, double door, adj shelves | 1
SRM-35 | Staff Room, 35m2 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
SRM-35 | Staff Room, 35m2 | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 2
SRM-35 | Staff Room, 35m2 | FFE | FIJO-156 | CUPBOARD: wall mount, single door, adj shelves | 1
SRM-35 | Staff Room, 35m2 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to sink
SRM-35 | Staff Room, 35m2 | FFE | FIDI-030 | DISPENSER: soap | 1 | to sink
SRM-35 | Staff Room, 35m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
SRM-35 | Staff Room, 35m2 | FFE | FSPV-005 | FIRE EXTINGUISHER: blanket | 1
SRM-35 | Staff Room, 35m2 | FFE | FIJO-106 | CUPBOARD: drawer unit, 4 equal | 1
SRM-35 | Staff Room, 35m2 | FFE | FQCW-019 | BIN: waste, general, 50L | 1 | quantity and type of bins to be based on local waste management policies
SRM-35 | Staff Room, 35m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
SRM-35 | Staff Room, 35m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1
SRM-35 | Staff Room, 35m2 | FFE | HYGE-003 | BOILING/CHILLED WATER UNIT: Combination | 1 | provide grille in joinery for ventilation
SRM-35 | Staff Room, 35m2 | FFE | HYTP-029 | TAPWARE: boiling/chilled water unit (comb) | 1
SRM-35 | Staff Room, 35m2 | FFE | FQCA-067 | OVEN: microwave, domestic | 2
SRM-35 | Staff Room, 35m2 | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
SRM-35 | Staff Room, 35m2 | FFE | FQCA-032 | FRIDGE: domestic, 300L | 1
SRM-35 | Staff Room, 35m2 | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
SRM-35 | Staff Room, 35m2 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
SRM-35 | Staff Room, 35m2 | FFE | FQTA-055 | TABLE: round, 900 diam, dining height | 4
SRM-35 | Staff Room, 35m2 | FFE | FQTA-075 | TABLE: square, 600W, low height | 3
SRM-35 | Staff Room, 35m2 | FFE | FQBS-042 | LOUNGE: 2 seat, 1200 wide | 2
SRM-35 | Staff Room, 35m2 | FFE | FIJO-164 | CUPBOARD: tall, double microwave | 1
SRM-35 | Staff Room, 35m2 | FFE | FIJO-212 | BULKHEAD: joinery | 3 | plasterboard bulkhead is also acceptable
SRM-35 | Staff Room, 35m2 | FFE | FQCW-062 | BIN: waste, recycling, comingled, 50L | 1 | quantity and type of bins to be based on local waste management policies
SRM-35 | Staff Room, 35m2 | FFE | FIJO-228 | JOINERY UNIT: sink shroud, accessible | 1
SRM-35 | Staff Room, 35m2 | FFE | HYBA-080 | SINK: single, with drainer, accessible | 1
SRM-35 | Staff Room, 35m2 | FFE | FIJO-226 | CUPBOARD: drawer unit, bin enclosure | 2 | inclusion will depend on waste management and infection control policies
SRM-35 | Staff Room, 35m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 x to television, 1 x to telephone
SRM-35 | Staff Room, 35m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
SRM-35 | Staff Room, 35m2 | SE | ITIN-070 | OUTLET: MATV | 1 | to television, may be replaced with data outlet in accordance with ICT policy
SRM-35 | Staff Room, 35m2 | SE | ELBO-015 | SWITCH: light, single | 1
SRM-35 | Staff Room, 35m2 | SE | ITCL-001 | ANNUNCIATOR: nurse call | 1 | optional, as required in patient care areas, audible
SRM-35 | Staff Room, 35m2 | SE | ELGP-109 | GPO: single, wall mounted | 5
SRM-35 | Staff Room, 35m2 | SE | ELGP-208 | GPO: double, wall mounted | 6
SRM-35 | Staff Room, 35m2 | SE | ITSE-038 | READER: access control, general | 1 | optional
SSTN-10 | Staff Station, 10m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
SSTN-10 | Staff Station, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
SSTN-10 | Staff Station, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
SSTN-10 | Staff Station, 10m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | acoustic ceiling treatments are also acceptable, including bulkheads
SSTN-10 | Staff Station, 10m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SSTN-10 | Staff Station, 10m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 2
SSTN-10 | Staff Station, 10m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
SSTN-10 | Staff Station, 10m2 | FFE | FIJO-178 | JOINERY UNIT: staff station | 1 | depth of bench and counter top to be confirmed based on service requirements and must consider ergonomics, staff safety and security, WHS requirements and Australian Standards
SSTN-10 | Staff Station, 10m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1
SSTN-10 | Staff Station, 10m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 2 | under counter top
SSTN-10 | Staff Station, 10m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 2
SSTN-10 | Staff Station, 10m2 | FFE | ITNE-023 | PRINTER: networked | 1
SSTN-10 | Staff Station, 10m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2
SSTN-10 | Staff Station, 10m2 | FFE | FIJO-099 | CAC: cable access cap | 2
SSTN-10 | Staff Station, 10m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
SSTN-10 | Staff Station, 10m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 2 | location of outlets to meet local operational/WHS requirements
SSTN-10 | Staff Station, 10m2 | SE | ELBO-003 | CONTROL PANEL: lighting | 1 | optional, location dependent on project requirements
SSTN-10 | Staff Station, 10m2 | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | optional, location dependent on project requirements
SSTN-10 | Staff Station, 10m2 | SE | LIFX-008 | LIGHT: task light, built in | 2 | under counter top
SSTN-10 | Staff Station, 10m2 | SE | ELBO-015 | SWITCH: light, single | 2 | to task lighting
SSTN-10 | Staff Station, 10m2 | SE | MGFP-001 | ALARM PANEL: medical gases | 1 | optional, location dependent on project requirements
SSTN-10 | Staff Station, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 3 | 2 x for BYO devices, to be located above desk for easy access, location of outlets to meet local operational/WHS requirements
SSTN-10 | Staff Station, 10m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | location of outlets to meet local operational/WHS requirements
SSTN-10 | Staff Station, 10m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1 | under bench
SSTN-10 | Staff Station, 10m2 | SE | ELGP-130 | GPO: UPS, single | 1 | to printer; UPS to be provided if downtime computer required
SSTN-10 | Staff Station, 10m2 | SE | ELGP-235 | GPO: UPS, double, wall mounted | 1 | UPS to be provided if downtime computer required, location of outlets to meet local operational/WHS requirements
SSTN-10 | Staff Station, 10m2 | SE | ELPR-007 | UPS: status indicator | 1 | optional
SSTN-14 | Staff Station, 14m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
SSTN-14 | Staff Station, 14m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
SSTN-14 | Staff Station, 14m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
SSTN-14 | Staff Station, 14m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | acoustic ceiling treatments are also acceptable, including bulkheads
SSTN-14 | Staff Station, 14m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SSTN-14 | Staff Station, 14m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 3
SSTN-14 | Staff Station, 14m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
SSTN-14 | Staff Station, 14m2 | FFE | FIJO-178 | JOINERY UNIT: staff station | 1 | depth of bench and counter top to be confirmed based on service requirements and must consider ergonomics, staff safety and security, WHS requirements and Australian Standards
SSTN-14 | Staff Station, 14m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 2
SSTN-14 | Staff Station, 14m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 2 | under counter top
SSTN-14 | Staff Station, 14m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 3
SSTN-14 | Staff Station, 14m2 | FFE | ITNE-023 | PRINTER: networked | 1
SSTN-14 | Staff Station, 14m2 | FFE | FQGE-056 | PRINTER: desktop, labels | 1 | optional, dependent on service requirements
SSTN-14 | Staff Station, 14m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 3
SSTN-14 | Staff Station, 14m2 | FFE | FIJO-099 | CAC: cable access cap | 3
SSTN-14 | Staff Station, 14m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 x optional, to labels printer if provided
SSTN-14 | Staff Station, 14m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 3 | location of outlets to meet local operational/WHS requirements
SSTN-14 | Staff Station, 14m2 | SE | ELBO-003 | CONTROL PANEL: lighting | 1 | optional, location dependent on project requirements
SSTN-14 | Staff Station, 14m2 | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | optional, location dependent on project requirements
SSTN-14 | Staff Station, 14m2 | SE | LIFX-008 | LIGHT: task light, built in | 3 | under counter top
SSTN-14 | Staff Station, 14m2 | SE | ELBO-015 | SWITCH: light, single | 2 | to task lighting
SSTN-14 | Staff Station, 14m2 | SE | MGFP-001 | ALARM PANEL: medical gases | 1 | optional, location dependent on project requirements
SSTN-14 | Staff Station, 14m2 | SE | ELGP-208 | GPO: double, wall mounted | 6 | 3 x for BYO devices, to be located above desk for easy access, location of outlets to meet local operational/WHS requirements
SSTN-14 | Staff Station, 14m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 3 | location of outlets to meet local operational/WHS requirements
SSTN-14 | Staff Station, 14m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1 | under bench
SSTN-14 | Staff Station, 14m2 | SE | ELGP-130 | GPO: UPS, single | 1 | to printer, UPS to be provided if downtime computer required
SSTN-14 | Staff Station, 14m2 | SE | ELGP-235 | GPO: UPS, double, wall mounted | 1 | UPS to be provided if downtime computer required, location of outlets to meet local operational/WHS requirements
SSTN-14 | Staff Station, 14m2 | SE | ELPR-007 | UPS: status indicator | 1 | optional
SSTN-20 | Staff Station, 20m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
SSTN-20 | Staff Station, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 18
SSTN-20 | Staff Station, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
SSTN-20 | Staff Station, 20m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | acoustic ceiling treatments are also acceptable, including bulkheads
SSTN-20 | Staff Station, 20m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SSTN-20 | Staff Station, 20m2 | FFE | FQBS-018 | CHAIR: office, ergonomic | 4
SSTN-20 | Staff Station, 20m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
SSTN-20 | Staff Station, 20m2 | FFE | FIJO-178 | JOINERY UNIT: staff station | 1 | depth of bench and counter top to be confirmed based on service requirements and must consider ergonomics, staff safety and security, WHS requirements and Australian Standards
SSTN-20 | Staff Station, 20m2 | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 2
SSTN-20 | Staff Station, 20m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 6 | under counter top
SSTN-20 | Staff Station, 20m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 4
SSTN-20 | Staff Station, 20m2 | FFE | ITNE-023 | PRINTER: networked | 1
SSTN-20 | Staff Station, 20m2 | FFE | FQGE-056 | PRINTER: desktop, labels | 1 | optional, dependent on service requirements
SSTN-20 | Staff Station, 20m2 | FFE | ITNE-049 | TELEPHONE: handset, desktop | 4
SSTN-20 | Staff Station, 20m2 | FFE | FIJO-099 | CAC: cable access cap | 5 | 1 x optional, to central monitor if provided
SSTN-20 | Staff Station, 20m2 | FFE | MMGE-187 | MONITOR: patient, central monitoring | 1 | optional, dependent on service requirements
SSTN-20 | Staff Station, 20m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3 | 2 x optional, to central monitor and labels printer if provided
SSTN-20 | Staff Station, 20m2 | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 4 | location of outlets to meet local operational/WHS requirements
SSTN-20 | Staff Station, 20m2 | SE | ELBO-003 | CONTROL PANEL: lighting | 1 | optional, location dependent on project requirements
SSTN-20 | Staff Station, 20m2 | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | optional, location dependent on project requirements
SSTN-20 | Staff Station, 20m2 | SE | LIFX-008 | LIGHT: task light, built in | 4 | under counter top
SSTN-20 | Staff Station, 20m2 | SE | ELBO-015 | SWITCH: light, single | 4 | to task lighting
SSTN-20 | Staff Station, 20m2 | SE | MGFP-001 | ALARM PANEL: medical gases | 1 | optional, location dependent on project requirements
SSTN-20 | Staff Station, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 8 | 4 x for BYO devices, to be located above desk for easy access, location of outlets to meet local operational/WHS requirements
SSTN-20 | Staff Station, 20m2 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 5 | 1 x optional, to central monitor if provided
SSTN-20 | Staff Station, 20m2 | SE | ITSE-006 | BUTTON: duress, fixed | 1 | under bench
SSTN-20 | Staff Station, 20m2 | SE | ELGP-130 | GPO: UPS, single | 1 | to printer, UPS to be provided if downtime computer required
SSTN-20 | Staff Station, 20m2 | SE | ELGP-235 | GPO: UPS, double, wall mounted | 1 | UPS to be provided if downtime computer required, location of outlets to meet local operational/WHS requirements
SSTN-20 | Staff Station, 20m2 | SE | ELPR-007 | UPS: status indicator | 1 | optional
STAD | Store - Accountable Drugs | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STAD | Store - Accountable Drugs | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
STAD | Store - Accountable Drugs | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
STAD | Store - Accountable Drugs | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
STAD | Store - Accountable Drugs | FA | DWGE-003 | DOOR: 910mm c/o, steel, strongroom, security door, paint | 1 | outer door; access and door release controls to be determined based on security strategy
STAD | Store - Accountable Drugs | FA | DWGE-002 | DOOR: 910mm c/o, steel, strongroom, grille gate door, paint | 1 | inner door; access and door release controls to be determined based on security strategy
STAD | Store - Accountable Drugs | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
STAD | Store - Accountable Drugs | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 1 | extent dependent on type and height storage solution
STAD | Store - Accountable Drugs | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
STAD | Store - Accountable Drugs | FFE | FIHR-002 | BRACKET: for CCTV camera, ceiling mounted | 1
STAD | Store - Accountable Drugs | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1
STAD | Store - Accountable Drugs | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
STAD | Store - Accountable Drugs | FFE | ITSE-012 | CAMERA: CCTV | 1
STAD | Store - Accountable Drugs | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | may be on wall or desk mounted bracket, an additional screen may be provided depending on project ICT requirements
STAD | Store - Accountable Drugs | FFE | MMGE-034 | FRIDGE: drugs, large | 1 | optional
STAD | Store - Accountable Drugs | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
STAD | Store - Accountable Drugs | FFE | ITNE-003 | BAR CODING: reader | 1
STAD | Store - Accountable Drugs | FFE | FIJO-212 | BULKHEAD: joinery | 1
STAD | Store - Accountable Drugs | FFE | FQSM-048 | SHELVING: storage solution, shelving, 600W x 400D x 2100H nom., wire, mobile | 3 | type and quantity of storage options to be determined based on service requirements
STAD | Store - Accountable Drugs | FFE | FQGE-093 | PRINTER: desktop, labels, pharmacy | 1
STAD | Store - Accountable Drugs | FFE | FQDW-051 | DESK: adj height, 1500W x 600D x 900H nom. | 1
STAD | Store - Accountable Drugs | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | optional, temperature monitoring, for refrigerator if provided
STAD | Store - Accountable Drugs | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
STAD | Store - Accountable Drugs | SE | MEGE-015 | VENTILATION PORT | 2
STAD | Store - Accountable Drugs | SE | LIFX-008 | LIGHT: task light, built in | 1 | under wall mounted joinery
STAD | Store - Accountable Drugs | SE | ELBO-015 | SWITCH: light, single | 2 | 1 x general room lighting, 1 x task light
STAD | Store - Accountable Drugs | SE | ELGP-208 | GPO: double, wall mounted | 3 | extent of emergency power to be determined based on service requirements
STAD | Store - Accountable Drugs | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | optional, for refrigerator if provided
STAD | Store - Accountable Drugs | SE | ITSE-038 | READER: access control, general | 1
STAD | Store - Accountable Drugs | SE | ITSE-003 | ALARM: intruder detection | 1 | to entry lock mechanism
STBK | Store - Bulk | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STBK | Store - Bulk | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
STBK | Store - Bulk | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
STBK | Store - Bulk | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
STBK | Store - Bulk | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1 | height dependent on movement of trolleys and pallets, crash rails are also acceptable
STBK | Store - Bulk | FA | DWGL-062 | DOOR: 1800mm c/o, double leaf, observation panel, solid core/glass, paint | 1 | lockable or secure staff access, hold open hardware may be required
STBK | Store - Bulk | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
STBK | Store - Bulk | FA | CLCN-009 | CORNICE: wall trim fixing | 1
STBK | Store - Bulk | FFE | FQSN-086 | SHELVING: floor standing, 900W x 400D x 2100H nom., metal | 2 | type and extent of storage solutions will be dependent on service requirements
STBK | Store - Bulk | FFE | FQSM-032 | TROLLEY: receiving | 3
STBK | Store - Bulk | FFE | FQCA-026 | FRIDGE: commercial, floor standing, double door | 1 | optional
STBK | Store - Bulk | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | optional, to refrigerator if provided and temperature monitoring is required
STBK | Store - Bulk | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
STBK | Store - Bulk | SE | ELGP-208 | GPO: double, wall mounted | 1
STBK | Store - Bulk | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | optional, to refrigerator if provided
STBK | Store - Bulk | SE | ITSE-038 | READER: access control, general | 1
STCL | Store - Cleaner | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STCL | Store - Cleaner | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
STCL | Store - Cleaner | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
STCL | Store - Cleaner | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
STCL | Store - Cleaner | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1 | height dependent on movement of trolleys and pallets, crash rails are also acceptable
STCL | Store - Cleaner | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
STCL | Store - Cleaner | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1 | lockable or secure staff access, hold open hardware may be required, vision panel may be provided
STCL | Store - Cleaner | FA | CLCN-009 | CORNICE: wall trim fixing | 1
STCL | Store - Cleaner | FFE | FQSN-012 | CABINET: cleaning equipment | 1 | lockable
STCL | Store - Cleaner | FFE | FIHR-029 | HOOK: equipment | 2 | type and extent of storage solutions will be dependent on service requirements
STCL | Store - Cleaner | FFE | FQSN-086 | SHELVING: floor standing, 900W x 400D x 2100H nom., metal | 3 | type and extent of storage solutions will be dependent on service requirements
STCL | Store - Cleaner | SE | ELBO-015 | SWITCH: light, single | 1
STCL | Store - Cleaner | SE | ELGP-208 | GPO: double, wall mounted | 2
STCL | Store - Cleaner | SE | ITSE-038 | READER: access control, general | 1
STEQ-14 | Store - Equipment, 14m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STEQ-14 | Store - Equipment, 14m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
STEQ-14 | Store - Equipment, 14m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
STEQ-14 | Store - Equipment, 14m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
STEQ-14 | Store - Equipment, 14m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 5 | height dependent on type and height of mobile equipment, crash rails are also acceptable
STEQ-14 | Store - Equipment, 14m2 | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
STEQ-14 | Store - Equipment, 14m2 | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1
STEQ-14 | Store - Equipment, 14m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
STEQ-14 | Store - Equipment, 14m2 | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 4 | optional, provision dependent on service requirements, to be rated for minimum 10kg per hook if provided
STEQ-14 | Store - Equipment, 14m2 | FFE | FIRT-012 | RAIL: equipment | 2 | optional, provision dependent on service requirements
STEQ-14 | Store - Equipment, 14m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
STEQ-14 | Store - Equipment, 14m2 | FFE | FQSN-082 | SHELVING: floor standing, 900W x 450D x 2100H nom., wire | 3 | adjustable shelving; type and quantity of storage options to be determined based on service requirements
STEQ-14 | Store - Equipment, 14m2 | SE | ELBO-015 | SWITCH: light, single | 1
STEQ-14 | Store - Equipment, 14m2 | SE | ELGP-208 | GPO: double, wall mounted | 13 | type, quantity and configuration of outlets dependent on operational and equipment requirements
STEQ-20 | Store - Equipment, 20m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STEQ-20 | Store - Equipment, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
STEQ-20 | Store - Equipment, 20m2 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
STEQ-20 | Store - Equipment, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
STEQ-20 | Store - Equipment, 20m2 | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 5 | height dependent on type and height of mobile equipment, crash rails are also acceptable
STEQ-20 | Store - Equipment, 20m2 | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
STEQ-20 | Store - Equipment, 20m2 | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1
STEQ-20 | Store - Equipment, 20m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
STEQ-20 | Store - Equipment, 20m2 | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 4 | optional, provision dependent on service requirements, to be rated for minimum 10kg per hook if provided
STEQ-20 | Store - Equipment, 20m2 | FFE | FIRT-012 | RAIL: equipment | 3 | optional, provision dependent on service requirements
STEQ-20 | Store - Equipment, 20m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
STEQ-20 | Store - Equipment, 20m2 | FFE | FQSN-082 | SHELVING: floor standing, 900W x 450D x 2100H nom., wire | 3 | adjustable shelving; type and quantity of storage options to be determined based on service requirements
STEQ-20 | Store - Equipment, 20m2 | SE | ELBO-015 | SWITCH: light, single | 1
STEQ-20 | Store - Equipment, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 16 | type, quantity and configuration of outlets dependent on operational and equipment requirements
STFS-10 | Store - Files, 10m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STFS-10 | Store - Files, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
STFS-10 | Store - Files, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
STFS-10 | Store - Files, 10m2 | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
STFS-10 | Store - Files, 10m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
STFS-10 | Store - Files, 10m2 | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | Lockable
STFS-10 | Store - Files, 10m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
STFS-10 | Store - Files, 10m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional.
STFS-10 | Store - Files, 10m2 | FFE | FQSN-085 | SHELVING: floor standing, 900W x 400D x 2100H nom., wire | 9
STFS-10 | Store - Files, 10m2 | FFE | FQGE-095 | STEP: safety | 1
STFS-10 | Store - Files, 10m2 | SE | ELBO-015 | SWITCH: light, single | 1
STFS-10 | Store - Files, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
STFS-10 | Store - Files, 10m2 | SE | ITSE-038 | READER: access control, general | 1
STFS-20 | Store - Files, 20m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STFS-20 | Store - Files, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
STFS-20 | Store - Files, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
STFS-20 | Store - Files, 20m2 | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
STFS-20 | Store - Files, 20m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
STFS-20 | Store - Files, 20m2 | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1
STFS-20 | Store - Files, 20m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
STFS-20 | Store - Files, 20m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
STFS-20 | Store - Files, 20m2 | FFE | FQSN-063 | SHELVING: compactus | 2 | static shelving is also acceptable
STFS-20 | Store - Files, 20m2 | FFE | FQGE-095 | STEP: safety | 1
STFS-20 | Store - Files, 20m2 | SE | ELBO-015 | SWITCH: light, single | 1
STFS-20 | Store - Files, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
STFS-20 | Store - Files, 20m2 | SE | ITSE-038 | READER: access control, general | 1
STGN | Store - General | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
STGN | Store - General | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
STGN | Store - General | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
STGN | Store - General | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
STGN | Store - General | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
STGN | Store - General | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
STGN | Store - General | FA | CLCN-009 | CORNICE: wall trim fixing | 1
STGN | Store - General | FFE | FQSN-007 | BIN: storage bins, wire mesh, stainless steel | 15 | optional
STGN | Store - General | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 3 | optional
STGN | Store - General | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
STGN | Store - General | FFE | FQSN-082 | SHELVING: floor standing, 900W x 450D x 2100H nom., wire | 6 | adjustable shelving; type and quantity of storage options to be determined based on service requirements
STGN | Store - General | SE | ELBO-015 | SWITCH: light, single | 1
STGN | Store - General | SE | ELGP-208 | GPO: double, wall mounted | 1
STRT | Stress Testing | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
STRT | Stress Testing | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
STRT | Stress Testing | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
STRT | Stress Testing | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
STRT | Stress Testing | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
STRT | Stress Testing | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
STRT | Stress Testing | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1
STRT | Stress Testing | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1
STRT | Stress Testing | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
STRT | Stress Testing | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 5
STRT | Stress Testing | FFE | MGFP-002 | FLOWMETER: medical air | 1 | optional
STRT | Stress Testing | FFE | FIHR-019 | BRACKET: for suction bottle | 2
STRT | Stress Testing | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
STRT | Stress Testing | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 1
STRT | Stress Testing | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
STRT | Stress Testing | FFE | FIHR-026 | HOOK: coat | 2 | optional
STRT | Stress Testing | FFE | FIJO-131 | CUPBOARD: tall, single door, adj shelves, lockable | 1
STRT | Stress Testing | FFE | FIRT-025 | TRACK: curtain, bedscreen, U shaped | 1
STRT | Stress Testing | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | to entry door
STRT | Stress Testing | FFE | FQCW-016 | BIN: waste, general, 20L | 1
STRT | Stress Testing | FFE | FIDI-019 | DISPENSER: gown/apron | 1
STRT | Stress Testing | FFE | FIDI-014 | DISPENSER: disposable glove | 1
STRT | Stress Testing | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
STRT | Stress Testing | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
STRT | Stress Testing | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
STRT | Stress Testing | FFE | FIDI-021 | DISPENSER: mask | 1
STRT | Stress Testing | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
STRT | Stress Testing | FFE | MGFP-004 | FLOWMETER: oxygen | 2
STRT | Stress Testing | FFE | MMHA-027 | TROLLEY: resuscitation, adult | 1
STRT | Stress Testing | FFE | MMGE-088 | SAFE: dangerous drugs, small | 1 | optional, for pharmacological stress testing
STRT | Stress Testing | FFE | FQBS-052 | STOOL: mobile, high, adj height | 2
STRT | Stress Testing | FFE | MMGE-002 | ADAPTER: suction | 2
STRT | Stress Testing | FFE | MMGE-100 | SUCTION BOTTLE | 2
STRT | Stress Testing | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
STRT | Stress Testing | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
STRT | Stress Testing | FFE | FQSM-003 | PEDESTAL: mobile, wire basket unit, 840H | 2
STRT | Stress Testing | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | refer to BHWS-B (similar)
STRT | Stress Testing | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
STRT | Stress Testing | FFE | ITCL-014 | CLOCK: analogue, electric | 1
STRT | Stress Testing | FFE | MMHA-002 | DEFIBRILLATOR | 1
STRT | Stress Testing | FFE | MMMI-325 | DEVICE: imaging, echocardiography machine | 1
STRT | Stress Testing | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
STRT | Stress Testing | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
STRT | Stress Testing | FFE | MMAH-027 | TREADMILL: automatic | 1
STRT | Stress Testing | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
STRT | Stress Testing | FFE | FIJO-212 | BULKHEAD: joinery | 1
STRT | Stress Testing | FFE | MMGE-179 | MACHINE: electrocardiography (ECG), c/w trolley and printer shelf | 1
STRT | Stress Testing | FFE | MMBE-069 | TROLLEY: patient, echocardiography | 1
STRT | Stress Testing | FFE | ITNE-071 | COMPUTER: desktop, with dual display screens | 1
STRT | Stress Testing | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 5 | 1 outlet for telephone
STRT | Stress Testing | SE | ELPR-006 | RCD: residual current device | 2
STRT | Stress Testing | SE | ELGP-101 | GPO: single, ceiling mounted | 1
STRT | Stress Testing | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel, + indicator button & light
STRT | Stress Testing | SE | ELBO-015 | SWITCH: light, single | 2 | optional, to optional change room
STRT | Stress Testing | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
STRT | Stress Testing | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel, optional
STRT | Stress Testing | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
STRT | Stress Testing | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
STRT | Stress Testing | SE | ELGP-109 | GPO: single, wall mounted | 1
STRT | Stress Testing | SE | ELGP-208 | GPO: double, wall mounted | 6
STRT | Stress Testing | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 3
STRT | Stress Testing | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
STRT | Stress Testing | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
STRT | Stress Testing | SE | ELBO-012 | SWITCH: light, dimmer | 1
STRT | Stress Testing | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4
TRMT | Treatment Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
TRMT | Treatment Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
TRMT | Treatment Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
TRMT | Treatment Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
TRMT | Treatment Room | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
TRMT | Treatment Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
TRMT | Treatment Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 5
TRMT | Treatment Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
TRMT | Treatment Room | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
TRMT | Treatment Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
TRMT | Treatment Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
TRMT | Treatment Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
TRMT | Treatment Room | FFE | FIHR-016 | BRACKET: for sharps bin | 1
TRMT | Treatment Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
TRMT | Treatment Room | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
TRMT | Treatment Room | FFE | FIHR-026 | HOOK: coat | 2 | optional
TRMT | Treatment Room | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
TRMT | Treatment Room | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1
TRMT | Treatment Room | FFE | FQGE-010 | CURTAIN: privacy screen | 1
TRMT | Treatment Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
TRMT | Treatment Room | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
TRMT | Treatment Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
TRMT | Treatment Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
TRMT | Treatment Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
TRMT | Treatment Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
TRMT | Treatment Room | FFE | MMGE-119 | TROLLEY: equipment | 1
TRMT | Treatment Room | FFE | MMBE-043 | TROLLEY: patient, examination | 1
TRMT | Treatment Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1
TRMT | Treatment Room | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
TRMT | Treatment Room | FFE | FQBS-052 | STOOL: mobile, high, adj height | 1
TRMT | Treatment Room | FFE | MMGE-002 | ADAPTER: suction | 1
TRMT | Treatment Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
TRMT | Treatment Room | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
TRMT | Treatment Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
TRMT | Treatment Room | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
TRMT | Treatment Room | FFE | MMGE-031 | DIAGNOSTIC SET: wall mounted | 1 | optional, may be portable/mobile on stand
TRMT | Treatment Room | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
TRMT | Treatment Room | FFE | ITNE-023 | PRINTER: networked | 1 | optional
TRMT | Treatment Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
TRMT | Treatment Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
TRMT | Treatment Room | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1 | optional; type, size and quantity of bins to be determined based on service requirements and waste management policies
TRMT | Treatment Room | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
TRMT | Treatment Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
TRMT | Treatment Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
TRMT | Treatment Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2
TRMT | Treatment Room | SE | ELPR-006 | RCD: residual current device | 4
TRMT | Treatment Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to exam light
TRMT | Treatment Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
TRMT | Treatment Room | SE | LIFX-008 | LIGHT: task light, built in | 1 | under wall mounted joinery
TRMT | Treatment Room | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 1 x task light, 1 x examination light
TRMT | Treatment Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
TRMT | Treatment Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
TRMT | Treatment Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1
TRMT | Treatment Room | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on services panel + indicator button & light
TRMT | Treatment Room | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on services panel, call button on the handset
TRMT | Treatment Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
TRMT | Treatment Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level, for patient trolley
TRMT | Treatment Room | SE | ELGP-208 | GPO: double, wall mounted | 2
TRMT | Treatment Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
TRMT | Treatment Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
TRMT | Treatment Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
TRMT | Treatment Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
TRMT | Treatment Room | SE | ELGP-105 | GPO: single, on services panel | 3
TRMT | Treatment Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent to extend under basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 8
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | WLWA-007 | WALL FINISH: splashback | 2 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to bench
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional, second door for egress
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | CLCN-009 | CORNICE: wall trim fixing | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQBS-030 | CHAIR: visitor, patient | 6
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-116 | CUPBOARD: tall, double door, adj shelves | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-140 | CUPBOARD: underbench, single door, adj shelves | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQGE-010 | CURTAIN: privacy screen | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-104 | CUPBOARD: drawer unit, 3 equal | 1 | for thermaplastic storage, width to be confirmed to suit size of material
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 4
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYBA-010 | BASIN: handwash, Type B | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYBA-069 | SINK: single, with drainer, 930mmL | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMAH-012 | DEVICE: continuous passive motion (CPM), hand and fingers | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMAH-021 | PAN: splint, molding | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | LIFX-026 | LIGHT: examination, wall mounted | 2 | to therapy tables, ceiling mounted is also acceptable
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYDR-008 | TRAP: plaster | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQGE-043 | MACHINE: sewing | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMLA-014 | DEVICE: wax bath | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMGE-150 | VIEWER: X-ray, 1 screen | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQCW-010 | BIN: waste, clinical, 80L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead also acceptable
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 2 | to therapy tables, dependent on ICT strategy
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | ITNE-066 | COMPUTER: all-in-one, wall mounted | 2 | to therapy tables, dependent on ICT strategy
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQTA-084 | TABLE: adjustable height, hand therapy | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQDW-049 | DESK: adj height, 1200W x 600D x 750H nom. | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-053 | DISPENSER: velcro | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELPR-006 | RCD: residual current device | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 2 x examination lights
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELGP-109 | GPO: single, wall mounted | 3 | 1 x optional to x-ray viewing box
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELGP-208 | GPO: double, wall mounted | 9 | extent of emergency power to be determined based on service requirements
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ITSE-006 | BUTTON: duress, fixed | 2
TRMT-POD | Treatment Room - Podiatry | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent to extend under basin
TRMT-POD | Treatment Room - Podiatry | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
TRMT-POD | Treatment Room - Podiatry | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
TRMT-POD | Treatment Room - Podiatry | FA | WLWA-007 | WALL FINISH: splashback | 2
TRMT-POD | Treatment Room - Podiatry | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
TRMT-POD | Treatment Room - Podiatry | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 4
TRMT-POD | Treatment Room - Podiatry | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
TRMT-POD | Treatment Room - Podiatry | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
TRMT-POD | Treatment Room - Podiatry | FA | CLCN-009 | CORNICE: wall trim fixing | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FIHR-019 | BRACKET: for suction bottle | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FIHR-016 | BRACKET: for sharps bin | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 2
TRMT-POD | Treatment Room - Podiatry | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FQGE-010 | CURTAIN: privacy screen | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
TRMT-POD | Treatment Room - Podiatry | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
TRMT-POD | Treatment Room - Podiatry | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
TRMT-POD | Treatment Room - Podiatry | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
TRMT-POD | Treatment Room - Podiatry | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
TRMT-POD | Treatment Room - Podiatry | FFE | MGFP-004 | FLOWMETER: oxygen | 1
TRMT-POD | Treatment Room - Podiatry | FFE | MMGE-139 | TROLLEY: procedure | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FQBS-051 | STOOL: mobile, adj height | 1
TRMT-POD | Treatment Room - Podiatry | FFE | MMGE-002 | ADAPTER: suction | 1
TRMT-POD | Treatment Room - Podiatry | FFE | MMGE-100 | SUCTION BOTTLE | 1
TRMT-POD | Treatment Room - Podiatry | FFE | HYBA-010 | BASIN: handwash, Type B | 1
TRMT-POD | Treatment Room - Podiatry | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
TRMT-POD | Treatment Room - Podiatry | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | computer may alternatively be on benchtop or wall mounted
TRMT-POD | Treatment Room - Podiatry | FFE | LIFX-013 | LIGHT: procedure, ceiling mounted, small | 1
TRMT-POD | Treatment Room - Podiatry | FFE | MMAH-010 | CHAIR: podiatry, electric | 1
TRMT-POD | Treatment Room - Podiatry | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1 | optional; type, size and quantity of bins to be determined based on service requirements and waste management policies
TRMT-POD | Treatment Room - Podiatry | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | computer may alternatively be on benchtop or wall mounted
TRMT-POD | Treatment Room - Podiatry | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead also acceptable
TRMT-POD | Treatment Room - Podiatry | FFE | FIDI-054 | DISPENSER: ppe combination | 1
TRMT-POD | Treatment Room - Podiatry | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 1
TRMT-POD | Treatment Room - Podiatry | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
TRMT-POD | Treatment Room - Podiatry | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
TRMT-POD | Treatment Room - Podiatry | SE | ELPR-006 | RCD: residual current device | 3
TRMT-POD | Treatment Room - Podiatry | SE | ELGP-101 | GPO: single, ceiling mounted | 1
TRMT-POD | Treatment Room - Podiatry | SE | LIFX-008 | LIGHT: task light, built in | 1
TRMT-POD | Treatment Room - Podiatry | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 1 x task light, 1 x examination light
TRMT-POD | Treatment Room - Podiatry | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
TRMT-POD | Treatment Room - Podiatry | SE | MGAS-062 | OUTLET: suction, on services panel | 1
TRMT-POD | Treatment Room - Podiatry | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1
TRMT-POD | Treatment Room - Podiatry | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
TRMT-POD | Treatment Room - Podiatry | SE | ELGP-208 | GPO: double, wall mounted | 3 | extent of emergency power to be determined based on service requirements
TRMT-POD | Treatment Room - Podiatry | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
TRMT-POD | Treatment Room - Podiatry | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
TRMT-POD | Treatment Room - Podiatry | SE | ITSE-006 | BUTTON: duress, fixed | 1
TRMT-POD | Treatment Room - Podiatry | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
TRMT-POD | Treatment Room - Podiatry | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
TRMT-POD | Treatment Room - Podiatry | SE | ELGP-205 | GPO: double, on services panel | 2 | extent of emergency power to be determined based on service requirements
UTRM | Uptake Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
UTRM | Uptake Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
UTRM | Uptake Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 5
UTRM | Uptake Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
UTRM | Uptake Room | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
UTRM | Uptake Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1 | automated, hold open
UTRM | Uptake Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
UTRM | Uptake Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
UTRM | Uptake Room | FFE | FIHR-016 | BRACKET: for sharps bin | 1
UTRM | Uptake Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
UTRM | Uptake Room | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
UTRM | Uptake Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
UTRM | Uptake Room | FFE | FIHR-029 | HOOK: equipment | 2
UTRM | Uptake Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1
UTRM | Uptake Room | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
UTRM | Uptake Room | FFE | MMGE-002 | ADAPTER: suction | 1
UTRM | Uptake Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
UTRM | Uptake Room | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
UTRM | Uptake Room | FFE | ITCL-012 | CAMERA: observation | 1
UTRM | Uptake Room | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | optional
UTRM | Uptake Room | FFE | MMMI-219 | PASS THROUGH: injection port | 1 | optional
UTRM | Uptake Room | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 1
UTRM | Uptake Room | FFE | MMBE-091 | CHAIR: recliner, clinical, electric | 1 | consider bariatric capacity as required
UTRM | Uptake Room | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 1
UTRM | Uptake Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room
UTRM | Uptake Room | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment screen, may be replaced with data outlet in accordance with ICT policy
UTRM | Uptake Room | SE | ELGP-101 | GPO: single, ceiling mounted | 3
UTRM | Uptake Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
UTRM | Uptake Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
UTRM | Uptake Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1
UTRM | Uptake Room | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1
UTRM | Uptake Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
UTRM | Uptake Room | SE | ELGP-109 | GPO: single, wall mounted | 1
UTRM | Uptake Room | SE | ELGP-208 | GPO: double, wall mounted | 1
UTRM | Uptake Room | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
UTRM | Uptake Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
UTRM | Uptake Room | SE | ELBO-012 | SWITCH: light, dimmer | 2 | inside and outside room, to be reviewed to suit operational requirements
UTRM | Uptake Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2
WAIT-10 | Waiting, 10m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
WAIT-10 | Waiting, 10m2 | FA | WLPR-006 | WALL PROTECTION: corner guards | 1 | coordinated with adjacent wall protection
WAIT-10 | Waiting, 10m2 | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | carpet tiles are also acceptable
WAIT-10 | Waiting, 10m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
WAIT-10 | Waiting, 10m2 | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 3 | as required to suit loose furniture layout; crash rails are also acceptable
WAIT-10 | Waiting, 10m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
WAIT-10 | Waiting, 10m2 | FA | DWWH-005 | WINDOW | 1 | desirable
WAIT-10 | Waiting, 10m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
WAIT-10 | Waiting, 10m2 | FFE | FQGE-002 | ARTWORK | 1 | optional
WAIT-10 | Waiting, 10m2 | FFE | FQBS-032 | CHAIR: waiting | 6
WAIT-10 | Waiting, 10m2 | FFE | DWWD-004 | BLIND: roller | 1 | desirable, associated with window
WAIT-10 | Waiting, 10m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 | to suit local infection control policies
WAIT-10 | Waiting, 10m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
WAIT-10 | Waiting, 10m2 | FFE | FQBS-027 | CHAIR: visitor, bariatric | 1 | quantity dependent on service requirements
WAIT-10 | Waiting, 10m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
WAIT-10 | Waiting, 10m2 | FFE | FIHR-045 | RACK: magazine/pamphlet | 1 | optional
WAIT-10 | Waiting, 10m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
WAIT-10 | Waiting, 10m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1
WAIT-10 | Waiting, 10m2 | FFE | FQTA-071 | TABLE: square, 500W, low height | 1
WAIT-10 | Waiting, 10m2 | SE | ITIN-070 | OUTLET: MATV | 1 | may be replaced with data outlet in accordance with ICT policy
WAIT-10 | Waiting, 10m2 | SE | ELGP-109 | GPO: single, wall mounted | 1 | to television
WAIT-10 | Waiting, 10m2 | SE | ELGP-208 | GPO: double, wall mounted | 1
WAIT-20 | Waiting, 20m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
WAIT-20 | Waiting, 20m2 | FA | WLPR-006 | WALL PROTECTION: corner guards | 1 | coordinated with adjacent wall protection
WAIT-20 | Waiting, 20m2 | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | carpet tiles are also acceptable
WAIT-20 | Waiting, 20m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
WAIT-20 | Waiting, 20m2 | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 3 | as required to suit loose furniture layout; crash rails are also acceptable
WAIT-20 | Waiting, 20m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
WAIT-20 | Waiting, 20m2 | FA | DWWH-005 | WINDOW | 1 | desirable
WAIT-20 | Waiting, 20m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
WAIT-20 | Waiting, 20m2 | FFE | FQGE-002 | ARTWORK | 1 | optional
WAIT-20 | Waiting, 20m2 | FFE | FQBS-032 | CHAIR: waiting | 14
WAIT-20 | Waiting, 20m2 | FFE | DWWD-004 | BLIND: roller | 1 | desirable, associated with window
WAIT-20 | Waiting, 20m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 | to suit local infection control policies
WAIT-20 | Waiting, 20m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
WAIT-20 | Waiting, 20m2 | FFE | FQBS-027 | CHAIR: visitor, bariatric | 2 | quantity dependent on service requirements
WAIT-20 | Waiting, 20m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
WAIT-20 | Waiting, 20m2 | FFE | FIHR-045 | RACK: magazine/pamphlet | 1 | optional
WAIT-20 | Waiting, 20m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
WAIT-20 | Waiting, 20m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1
WAIT-20 | Waiting, 20m2 | FFE | FQTA-071 | TABLE: square, 500W, low height | 2
WAIT-20 | Waiting, 20m2 | SE | ITIN-070 | OUTLET: MATV | 1 | may be replaced with data outlet in accordance with ICT policy
WAIT-20 | Waiting, 20m2 | SE | ELGP-109 | GPO: single, wall mounted | 1 | to television
WAIT-20 | Waiting, 20m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
WAIT-30 | Waiting, 30m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
WAIT-30 | Waiting, 30m2 | FA | WLPR-006 | WALL PROTECTION: corner guards | 2 | coordinated with adjacent wall protection
WAIT-30 | Waiting, 30m2 | FA | FLVY-003 | FLOOR FINISH: vinyl, seamless, coved, non slip | 1 | carpet tiles are also acceptable
WAIT-30 | Waiting, 30m2 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
WAIT-30 | Waiting, 30m2 | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 3 | as required to suit loose furniture layout; crash rails are also acceptable
WAIT-30 | Waiting, 30m2 | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
WAIT-30 | Waiting, 30m2 | FA | DWWH-005 | WINDOW | 1 | desirable
WAIT-30 | Waiting, 30m2 | FA | CLCN-009 | CORNICE: wall trim fixing | 1
WAIT-30 | Waiting, 30m2 | FFE | FQGE-002 | ARTWORK | 2 | optional
WAIT-30 | Waiting, 30m2 | FFE | FQBS-032 | CHAIR: waiting | 21
WAIT-30 | Waiting, 30m2 | FFE | DWWD-004 | BLIND: roller | 1 | desirable, associated with window
WAIT-30 | Waiting, 30m2 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2 | to suit local infection control policies
WAIT-30 | Waiting, 30m2 | FFE | FQCW-016 | BIN: waste, general, 20L | 1
WAIT-30 | Waiting, 30m2 | FFE | FQBS-027 | CHAIR: visitor, bariatric | 2 | quantity dependent on service requirements
WAIT-30 | Waiting, 30m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 2
WAIT-30 | Waiting, 30m2 | FFE | FIHR-045 | RACK: magazine/pamphlet | 1 | optional
WAIT-30 | Waiting, 30m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | optional
WAIT-30 | Waiting, 30m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 2
WAIT-30 | Waiting, 30m2 | FFE | FQTA-071 | TABLE: square, 500W, low height | 3
WAIT-30 | Waiting, 30m2 | SE | ITIN-070 | OUTLET: MATV | 2 | Optional. May be replaced with data outlet, in accordance with ICT policy if required.
WAIT-30 | Waiting, 30m2 | SE | ELGP-109 | GPO: single, wall mounted | 2
WAIT-30 | Waiting, 30m2 | SE | ELGP-208 | GPO: double, wall mounted | 2
WAIT-S | Waiting, Sub | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | coordinated with adjacent area, carpet tiles are also acceptable
WAIT-S | Waiting, Sub | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 3
WAIT-S | Waiting, Sub | FA | WLPR-006 | WALL PROTECTION: corner guards | 1 | coordinated with adjacent wall protection
WAIT-S | Waiting, Sub | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
WAIT-S | Waiting, Sub | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 3 | as required to suit loose furniture layout; crash rails are also acceptable
WAIT-S | Waiting, Sub | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | coordinated with adjacent area
WAIT-S | Waiting, Sub | FA | CLCN-009 | CORNICE: wall trim fixing | 1 | coordinated with adjacent area
WAIT-S | Waiting, Sub | FFE | FQBS-032 | CHAIR: waiting | 2
WAIT-S | Waiting, Sub | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 | to suit local infection control policies
WAIT-S | Waiting, Sub | FFE | FQBS-027 | CHAIR: visitor, bariatric | 1 | quantity dependent on service requirements
WAIT-S | Waiting, Sub | FFE | FQTA-071 | TABLE: square, 500W, low height | 1
WAIT-S | Waiting, Sub | FFE | FIDI-056 | HOLDER: document, A4 | 1
WAIT-S | Waiting, Sub | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | optional
WAIT-S | Waiting, Sub | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | optional
WAIT-S | Waiting, Sub | SE | ELGP-208 | GPO: double, wall mounted | 1
WCAC | Toilet - Accessible | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
WCAC | Toilet - Accessible | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | tiles are also acceptable
WCAC | Toilet - Accessible | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
WCAC | Toilet - Accessible | FA | WLWA-007 | WALL FINISH: splashback | 1
WCAC | Toilet - Accessible | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
WCAC | Toilet - Accessible | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
WCAC | Toilet - Accessible | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | with privacy latch, emergency release required if inward opening door
WCAC | Toilet - Accessible | FA | CLCN-009 | CORNICE: wall trim fixing | 1
WCAC | Toilet - Accessible | FFE | FQBS-035 | CHANGE TABLE: baby | 1 | optional
WCAC | Toilet - Accessible | FFE | FIHR-026 | HOOK: coat | 2
WCAC | Toilet - Accessible | FFE | FQCW-016 | BIN: waste, general, 20L | 1
WCAC | Toilet - Accessible | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
WCAC | Toilet - Accessible | FFE | FIDI-030 | DISPENSER: soap | 1
WCAC | Toilet - Accessible | FFE | FIDI-036 | DISPENSER: toilet paper | 1
WCAC | Toilet - Accessible | FFE | FIRT-007 | GRAB RAIL: WC, accessible use | 1
WCAC | Toilet - Accessible | FFE | FIRT-006 | GRAB RAIL: WC | 1
WCAC | Toilet - Accessible | FFE | FIBM-006 | MIRROR: fixed, glass | 1
WCAC | Toilet - Accessible | FFE | FQCW-037 | BIN: waste, sanitary | 1
WCAC | Toilet - Accessible | FFE | HYBA-005 | BASIN: handwash, accessible | 1
WCAC | Toilet - Accessible | FFE | HYWC-001 | TOILET SUITE: accessible | 1 | Large format flush buttons to cistern
WCAC | Toilet - Accessible | FFE | HYTP-035 | TAPWARE: mixer, basin, accessible | 1
WCAC | Toilet - Accessible | FFE | FIGE-011 | BACKREST: accessible toilet | 1
WCAC | Toilet - Accessible | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | optional
WCAC | Toilet - Accessible | SE | HYDR-004 | DRAIN: floor waste | 1
WCAC | Toilet - Accessible | SE | ELBO-015 | SWITCH: light, single | 1
WCAC | Toilet - Accessible | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1
WCAC | Toilet - Accessible | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | optional
WCPT | Toilet - Patient | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
WCPT | Toilet - Patient | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | tiles are also acceptable
WCPT | Toilet - Patient | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
WCPT | Toilet - Patient | FA | WLWA-007 | WALL FINISH: splashback | 1
WCPT | Toilet - Patient | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
WCPT | Toilet - Patient | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
WCPT | Toilet - Patient | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | with privacy latch, emergency release required if inward opening door
WCPT | Toilet - Patient | FA | CLCN-009 | CORNICE: wall trim fixing | 1
WCPT | Toilet - Patient | FFE | FIHR-026 | HOOK: coat | 2
WCPT | Toilet - Patient | FFE | FQCW-016 | BIN: waste, general, 20L | 1
WCPT | Toilet - Patient | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
WCPT | Toilet - Patient | FFE | FIDI-030 | DISPENSER: soap | 1
WCPT | Toilet - Patient | FFE | FIDI-036 | DISPENSER: toilet paper | 1
WCPT | Toilet - Patient | FFE | FIRT-006 | GRAB RAIL: WC | 1
WCPT | Toilet - Patient | FFE | FIRT-008 | GRAB RAIL: WC, hinged | 1 | optional
WCPT | Toilet - Patient | FFE | FIBM-006 | MIRROR: fixed, glass | 1
WCPT | Toilet - Patient | FFE | FQCW-037 | BIN: waste, sanitary | 1
WCPT | Toilet - Patient | FFE | HYBA-011 | BASIN: handwash, Type C | 1
WCPT | Toilet - Patient | FFE | HYWC-006 | TOILET SUITE: patient | 1
WCPT | Toilet - Patient | FFE | HYTP-035 | TAPWARE: mixer, basin, accessible | 1
WCPT | Toilet - Patient | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
WCPT | Toilet - Patient | SE | HYDR-004 | DRAIN: floor waste | 1
WCPT | Toilet - Patient | SE | ELBO-015 | SWITCH: light, single | 1
WCPT | Toilet - Patient | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | + indicator button & light
WCPT | Toilet - Patient | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
WCPU | Toilet - Public | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
WCPU | Toilet - Public | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | tiles are also acceptable
WCPU | Toilet - Public | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
WCPU | Toilet - Public | FA | WLWA-007 | WALL FINISH: splashback | 1
WCPU | Toilet - Public | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
WCPU | Toilet - Public | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
WCPU | Toilet - Public | FA | DWSC-001 | DOOR: 850mm c/o, single leaf, solid core, paint | 1 | with privacy latch, emergency release required if inward opening door
WCPU | Toilet - Public | FA | CLCN-009 | CORNICE: wall trim fixing | 1
WCPU | Toilet - Public | FFE | FIHR-026 | HOOK: coat | 1
WCPU | Toilet - Public | FFE | FQCW-016 | BIN: waste, general, 20L | 1
WCPU | Toilet - Public | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
WCPU | Toilet - Public | FFE | FIDI-030 | DISPENSER: soap | 1
WCPU | Toilet - Public | FFE | FIDI-036 | DISPENSER: toilet paper | 1
WCPU | Toilet - Public | FFE | FIBM-006 | MIRROR: fixed, glass | 1
WCPU | Toilet - Public | FFE | FQCW-037 | BIN: waste, sanitary | 1
WCPU | Toilet - Public | FFE | HYBA-011 | BASIN: handwash, Type C | 1
WCPU | Toilet - Public | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1
WCPU | Toilet - Public | FFE | HYWC-008 | TOILET SUITE: standard | 1
WCPU | Toilet - Public | SE | HYDR-004 | DRAIN: floor waste | 1
WCPU | Toilet - Public | SE | ELBO-015 | SWITCH: light, single | 1
WCST | Toilet - Staff | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5 | to be tiles if tiled floor
WCST | Toilet - Staff | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | tiles are also acceptable
WCST | Toilet - Staff | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
WCST | Toilet - Staff | FA | WLWA-007 | WALL FINISH: splashback | 1
WCST | Toilet - Staff | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
WCST | Toilet - Staff | FA | CLTI-014 | CEILING: plasterboard, water resistant, drop-in tiles, prefinished, 600 x 1200 | 1
WCST | Toilet - Staff | FA | DWSC-001 | DOOR: 850mm c/o, single leaf, solid core, paint | 1 | with privacy latch, emergency release required if inward opening door
WCST | Toilet - Staff | FA | CLCN-009 | CORNICE: wall trim fixing | 1
WCST | Toilet - Staff | FFE | FIHR-026 | HOOK: coat | 1
WCST | Toilet - Staff | FFE | FQCW-016 | BIN: waste, general, 20L | 1
WCST | Toilet - Staff | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
WCST | Toilet - Staff | FFE | FIDI-030 | DISPENSER: soap | 1
WCST | Toilet - Staff | FFE | FIDI-036 | DISPENSER: toilet paper | 1
WCST | Toilet - Staff | FFE | FIBM-006 | MIRROR: fixed, glass | 1
WCST | Toilet - Staff | FFE | FQCW-037 | BIN: waste, sanitary | 1
WCST | Toilet - Staff | FFE | HYBA-011 | BASIN: handwash, Type C | 1
WCST | Toilet - Staff | FFE | HYTP-036 | TAPWARE: mixer, basin, lever action | 1
WCST | Toilet - Staff | FFE | HYWC-008 | TOILET SUITE: standard | 1
WCST | Toilet - Staff | SE | HYDR-004 | DRAIN: floor waste | 1
WCST | Toilet - Staff | SE | ELBO-015 | SWITCH: light, single | 1
WCST | Toilet - Staff | SE | ITSE-038 | READER: access control, general | 1 | optional, may be provided if room located in a public area