## Room Data
Standard Components - Room Data Sheet
Date of issue: 26.04.2021
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
ADLK | ADL Kitchen | 12 | up to 24 hours | up to 2 patients, 1-2 staff |  | Acoustics: sound attenuation level high. | The ADL Kitchen provides an accessible space for the training and assessment of patients in Activities of Daily Living (ADL) tasks. The space is to be functional for a range of patients including wheelchair users. The kitchen can include a combination of fittings and appliances such as a sink, cooktop, oven/grill, microwave and refrigerator to support common ADL training and assessment tasks as per service requirements. Room exhaust to be provided to range hood over cooktop. Additional Considerations: An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. If multiple ADL rooms are combined into an ADL suite nurse call buttons can be rationalised into a central location. The room may be enclosed with a lockable door depending on security requirements. External outlook and access to natural light is desirable.

## Items
Standard Components - Items in room
Date of issue: 26.04.2021
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
ADLK | ADL Kitchen | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent around sink
ADLK | ADL Kitchen | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 8
ADLK | ADL Kitchen | FA | WLWA-007 | WALL FINISH: splashback | 2
ADLK | ADL Kitchen | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
ADLK | ADL Kitchen | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | plasterboard flush set ceiling is also acceptable
ADLK | ADL Kitchen | FA | DWWH-005 | WINDOW | 1 | desirable
ADLK | ADL Kitchen | FA | CLCN-009 | CORNICE: wall trim fixing | 1
ADLK | ADL Kitchen | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 2
ADLK | ADL Kitchen | FFE | DWWD-004 | BLIND: roller | 1
ADLK | ADL Kitchen | FFE | FIJO-131 | CUPBOARD: tall, single door, adj shelves, lockable | 1
ADLK | ADL Kitchen | FFE | FQSM-040 | PEDESTAL: mobile, drawer unit, lockable, 660H | 1 | under bench
ADLK | ADL Kitchen | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 1
ADLK | ADL Kitchen | FFE | FIJO-156 | CUPBOARD: wall mount, single door, adj shelves | 2
ADLK | ADL Kitchen | FFE | FQCW-016 | BIN: waste, general, 20L | 1
ADLK | ADL Kitchen | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to sink
ADLK | ADL Kitchen | FFE | FIDI-030 | DISPENSER: soap | 1 | to sink
ADLK | ADL Kitchen | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
ADLK | ADL Kitchen | FFE | FQCA-046 | KETTLE: electric | 1
ADLK | ADL Kitchen | FFE | FQCA-067 | OVEN: microwave, domestic | 1 | on bench
ADLK | ADL Kitchen | FFE | FQCA-063 | OVEN: electric | 1 | retractable/side opening doors
ADLK | ADL Kitchen | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
ADLK | ADL Kitchen | FFE | FQCA-072 | RANGEHOOD | 1 | over cooktop
ADLK | ADL Kitchen | FFE | FQCA-031 | FRIDGE: domestic, 250L | 1
ADLK | ADL Kitchen | FFE | HYBA-041 | SINK: double, with drainer, half size rinse bowl | 1 | heat lined, warm water
ADLK | ADL Kitchen | FFE | FQCA-080 | TOASTER: domestic | 1
ADLK | ADL Kitchen | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 2
ADLK | ADL Kitchen | FFE | FIJO-205 | CUPBOARD: drawer unit, single drawer, lockable | 6
ADLK | ADL Kitchen | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead also acceptable
ADLK | ADL Kitchen | FFE | FIJO-213 | PANEL: joinery carcass end | 3
ADLK | ADL Kitchen | FFE | FIJO-223 | CUPBOARD: tall, oven cabinet | 1
ADLK | ADL Kitchen | FFE | FQCA-089 | COOKTOP: induction | 1 | controls to be located close to front
ADLK | ADL Kitchen | FFE | FIJO-224 | PANEL: joinery fixed panel | 5 | to sink and cooktop
ADLK | ADL Kitchen | SE | ELPR-006 | RCD: residual current device | 1
ADLK | ADL Kitchen | SE | ELBO-007 | SWITCH: emergency isolation | 1 | to cooktop and oven
ADLK | ADL Kitchen | SE | ELBO-015 | SWITCH: light, single | 1
ADLK | ADL Kitchen | SE | ELGP-109 | GPO: single, wall mounted | 2 | to range hood and fridge
ADLK | ADL Kitchen | SE | ELGP-208 | GPO: double, wall mounted | 4
ADLK | ADL Kitchen | SE | ELGE-001 | DIRECT CONNECTION: power | 2 | to cooktop and oven