## Room Data
Standard Components - Room Data Sheet
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
AHBF | After Hours Blood Fridge | 8 | Available 24 Hours | 1-2 staff, intermittently |  |  | The After Hours Blood Fridge is a discreet area for refrigerated storage of blood and blood products for access by authorised personnel after hours where a 24 hour pathology service is not available. This room is typically located on the periphery of a Pathology unit and may have direct, secure access to the Pathology laboratory with ready access to the transfusion service area. Additional Considerations: Refrigerator temperature monitoring is required and is to be connected to the building management system (BMS). The alarm is to provide alerts where 24-hour monitoring is undertaken, within the unit and/or remotely. Temperature monitoring is to meet ANZSBT guidelines. Additional fridges/freezers or lockable sections within the after hours fridge may be provided to support Pathology laboratory workflows. Provision of smart fridges for blood product tracking and other ICT requirements will be based on local policies and service requirements. Door width to allow for installation of fridge, including considerations for future proofing as equipment is replaced.

## Items
Standard Components - Items in room
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
AHBF | After Hours Blood Fridge | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
AHBF | After Hours Blood Fridge | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
AHBF | After Hours Blood Fridge | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
AHBF | After Hours Blood Fridge | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
AHBF | After Hours Blood Fridge | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
AHBF | After Hours Blood Fridge | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | optional, to laboratory if adjacent
AHBF | After Hours Blood Fridge | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1
AHBF | After Hours Blood Fridge | FA | CLCN-009 | CORNICE: wall trim fixing | 1
AHBF | After Hours Blood Fridge | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
AHBF | After Hours Blood Fridge | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
AHBF | After Hours Blood Fridge | FFE | ITNE-003 | BAR CODING: reader | 1
AHBF | After Hours Blood Fridge | FFE | FIJO-017 | BENCH: 600D, open under | 1 | laboratory grade
AHBF | After Hours Blood Fridge | FFE | MMLA-053 | FRIDGE: blood products, single door | 1 | sections may be lockable; BMS monitored, audible alarm 24hr; size is to suit service requirements
AHBF | After Hours Blood Fridge | FFE | FIJO-213 | PANEL: joinery carcass end | 1
AHBF | After Hours Blood Fridge | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to fridge for temperature monitoring
AHBF | After Hours Blood Fridge | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
AHBF | After Hours Blood Fridge | SE | ITBU-001 | ALARM: equipment monitoring, temperature | 1 | may be integral to fridge
AHBF | After Hours Blood Fridge | SE | ELBO-015 | SWITCH: light, single | 1
AHBF | After Hours Blood Fridge | SE | ELGP-208 | GPO: double, wall mounted | 2
AHBF | After Hours Blood Fridge | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to fridge
AHBF | After Hours Blood Fridge | SE | ITSE-038 | READER: access control, general | 2 | 1 to entry, 1 is optional to laboratory door if provided