## Room Data
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Acoustics | Description / Special Requirements
ANAE-16 | Anaesthetic Preparation Room, 16m² | 16 | Available 24 Hours | 1 patient; 1 parent; up to 3 staff | Acoustic privacy is required | The Anaesthetic Preparation Room is for holding patients on mobile beds or trolleys prior to operative procedures. Local, regional or general anaesthesia may be administered in this area. If serving an Operating Room with cardiac protected power, the Anaesthetic Preparation Room should also have cardiac protected power.

## Items
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | DWGL-051 | DOOR: 1600mm c/o, double leaf, observation panel, solid core/glass, paint | 1 | auto release, delayed action, magnetic lead flap to be installed over glass if laser is in use | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | DWGL-053 | DOOR: 1600mm c/o, double leaf, sliding, observation panel, solid core/glass, paint | 1 | manual release, delayed action | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | DWPR-011 | DOOR PROTECTION: protection plate, to full height | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ELBO-015 | SWITCH: light, single | 3 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 5 | to bench, computer | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ELPR-006 | RCD: residual current device | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIDI-014 | DISPENSER: disposable glove | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIDI-021 | DISPENSER: mask | 1 | to entry | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIHR-019 | BRACKET: for suction bottle | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIHR-062 | RACK: wire basket | 4 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIHR-063 | BRACKET: for all-in-one computer, with keyboard shelf, wall mounted | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 2 | optional | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIJO-212 | BULKHEAD: joinery | 1 | optional, install over joinery when used, plasterboard bulkhead is a suitable alternative | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIJO-213 | PANEL: joinery carcass end | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FIRT-020 | TRACK: ceiling mounted, IV | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FQCW-004 | BIN: waste, clinical, 120L | 1 | optional, dependent on waste management policy | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FQCW-016 | BIN: waste, general, 20L | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | FQSM-003 | PEDESTAL: mobile, wire basket unit, 840H | 3 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1 |  | No | No | Yes | No | Yes | Yes | Yes | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | HYTP-042 | TAPWARE: sensor | 1 |  | No | No | Yes | No | Yes | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITCL-005 | BUTTON: nurse call, orderly / porter | 1 | optional | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | ITCL-018 | CLOCK: clock system slave, digital | 1 | mount flush with wall | Yes | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 |  | No | Yes | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 |  | No | Yes | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 |  | No | Yes | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | ITNE-052 | TELEPHONE: handset, wall mount | 1 |  | No | Yes | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | ITNE-066 | COMPUTER: all-in-one, wall mounted | 1 |  | Yes | Yes | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1 | wording on light to jurisdictional requirements | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | LIFX-003 | LIGHT: illuminated sign, "Laser in Use" | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | LIFX-013 | LIGHT: procedure, ceiling mounted, small | 1 |  | Yes | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | LIFX-016 | LIGHT: pelmet light | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | MEEX-008 | EXHAUST: low level | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | MGAS-062 | OUTLET: suction, on services panel | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MGFP-002 | FLOWMETER: medical air | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MGFP-004 | FLOWMETER: oxygen | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMBE-044 | TROLLEY: patient, transport | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMGE-002 | ADAPTER: suction | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMGE-014 | CHAIR: anaesthetist | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMGE-039 | HANGER: IV, for ceiling track | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMGE-088 | SAFE: dangerous drugs, small | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMGE-100 | SUCTION BOTTLE | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMGE-142 | TROLLEY: sharps bin | 2 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FFE | MMOP-040 | MACHINE: anaesthetic | 0 | space allowance only | Yes | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | recessed | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No
ANAE-16 | Anaesthetic Preparation Room, 16m² | FA | WLWA-023 | WALL FINISH: vinyl, to 2100 AFFL | 1 |  | No | No | No | No | No | No | No | No | No