## Room Data
Standard Components - Room Data Sheet
Date of issue: 26.04.2021
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
AUD-TR | Audiology Testing Room | 14 | 8am to 6pm, daily | 1 patient, 1 visitor/carer, 1-2 staff |  | Acoustic isolation is required | The Audiology Testing Room provides an acoustically treated, soundproof space for clinical assessments to be carried out under controlled conditions. Testing Rooms may be modular proprietary items or constructed in situ and are generally designed as rooms within rooms in order to isolate the Testing Room from the rest of the building as much as possible. All designs should be reviewed and confirmed by an acoustic design expert. Additional Consideration: The Audiology Testing Room currently indicates a layout (with an adjacent control room) that is suitable for paediatric testing. Layouts will vary to suit service requirements and the provided layout will need to be altered to suit services for older children and adults. The audiometric booth is to be ventilated and connected to air-conditioning, power and data. Noise from mechanical services is to be minimised with air-conditioning to be designed for low noise operation to meet ambient noise requirements for testing. The floor of the Treatment Room should be flush with the floor of the Control Room, or the adjacent corridor where a Control Room is not provided, for access by patients in wheelchairs.

## Items
Standard Components - Items in room
Date of issue: 26.04.2021
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
AUD-TR | Audiology Testing Room | FA | FLCP-006 | FLOOR FINISH: carpet, carpet tiles, 90/10 wool/nylon | 1
AUD-TR | Audiology Testing Room | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1
AUD-TR | Audiology Testing Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | acoustic seals to suit testing requirements
AUD-TR | Audiology Testing Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
AUD-TR | Audiology Testing Room | FFE | FQBS-012 | CHAIR: child | 1
AUD-TR | Audiology Testing Room | FFE | FQBS-014 | CHAIR: client | 2
AUD-TR | Audiology Testing Room | FFE | FQBS-039 | LOUNGE: 1 seat | 1
AUD-TR | Audiology Testing Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 1
AUD-TR | Audiology Testing Room | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 1
AUD-TR | Audiology Testing Room | FFE | FQDW-001 | DESK: adj height, 1500W x 750D x 750H nom. | 1
AUD-TR | Audiology Testing Room | FFE | FIHR-029 | HOOK: equipment | 2 | for headphones
AUD-TR | Audiology Testing Room | FFE | FQTA-003 | TABLE: children's play | 1
AUD-TR | Audiology Testing Room | FFE | MMAH-005 | BOOTH: audiometric | 1 | with one way window (non-reflective)
AUD-TR | Audiology Testing Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
AUD-TR | Audiology Testing Room | FFE | MMAH-036 | CONSOLE: control, audiometric booth | 1
AUD-TR | Audiology Testing Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
AUD-TR | Audiology Testing Room | FFE | ITAV-061 | SPEAKER: audiology | 2 | wall mounted, calibrated for testing space
AUD-TR | Audiology Testing Room | FFE | MMAH-043 | DEVICE: visual reinforcement audiometry | 2 | wall mounted, may be provided as a screen or puppet box, inclusion is dependent on service requirements
AUD-TR | Audiology Testing Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
AUD-TR | Audiology Testing Room | SE | ELPR-006 | RCD: residual current device | 1
AUD-TR | Audiology Testing Room | SE | ELGP-208 | GPO: double, wall mounted | 6 | provision of emergency power to be determined based on service requirements
AUD-TR | Audiology Testing Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
AUD-TR | Audiology Testing Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
AUD-TR | Audiology Testing Room | SE | MMSP-032 | SERVICES PANEL: instrument jack | 4 | optional, as required for remote operation of equipment