## Room Data
Standard Components - Room Data Sheet
Date of issue: 12.07.2023
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
BIRM-A | Birthing Room, Without Pool | 34 | 24 Hours | 1 mother; 1-2 support persons; up to 4 staff |  | Sound attenuation level high, Acoustic privacy is required. | The Birthing Room - Without Pool provides area and amenities for the entire birthing process including mother's preparation, active birthing exercises, relaxation and analgesia during labour, birthing, infant resuscitation, postnatal recovery and observation. The room includes an active birthing/bed zone, an area to accommodate a support person, a staff observation and write up area, an infant resuscitation area, an equipment store, and a hand hygiene area with type A handwashing basin. Additional Considerations: The room will have direct access to an Ensuite - Birthing. Air should not be recirculated to other areas from the Birthing Room. Negative pressure air conditioning is required for clearance of anaesthetic gases. Clinical items such as medical gases and services as well as neonatal resuscitation equipment may be concealed to support the provision of a calm environment. Even while concealed, items are to be within easy reach of the active birthing zone. The availability of proprietary items for in-room infant resuscitation may change and this can inform the use of built in or mobile equipment. Overall dimensions and configuration of storage for the resuscitaire equipment are to be evaluated based on the selection of equipment and it is recommended that a test-fit is conducted, ensuring appropriate clearances are provided for emergency procedures to be undertaken, in order to confirm room layout requirements. Equipment storage to be configured to suit service requirements including models for sharing equipment between multiple rooms. Storage requirements within the room will vary between facilities depending on models of centralising or decanting and restocking consumables. Configuration and extent of storage to be determined based on service requirements. Various supportive fixtures for active birthing are to be considered, for example, a wall mounted ladder, bench or mantlepiece for support and leaning, or ceiling mounted hooks for supportive fabrics to be hung. Consider mounting heights and providing support for various postures at multiple heights. A mother's entire weight may be supported on this item, therefore the strength of the item and fixings, including in-wall support, to be equivalent to heavy duty grab rails. A fixed workstation for staff write-up and observation has been shown. Sufficient space for dual monitors may be provided as required to suit operational, EMR and ICT requirements. A workstation on wheels is also acceptable. An examination light is required for suturing and small procedures. It is preferred that the light be as unobtrusive as possible and be able to be pushed away so that it is not dominating the area. Lighting options to be provided, including dimmable room lighting, reading lights and night lights to provide flexible lighting across the space. Coloured lighting and calming light artworks may be considered by projects. Systems, including portable items, to facilitate the selection and control of music by the mother are to be considered. Acoustics treatment may be required to support this without disruption to adjacent rooms. Provision of a telephone for patient use will be dependent on the operational model for transferring calls. Natural light and outlook essential. Window sill to be low enough to see view while laying in bed. Treatment, to windows to allow blinds to be kept open as desired but privacy to be maintained, may need to be considered depending on overlook from nearby buildings.

## Items
Standard Components - Items in room
Date of issue: 12.07.2023
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
BIRM-A | Birthing Room, Without Pool | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
BIRM-A | Birthing Room, Without Pool | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
BIRM-A | Birthing Room, Without Pool | FA | WLPR-006 | WALL PROTECTION: corner guards | 3
BIRM-A | Birthing Room, Without Pool | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
BIRM-A | Birthing Room, Without Pool | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
BIRM-A | Birthing Room, Without Pool | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FA | WLPR-002 | WALL PROTECTION: bed protection wall panel, composite, laminate finish | 1
BIRM-A | Birthing Room, Without Pool | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
BIRM-A | Birthing Room, Without Pool | FA | DWWH-005 | WINDOW | 2
BIRM-A | Birthing Room, Without Pool | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | with hold open device to support bed movement as required
BIRM-A | Birthing Room, Without Pool | FA | DWGE-015 | DOOR: concertina folding leaf | 1
BIRM-A | Birthing Room, Without Pool | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
BIRM-A | Birthing Room, Without Pool | FFE | MGFP-002 | FLOWMETER: medical air | 2 | 1 to bedhead, 1 to resuscitaire
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-001 | BASSINET | 1 | optional
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | 1 to bedhead cupboard, 1 to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-016 | BRACKET: for sharps bin | 1 | mounted on trolley
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | FFE | DWWD-004 | BLIND: roller | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-156 | CUPBOARD: wall mount, single door, adj shelves | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-019 | DISPENSER: gown/apron | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-014 | DISPENSER: disposable glove | 2 | 1 to basin, 1 to room
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQSN-051 | LOCKER: 2 drawers, open under, lockable, mobile | 1 | type and extent of storage for patient property to suit operational policies
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-044 | LOUNGE: 2 seat, sofa bed | 1
BIRM-A | Birthing Room, Without Pool | FFE | FQTA-013 | TABLE: overbed | 1
BIRM-A | Birthing Room, Without Pool | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | 1 to bedhead, 1 to resuscitaire
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-011 | MIXER: oxygen and nitrous oxide | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-194 | SHELF: laminate | 2 | to equipment store cupboard, type and quantity of storage options to be determined based on service requirements
BIRM-A | Birthing Room, Without Pool | FFE | FQBS-051 | STOOL: mobile, adj height | 1 | for suturing and small procedures
BIRM-A | Birthing Room, Without Pool | FFE | MMGE-002 | ADAPTER: suction | 2 | 1 to bedhead, 1 to resuscitaire
BIRM-A | Birthing Room, Without Pool | FFE | MMGE-100 | SUCTION BOTTLE | 2 | 1 to bedhead cupboard, 1 to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | FFE | FQSM-016 | TROLLEY | 1 | to equipment store cupboard
BIRM-A | Birthing Room, Without Pool | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1
BIRM-A | Birthing Room, Without Pool | FFE | MMBE-005 | BED: patient, birthing, electric | 1
BIRM-A | Birthing Room, Without Pool | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
BIRM-A | Birthing Room, Without Pool | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
BIRM-A | Birthing Room, Without Pool | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | to staff write-up, dual monitors may be provided as required to suit operational, EMR and ICT requirements; a workstation on wheels is also acceptable
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-015 | RESUSCITAIRE: infant, mobile | 1 | Mobile resuscitaire shown. Item may be wall mounted where suitable product is available.
BIRM-A | Birthing Room, Without Pool | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
BIRM-A | Birthing Room, Without Pool | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | optional, to patient locker
BIRM-A | Birthing Room, Without Pool | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-129 | CUPBOARD: tall, single door, adj shelves | 1 | type and extent of storage to suit service requirements
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-162 | CUPBOARD: tall, single door, to conceal services | 1
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-209 | CUPBOARD: tall, single door, adj shelves and clothes rail, recessed | 1 | wardrobe, type and extent of storage for patient property to suit operational policies
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-212 | BULKHEAD: joinery | 3
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-213 | PANEL: joinery carcass end | 1 | below FIJO-156 to separate dirty linen trolley from hand wash, storage for clean and dirty linen to suit local operational policies
BIRM-A | Birthing Room, Without Pool | FFE | FIHR-068 | BRACKET: for alcohol based hand rub, attached to bed | 1
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-031 | DEVICE: patient, fetal, cardiotocography (CTG), on trolley | 1 | to equipment store cupboard
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-043 | SUPPORT FIXTURE: birthing, hook, ceiling mounted | 1 | optional, provision dependent on birthing aids/supportive equipment to be used
BIRM-A | Birthing Room, Without Pool | FFE | FIJO-233 | CUPBOARD: tall, infant resuscitaire | 1 | Configuration of joinery to suit resuscitare equipment selected, including consideration of maneuverability and provision of sufficient space for emergency resuscitation procedures. Confirm design with users at project level.
BIRM-A | Birthing Room, Without Pool | FFE | MMMA-041 | SUPPORT FIXTURE: birthing, ladder, wall mounted | 1
BIRM-A | Birthing Room, Without Pool | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to staff write-up for wall mounted phone
BIRM-A | Birthing Room, Without Pool | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | SE | ELPR-006 | RCD: residual current device | 3 | 2 to bedhead, 1 to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELPR-006 | RCD: residual current device | 2 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment system, may be replaced with data outlet depending on ICT policies
BIRM-A | Birthing Room, Without Pool | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to patient entertainment system
BIRM-A | Birthing Room, Without Pool | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | with indicator button and light, to bedhead
BIRM-A | Birthing Room, Without Pool | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | LIFX-011 | LIGHT: reading light | 1 | to bedhead
BIRM-A | Birthing Room, Without Pool | SE | LIFX-017 | LIGHT: night light | 1 | to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ELBO-015 | SWITCH: light, single | 1 | to bedhead MSP for reading light
BIRM-A | Birthing Room, Without Pool | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | with indicator button and light, to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light, to bedhead
BIRM-A | Birthing Room, Without Pool | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level for patient bed
BIRM-A | Birthing Room, Without Pool | SE | ELGP-208 | GPO: double, wall mounted | 3 | 2 to staff write-up, 1 to lounge
BIRM-A | Birthing Room, Without Pool | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
BIRM-A | Birthing Room, Without Pool | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
BIRM-A | Birthing Room, Without Pool | SE | ITSE-006 | BUTTON: duress, fixed | 1 | to staff write-up
BIRM-A | Birthing Room, Without Pool | SE | MMSP-026 | SERVICES PANEL: wall mounted | 2 | 1 to bedhead, 1 inside bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELBO-012 | SWITCH: light, dimmer | 2 | 1 to entry door, 1 to staff write up for general room lights
BIRM-A | Birthing Room, Without Pool | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | 1 to bedhead, 1 to bedhead cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELGP-122 | GPO: emergency power, single, on services panel | 6 | to infant resuscitaire cupboard
BIRM-A | Birthing Room, Without Pool | SE | ELGP-122 | GPO: emergency power, single, on services panel | 8 | 5 to bedhead, 3 to bedhead cupboard