## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
CLAB | Catheter Laboratory Procedure Room | 42 |  | 1 patient, up to 6 staff | Available 24 Hours | Acoustic privacy is required | Procedure room where Cardiac Catheter procedures and electrophysiology studies are undertaken in a controlled environment. A Control room will be located adjacent and may serve more than one procedure room. A clinical scrub area should be located immediately adjacent to the room staff entry. Acoustics: Acoustic privacy is required Cardiac Protected Electrical Area Additional Design Considerations: - Radiation shielding as advised by Radiation Consultant - Unistrut Ceiling able to carry moving loads may be required - Cable ducts from Control Room and Computer/Equipment room may be required - &apos;X-Ray in Use&apos; lights are required outside each entry door - Power to Angiography equipment should be on emergency supply

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
CLAB | Catheter Laboratory Procedure Room | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 4020 | Wall protection | 1 | equal to 'Acrovyn' | Composite | Prefinished PVC | Crash rail at 900 AFFL | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 5010 | Ceiling - Plasterboard | 1 | at 3000 mm AFFL | Plasterboard | Paint, washable | Flush set, suspended | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 6015 | Cornice | 1 |  | Plasterboard | Paint | Flush coved between wall & ceiling | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 7025 | Observation | 1 | to control room | Lead glass | Clear | Internal window | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 8010 | Door protection | 1 | equal to 'Acrovyn' | Composite | Prefinished PVC | Protection plate to 900 AFFL | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 8040 | Door protection | 1 | equal to 'Acrovyn' | Composite | Prefinished PVC | To door frame | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 9005 | Door - 910mm | 1 | staff entry/ exit | Solid Core/ Lead glass | Paint | Single leaf, observation panel, lead shielding | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | Fabrics | 9130 | Door - 1600mm | 1 | patient entry/ exit, 1600mm clear opening | Solid Core | Paint | 1 1/2 leaf, with lead shielding | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FE | FE-32650 | Light - operating, single, 550 diam nominal | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FE | FE-36050 | Monitor - cardiac | 3 | on ceiling support frame |  |  |  | Yes | Yes | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FE | FE-66460 | IMG - Catheter Lab imaging unit | 1 | with imaging & cardiac monitors |  |  |  | Yes | Yes | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FE | FE-66464 | IMG - Catheterisation table | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FE | FE-9700 | Clock - analogue, electric | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-1050 | Bench - laminate | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-1465 | Bracket - suction bottle | 2 | below services panel |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-150 | Air flowmeter | 1 | services panel |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-16250 | Oxygen flowmeter | 2 | services panel |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-21950 | Suction adapter | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-22000 | Suction bottle | 2 | on bracket |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-4410 | Cupboard - under bench | 2 | adjustable shelves |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-4430 | Cupboard - drawer unit | 1 | under bench |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-4493 | Cupboard - catheter hanging | 1 | sliding glass doors, shelving to one side |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | FF | FF-5824 | Waste bin - general | 1 | under bench |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-1006 | Voice / Data outlet: double | 3 | 1 outlet on each panel for telephone |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-1040 | Intercom | 1 | handsfree to control room |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-11000 | Wall panel | 2 | patient services panels |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-11060 | Special | 1 | to angiography table |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-12030 | Lead lining | 1 | to walls and doors as specified |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-5000 | Airconditioning | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-6001 | General: colour corrected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-6110 | Downlights: dimmable | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-6200 | Light switches | 3 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-7000 | Oxygen (O2) | 2 | patient services panel on table |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-7010 | Nitrous Oxide (N2O) | 1 | patient services panel on table |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-7020 | Medical Air (MA) | 1 | patient services panel on table |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-7040 | Suction | 2 | patient services panel on table |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-7070 | Scavenge | 1 | patient services panel on table |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-8010 | Staff/ Nurse assist call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-8080 | Emergency call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-9000 | GPO: Single | 7 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-9001650 | GPO:Double RCD | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-9010 | GPO: Emergency power, single (V) | 6 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-9011 | GPO: Emergency power, double (H) | 3 | 2 above bench; 1-wall panel |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-9015 | GPO: Cleaner | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-9030 | Cardiac protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CLAB | Catheter Laboratory Procedure Room | SE | S-9060 | GPO: Special: for equipment | 1 | 20 amp, patient services panel |  |  |  | No | No | No | No | No | No | No | No | No