## Room Template Data
Room Data Sheet - File for Import
2026-10-29
No. of Standard Components = 1
Code | Template Name | Standard Component Set | Standard Area | Modeled Ceiling Height | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Panorama - URL for Project Use | Briefing - Hours of Operation | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experience system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statutory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magnetic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: low | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium | Performance Requirements - Acoustics - NOISE SENSITIVITY: high | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | 10 | 2700 |  | 1 | 29.10.2026 |  | Business hours (service may operate with extended hours, on weekends and/or for emergency and urgent care situations) | 1-2 staff intermittently | The Decontamination Room - Satellite, On-unit Reprocessing provides the space and facilities for the decontamination, rinsing and processing of reusable medical devices (RMDs) in a washer disinfector, including oiling and lubrication of handpieces as required, within a clinical unit. | Planning Context ~ Room must be configured to meet the requirements of AS/NZ 5369, providing clear circulation and support the unidirectional flow of dirty instruments through the decontamination process, including manual cleaning, ultrasonic cleaning, rinsing, loading into baskets or cassettes and processing through the washer disinfector. ~ Where reprocessing equipment does not provide pass-through capability, maintain effective segregation between dirty and clean activities through the room layout and operational workflow. ~ The Decontamination Room may be physically separated from, or open to, the Packing and Sterilizing Room. The configuration should be determined at project level based on the reprocessing workflow, equipment configuration and applicable requirements. Where the rooms are combined, dirty and clean zones and associated PPE requirements must be clearly identified. Where the rooms are enclosed and separated and the washer disinfector is not pass-through, an openable transfer hatch/window should be provided to enable racks to be transferred to bench space within the Packing and Sterilizing Room. ~ Consider staff circulation between the Decontamination and Sterilizing Rooms, including requirements for hand hygiene and changing PPE before moving from dirty to clean activities. ~ Provision of automated, hands-free door operation for delivery and movement of instruments may be considered at project level where this supports the selected workflow and local IPC requirements. ~ The workflow may include oiling or lubrication of handpieces and other RMDs following disinfection, where required. Furniture, Fixtures and Equipment (FF&E) Selection ~ Confirm the washer disinfector type, including pass-through, standing or benchtop configurations, based on local jurisdictional policies, operational service requirements and the volume of instruments and number of DIN trays to be processed. ~ Requirements for drying manually cleaned instruments are to be determined at project level based on equipment selection, and service requirements. Information and Communication Technology (ICT) ~ Confirm any in room requirements for instrument tracking, including location, process and configuration, in line with AS/NZ 5369, local policies and service requirements. Waste Management ~ Confirm the requirement and size of a sharps bin based on local jurisdictional policies and service requirements. Sharps disposal primarily occurs within the dental surgery as part of pre-treatment; however, a small sharps bin may be required for broken instruments identified during examination and cleaning. Engineering Services ~ Confirm requirements for RO water or water of suitable quality to decontamination sinks and washer disinfector in line with AS/NZ 5369, local policies and service requirements. Security ~ Where there is potential for unauthorised access to staff-only areas, consider access control and staff safety requirements, including swipe card access and/or a fixed duress button. | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE

## Items In Room Template
Room Data Sheet - File for Import
2026-10-29
Code | Template Name | Standard Component Set | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | to room side
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | to corridor side, extent dependent on movement of beds and mobile equipment in adjacent area
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | D+W | AFDPR-056.01 | DOOR FRAME PROTECTION: full wrap, to 900H | 1 | 3 | Internal - Door 1 | optional, provision and extent dependent on frame material/finish and movement of beds and mobile equipment in adjacent area
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | CLCN-031 | CORNICE: square set | 1 | 1 | Ceiling Cornice
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | CLFI-002 | CEILING FINISH: paint, clinical areas | 1 | 1 | Ceiling Finish 2
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | CLFS-013 | CEILING: flush set, suspended, moisture resistant | 1 | 1 | Ceiling Finish 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | D+W | DOHI-043.01 | DOOR: hinged, 1 leaf, 900 clear opening, solid, wide vision panel | 1 | 1 | Internal - Door 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | ELGP-101 | GPO: single, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | ELGP-301 | ISOLATOR: one phase, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | ELSW-001 | SWITCH: light | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FIDI-101 | DISPENSER: PPE, disposable masks, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FIDI-106 | DISPENSER: PPE, disposable gloves, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FIDI-231 | DISPENSER: paper towel, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FIDI-241 | DISPENSER: alcohol-based hand rub, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FIDI-256 | DISPENSER: soap, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FIHR-111 | BRACKET: sharps bin, wall mounted | 2 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | FLVY-112 | FLOOR FINISH: vinyl, seamless, non-slip, safety | 1 | 1 | Floor Finish 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FQWS-001 | BIN: sharps, clinical, extra small | 2 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FQWS-051 | BIN: general waste, 20L | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | FQWS-233 | BIN: clinical waste, 50L | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | HYBA-121 | BASIN: type B, handwashing | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | HYGE-181 | VALVE: backflow prevention, wall recessed | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | HYTP-067 | TAPWARE: basin, tap set, wall mounted, lever handles | 1 | 1 |  | selection of basin and tapware should be coordinated to ensure compatibility
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | HYTP-172 | TAPWARE: sink, tap set, fixture mounted, lever handles | 2 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | HYTP-231 | TAPWARE: sink, spray hose, hand-held, connection and bracket on fixture | 2 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | HYTP-431 | OUTLET: water, cold | 3 | 1 |  | 2 to sinks, 1 to basin
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | HYTP-432 | OUTLET: water, hot | 2 | 1 |  | to sinks
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | HYTP-433 | OUTLET: water, warm | 1 | 1 |  | to basin
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | JOBE-001 | BENCH: 450D, laminate | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | SER | MGAS-153 | OUTLET: compressed air, wall mounted | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | MMDN-102 | PEDESTAL: dental, mobile, 660H nominal, 5 equal drawers | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | MMDN-181 | DEVICE: dental, handpiece oiling | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | MMIC-132 | ULTRASONIC CLEANER: instrument, benchtop | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | MMIC-187 | WASHER DISINFECTOR: batch, single door | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FFE | MMIC-444 | WORKSTATION: final rinse, straight, adjustable height, 900H nom, 750D, with sink, reverse osmosis | 1 | 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | D+W | WIOP-017 | WINDOW: operable, hung, internal, single glazed, sill at 900H | 1 | 1 | Internal - Window 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | WLFI-002 | WALL FINISH: paint, clinical areas | 1 | 1 | Wall Finish 1
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | WLFI-011.06 | WALL FINISH: vinyl, to 2100 AFFL | 1 | 1 | Wall Finish 2
DECON-SAT2 | Decontamination Room - Satellite, On-unit Reprocessing | Oral Health | FIN | WLFI-101 | SPLASHBACK: vinyl | 1 | 1 | Wall Finish 3