## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.06.2021
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
DENSR-1 | Dental Surgery, Type 1 | 14.5 | Available 24 Hours | 1 patient, 1-2 staff, 1-2 visitors |  | Sound attenuation level high, Acoustic privacy is required. | The Dental Surgery provides for dental consultation, examination and treatment for patients. This Standard Component provides a configuration for switching between a a right-handed and left-handed set-up and should not be mirrored. Dental surgery layouts will vary and will be dependent on existing/proposed workflows and the equipment procured. Additional Design Considerations: Services will be required for the dental chair unit; these will be dependent on the chair model and supplier and include dental air, dental suction, filtered water, power and drainage. Facilities for viewing physical x-ray films brought in by patients may be provided in a central location as required. An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. Radiation shielding to be advised by Radiation Consultant. An external outlook is desirable.

## Items
Standard Components - Items in room
Date of issue: 28.06.2021
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
DENSR-1 | Dental Surgery, Type 1 | FA | FLVY-004 | FLOOR FINISH: vinyl, seamless, coved, non-slip, safety | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
DENSR-1 | Dental Surgery, Type 1 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 7
DENSR-1 | Dental Surgery, Type 1 | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
DENSR-1 | Dental Surgery, Type 1 | FA | WLWA-007 | WALL FINISH: splashback | 1
DENSR-1 | Dental Surgery, Type 1 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
DENSR-1 | Dental Surgery, Type 1 | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
DENSR-1 | Dental Surgery, Type 1 | FA | DWWH-005 | WINDOW | 1 | desirable
DENSR-1 | Dental Surgery, Type 1 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
DENSR-1 | Dental Surgery, Type 1 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | DWWD-007 | BLIND: venetian, integral | 1 | to external window if provided
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-019 | BRACKET: for suction bottle | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-016 | BRACKET: for sharps bin | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1 | below sink
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 2 | optional, extent of storage dependent on service requirements and approach to consumables storage
DENSR-1 | Dental Surgery, Type 1 | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-014 | DISPENSER: disposable glove | 2 | 1 x to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
DENSR-1 | Dental Surgery, Type 1 | FFE | FIDI-021 | DISPENSER: mask | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MGFP-004 | FLOWMETER: oxygen | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FQBS-051 | STOOL: mobile, adj height | 2
DENSR-1 | Dental Surgery, Type 1 | FFE | MMGE-002 | ADAPTER: suction | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMGE-100 | SUCTION BOTTLE | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-016 | TROLLEY: dental | 2
DENSR-1 | Dental Surgery, Type 1 | FFE | HYBA-010 | BASIN: handwash, Type B | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | ITNE-063 | DISPLAY SCREEN: for computer | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | optional, to sink if provided
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-002 | CHAIR: dental, electric | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-018 | X-RAY UNIT: dental, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-020 | LIGHT: dental, ceiling mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 2
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-099 | CAC: cable access cap | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-212 | BULKHEAD: joinery | 2 | optional, over wall mounted joinery if provided; plasterboard bulkhead is also acceptable
DENSR-1 | Dental Surgery, Type 1 | FFE | FIJO-213 | PANEL: joinery carcass end | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | HYBA-079 | SINK: single, dental | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-091 | RACK: lead apron, single | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-022 | CONSOLE: dentist, chair mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-021 | CONSOLE: dental assistant, chair mounted | 1
DENSR-1 | Dental Surgery, Type 1 | FFE | MMDN-025 | MONITOR: dentist, chair mounted | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | FFE | FIHR-088 | BRACKET: for monitor, desk mounted, articulated, adj height | 1 | articulated with smooth action, reach to be sufficient for left handed and right handed setup
DENSR-1 | Dental Surgery, Type 1 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ELPR-006 | RCD: residual current device | 2
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | SE | HYTP-008 | OUTLET: cold water | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | HYDR-001 | DRAIN: special requirements | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | LIFX-008 | LIGHT: task light, built in | 2 | optional, under wall mounted joinery if provided
DENSR-1 | Dental Surgery, Type 1 | SE | ELBO-015 | SWITCH: light, single | 1 | 1 x general room lighting, 2 x optional, to task lighting if provided
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-062 | OUTLET: suction, on services panel | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | optional, depending on service requirements
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-208 | GPO: double, wall mounted | 6
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to x-ray unit
DENSR-1 | Dental Surgery, Type 1 | SE | ITSE-006 | BUTTON: duress, fixed | 1
DENSR-1 | Dental Surgery, Type 1 | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 1
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-020 | OUTLET: medical air (MA) | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | MGAS-090 | OUTLET: vacuum | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | ITIN-010 | OUTLET: data, single RJ45 | 1 | connection only, to dental chair, dependent on service requirements and vendor specifications
DENSR-1 | Dental Surgery, Type 1 | SE | ELGE-004 | DIRECT CONNECTION: power, emergency | 1 | to dental chair