## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Briefing
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics | Description / Special Requirements
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
ENS-BA | Ensuite - Bariatric | 7 |  | 1 patient; 1-2 staff assisting | 24 Hours | Sound attenuation level high, Acoustic privacy is required. | An Ensuite - Bariatric is designed to meet the needs of bariatric (obese or morbidly obese) patients allowing the patient to wash, shave, groom themselves, shower and use the toilet, either independently or with assistance. The Ensuite - Bariatric is associated with the 1 Bed Room - Bariatric. The area and layout must accommodate two assisting nurses and a shower commode with bariatric capability. The Ensuite and Bedroom will need to be planned as an integral suite so the use of a ceiling mounted hoist is optimised across the two rooms. The choice of Ensuite will depend on the patient type, the Operational Policy and floor layout of the facility. Acoustics: Sound attenuation level - High. Acoustic privacy required. Body protected electrical area Additional Design Requirements: - The shower floor must not have a raised hob - The entrance is to be flush with the adjoining flooring - The door is to be fitted with escape hardware

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
ENS-BA | Ensuite - Bariatric | FE | FE-47400 | Shower - hand-held | 1 |  |  |  |  | No | No | Yes | No | Yes | Yes | Yes | No | No
ENS-BA | Ensuite - Bariatric | FE | FE-5603 | Basin - ensuite | 1 |  |  |  |  | No | No | Yes | No | Yes | Yes | Yes | No | No
ENS-BA | Ensuite - Bariatric | FE | FE-5602 | Toilet suite - bariatric | 1 |  |  |  |  | No | No | Yes | No | Yes | No | Yes | No | No
ENS-BA | Ensuite - Bariatric | FE | FF-39155 | Patient lifter, ceiling mounted | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FE | FE-40406 | Tapware - shower mixer | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-13700 | Mirror | 1 | over basin |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-19100 | Waste bin - sanitary | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-20603 | Soap / shampoo holder - recessed | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-23700 | Towel rail | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-3400 | Coathook | 2 | 1 @ 1200 AFFL for wheelchair users |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-4800 | Curtain track - shower (90deg) | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-5300 | Curtain - shower | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-6154 | Dispenser - paper/ hand towel | 1 | optional |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-6156 | Dispenser - soap | 2 | optional |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-6160 | Dispenser - toilet paper | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-8450 | Grab rail - Shower | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-8500 | Grab rail - WC, hinged | 2 | optional |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-5673 | Chair - shower, bariatric | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | FF | FF-5824 | Waste bin - general | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-9001 | GPO: Double | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE-Req | S-9040 | Body protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-4033 | Floor waste | 2 | 1 near wc/basin, 1 for shower |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE-Req | S-5015 | Ventilation | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-5015 | Ventilation | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE-Req | S-5019 | Exhaust | 1 | for heat/ steam |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-5019 | Exhaust | 1 | for heat/ steam |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE-Req | S-6000 | General fluorescent | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-6000 | General fluorescent | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-6200 | Light switches | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-8000 | Patient/ Staff call | 2 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-8070 | Patient call pendant | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-8080 | Emergency call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | SE | S-7000 | Oxygen (O2) | 1 | optional |  |  |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 3070 | Wall finish | 1 | to wet areas; tiles also acceptable | Vinyl | Prefinished | Wall vinyl full height | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 2525 | Wall protection | 1 |  |  |  | Bed protection wall panel | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 5023 | Ceiling - Plasterboard | 1 |  | Plasterboard | Prefinished | Water resistant, drop in tiles, 600x1200 | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 9110 | Door - 1400mm | 1 | 1200mm clear opening, outward opening, privacy latch, emergency release | Solid Core | Paint | 1 1/2 leaf | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 2530 | Door protection | 1 |  | Composite | Prefinished PVC | Protection plate to 900 AFFL | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 2529 | Door protection | 1 |  | Composite | Prefinished PVC | To door frame | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 6030 | Cornice | 1 |  | Aluminium | Powdercoat | 24mm shadow angle | No | No | No | No | No | No | No | No | No
ENS-BA | Ensuite - Bariatric | Fabrics | 2000 | Skirting | 1 | or tiles for a tiled floor | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No