## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.06.2021
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
LINAC | Linear Accelerator Treatment Room | 165 | 8am to 6pm, Monday – Friday (extended hours, weekends and/or emergency access may be required) | 1 patient, 1-4 staff |  | Acoustic privacy is required | The Linear Accelerator Treatment Room provides space and facilities for the treatment of patients requiring external beam radiation therapy (EBRT). A collocated Control Room is required for the remote operation of equipment. Cable pathways for services between the Control Room, Modulator Room and Treatment Room must be established and coordinated prior to construction and developed based on vendor and physics requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access. The Treatment Room must be adequately shielded to ensure that radiation doses to surrounding areas do not exceed radiation dose limits and satisfy design/dose constraints, including the provision of a maze entry or shielded door. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. The modulator may need to be accommodated within the Treatment Room. If so, this will be concealed by a wall, with a door for access. Depending on the quality of the power supply at a site, a power conditioner may be required and is typically located with the modulator. Additional noise from this equipment should also be accounted for with acoustic treatment. Visibility of the Treatment Room from exit and via observation cameras is required to ensure all persons are clear of the room before treatment commences. A last person out button is to be provided, as well as a sensor gate and/or physical gate at the entry of the maze. A Beam-On/Off light system is required at the entry to be illuminated when treatment begins/ends. Additional Considerations: Additional medical gases may be provided (including additional oxygen, additional suction, nitrous oxide and scavenge) depending on clinical requirements such as the provision of anaesthetic services. Acoustic requirements will vary depending on the equipment procured and its proximity to the patient treatment area. For patient and staff comfort, acoustics should be considered to reduce the impact of noise from the equipment. Dimmable lights are to be provided and are to be placed to avoid glare to patients. Back lit ceiling artwork may be installed for patient comfort and speakers and connections should be provided to allow music to be played and controlled individually for each bunker. Air conditioning is to be provided to control the room temperature and humidity to meet the equipment suppliers; requirements; high equipment loads shall be accounted for. Air conditioning equipment shall be located outside of this room with ductwork to be installed to follow the entry maze with any penetrations considered to ensure radiation protection. Where possible, it is preferred that ductwork in the ceiling is designed to the perimeter of room for future flexibility to provide ceiling mounted equipment. Cooling shall be provided to the linear accelerator equipment in the form of chilled water to meet the suppliers requirements, independent chiller equipment is preferred. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. Small UPS units are normally required as a minimum to support control and operating equipment. In some cases, these are supplied with the medical equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. A mix of respiratory gating, surface guidance, motion management and/or patient position technologies may be provided depending on clinical/service requirements. The provision of these technologies is to be confirmed by the project and may have additional requirements including additional power/data, water cooling, heat load, possible increased set down, increased number of trenches/conduits, space requirements for computer cabinets in control room and additional penetrations through the suspended ceiling. Requirements to be confirmed with vendors. Allow sufficient distance between the isocentre of the machine and primary shielding wall for the flexibility to provide specialised treatments, e.g. Total Body Irradiation (TBI), depending on service requirements.

## Items
Standard Components - Items in room
Date of issue: 28.06.2021
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
LINAC | Linear Accelerator Treatment Room | FA | FLVY-001 | FLOOR FINISH: vinyl, seamless, coved, anti-static | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
LINAC | Linear Accelerator Treatment Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 18
LINAC | Linear Accelerator Treatment Room | FA | WLPR-006 | WALL PROTECTION: corner guards | 9
LINAC | Linear Accelerator Treatment Room | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 14
LINAC | Linear Accelerator Treatment Room | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 16
LINAC | Linear Accelerator Treatment Room | FA | CLFS-005 | CEILING: plasterboard, flush set, suspended, paint, scrub resistant | 1
LINAC | Linear Accelerator Treatment Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 | inset area above machine as requried
LINAC | Linear Accelerator Treatment Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
LINAC | Linear Accelerator Treatment Room | FFE | DWGE-013 | GATE: single leaf, steel, prefinished | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-016 | BRACKET: for sharps bin | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-006 | DISPENSER: cleaning wipes | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 3
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-140 | CUPBOARD: underbench, single door, adj shelves | 1
LINAC | Linear Accelerator Treatment Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-110 | TROLLEY: clean linen, small | 1
LINAC | Linear Accelerator Treatment Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1
LINAC | Linear Accelerator Treatment Room | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-002 | ADAPTER: suction | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
LINAC | Linear Accelerator Treatment Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
LINAC | Linear Accelerator Treatment Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITCL-012 | CAMERA: observation | 6 | 1 x optional, for total body irradiation or visualisation of additional monitoring equipment, pan, tilt and zoom (PTZ) capability is required, quantity to be based on equipment and service requirements
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | workstation on wheels is also acceptable
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-004 | LINEAR ACCELERATOR | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 2 | for patient's own music, individual to each bunker from control room, hearing loop required to speakers
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-058 | BENCH TOP: 600D, joinery under | 3
LINAC | Linear Accelerator Treatment Room | FFE | ITAV-010 | MICROPHONE: ceiling mounted | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-165 | CUPBOARD: tall, applicator storage, radiotherapy | 1 | custom joinery is required for vendor specific electron cones, including sliding rails, shelving to be reinforced for electron shields
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-166 | CUPBOARD: tall, shelving, horizontal patient device storage, radiotherapy | 3 | wide span shelving to be reinforced
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-167 | CUPBOARD: tall, shelving, vertical patient device storage, radiotherapy | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead is also acceptable
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-066 | COMPUTER: all-in-one, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | to wall mounted computer, may be wireless
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 6 | 1 x optional, to additional patient observation camera if provided
LINAC | Linear Accelerator Treatment Room | FFE | ITNE-073 | DISPLAY SCREEN: patient treatment information | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-006 | ACCESS PANEL: medical physics port | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-013 | DISPLAY SCREEN: radiotherapy, ceiling mounted | 2
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-015 | LASER: radiotherapy, ceiling mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-016 | LASER: radiotherapy, lateral laser | 2
LINAC | Linear Accelerator Treatment Room | FFE | FIJO-231 | CUPBOARD: radiotherapy, laser enclosure | 3
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-017 | LASER: radiotherapy, sagittal laser | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-008 | DEVICE: surface guided radiation therapy system | 3 | as required for respiratory gating / motion management / surface guidance systems, to suit service requirements
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-011 | DEVICE: radiotherapy, daily QA | 1
LINAC | Linear Accelerator Treatment Room | FFE | MMRA-012 | GATE: light gate, radiotherapy bunker | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-093 | BRACKET: keyboard shelf, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | FFE | FIHR-089 | BRACKET: surface guided radiation therapy system, ceiling mounted | 3
LINAC | Linear Accelerator Treatment Room | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 11 | 1 x optional, to additional patient observation camera if provided
LINAC | Linear Accelerator Treatment Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3
LINAC | Linear Accelerator Treatment Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
LINAC | Linear Accelerator Treatment Room | SE | ELPR-006 | RCD: residual current device | 5
LINAC | Linear Accelerator Treatment Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
LINAC | Linear Accelerator Treatment Room | SE | ELBO-001 | BUTTON: emergency stop | 4
LINAC | Linear Accelerator Treatment Room | SE | LIFX-004 | LIGHT: illuminated sign, "Beam On-Beam Off" | 1
LINAC | Linear Accelerator Treatment Room | SE | ELBO-010 | SWITCH: light, 2-way | 2
LINAC | Linear Accelerator Treatment Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
LINAC | Linear Accelerator Treatment Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1
LINAC | Linear Accelerator Treatment Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
LINAC | Linear Accelerator Treatment Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 4
LINAC | Linear Accelerator Treatment Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 2
LINAC | Linear Accelerator Treatment Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
LINAC | Linear Accelerator Treatment Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
LINAC | Linear Accelerator Treatment Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
LINAC | Linear Accelerator Treatment Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3
LINAC | Linear Accelerator Treatment Room | SE | ELGP-125 | GPO: emergency power, single, ceiling mounted | 12 | 1 x to ceiling mounted laser, 1 x optional, to additional patient observation camera if provided
LINAC | Linear Accelerator Treatment Room | SE | ELGP-220.1 | GPO: emergency power, double (V) | 5 | to wall mounted lasers
LINAC | Linear Accelerator Treatment Room | SE | ELBO-019 | BUTTON: last person out | 1
LINAC | Linear Accelerator Treatment Room | SE | LIFX-038 | LIGHT: backlit feature panels | 2