## RDS
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
MAMMO | Mammography | 16 | Available 24 Hours | 1 patient; Up to 2 staff assisting |  | Acoustic privacy is required | X-ray room for mammography imaging procedures, with change area within the room. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Cable ducts between X-ray table and control area may be required. 'X-ray in Use' lights required outside each entry door.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
MAMMO | Mammography | FA | CLCN-002 | CORNICE: aluminium, 24mm shadow angle, powdercoat | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | DWSC-026 | DOOR: 1600mm c/o, 1 1/2 leaf, solid core, lead shielding, paint | 1 | min. 1400mm clear opening | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ELBO-015 | SWITCH: light, single | 3 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ELGE-001 | DIRECT CONNECTION: power | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ELGP-106 | GPO: single, special, for equipment | 1 | Mammography unit | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 2 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ELGP-208 | GPO: double, wall mounted | 3 | 2 to desk area | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ELGP-209 | GPO: double, wall mounted, RCD protected | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIBM-006 | MIRROR: fixed, glass | 1 | to change area | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIDI-014 | DISPENSER: disposable glove | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIHR-026 | HOOK: coat | 2 | to change area, 1 @1200H | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIJO-116 | CUPBOARD: tall, double door, adj shelves | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 1 | above desk, optional | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1 | to change area | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to basin | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FQBS-014 | CHAIR: client | 1 | to change area | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FQBS-018 | CHAIR: office, ergonomic | 1 | to desk | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FQDW-001 | DESK: adj height, 1100W x 600D x 750H nom. | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | to change area | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | refer to BHWS-B (similar) | No | No | Yes | No | Yes | Yes | Yes | No | No
MAMMO | Mammography | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin | No | No | Yes | No | Yes | No | No | No | No
MAMMO | Mammography | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ITCL-009 | BUTTON: nurse call, staff / nurse assist | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to desk, 1 outlet for telephone | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 | to desk | No | Yes | No | No | No | No | No | No | No
MAMMO | Mammography | SE | LIFX-008 | LIGHT: task light, built in | 1 | to desk area | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | MMGE-018 | CHAIR: mammography | 1 | Hausted' | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | MMGE-152 | VIEWER: X-ray, 2 screen | 1 | above desk | Yes | No | No | No | No | No | No | No | No
MAMMO | Mammography | FFE | MMMI-156 | SYSTEM: imaging, mammography | 1 | with control panel | Yes | Yes | No | No | No | No | No | No | No
MAMMO | Mammography | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No
MAMMO | Mammography | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin | No | No | No | No | No | No | No | No | No