## Room Template Data
MRI Control Room (MRI-CTRL) - Room Template Properties and Room Template Data
2025-05-15
No. of Standard Components = 1
Code | Template Name | Standard Area | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Briefing - Hours of Operation - Duration - Duration | Briefing - Hours of Operation - Duration - extended/limited hours | Briefing - Hours of Operation - Hours/days - Open time | Briefing - Hours of Operation - Hours/days - Close time | Briefing - Hours of Operation - Hours/days - Days | Briefing - Hours of Operation - Comment | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experiance system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statuatory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magentic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: sensitive | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
MRI-CTRL | MRI Control Room | 14 | MRICR | 2 | 15.05.2025 | Up to 24 hours | FALSE |  |  |  |  | up to 3 staff | The MRI Control Room provides the space and equipment for remote control of MRI equipment, and the amenities for review of images and reporting. Direct visibility of the patient through an observation window and direct access to the imaging room is required. | ~ The MRI Control Room is located within Zone III (the outer controlled area) as set out in the RANZCR MRI Safety Guidelines. ~ The entrance of the MRI Imaging Room must be visible from the MRI Control Room. ~ Magnetic and Radiofrequency shielding is to be as advised by specialist consultant. ~ Cable ducts for services between the imaging unit, the equipment room and the control room will be required and must be coordinated with the MME vendor. Additional building services connections and outlets are to be coordinated to suit equipment requirements. ~ Waveguide ports, also known as penetration panels, are specialised openings in the walls designed to allow the passage of various items like tubing and cables without compromising the magnetic and radiofrequency shielding. Waveguide ports are required from the MRI Control Room to the MRI Imaging Room. ~ The size, quantity and configuration of computer display screens is to be finalised based on the specific image display and processing task requirements (e.g. number of application windows to be viewed concurrently) and ergonomic considerations. Providing flexibility of monitor positioning is vital to ensure adjustments are possible to meet the ergonomic requirements of different users. ~ Provision of height adjustable workstations or fixed joinery is to be in line with service requirements and local WHS policies. ~ Hours of operation depend on the service in which this room is located. Ambulatory medical imaging services may only operate during business hours or extended business hours, some small hospital services may have an on-call, after-hours service. Large units are likely to provide a 24-hour service, especially where a satellite service is provided with an emergency department. ~ Requirement for fixed and/or mobile duress to be confirmed to suit service requirements. ~ An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. ~ Provision of an intercom for staff communication within the department is dependent on service requirements. This function may be provided as part of the telephone system. ~ The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements and confirmed based on risk assessment considering the impact of a power outage on patient care/safety. | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Items In Room Template
Standard Components - Items in Room Template
2025-05-15
No. of Standard Components = 1
Code | Template Name | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
MRI-CTRL | MRI Control Room | D+W | AHWCV-001 | WINDOW COVERING: blind, roller, single | 1 | 1 | Internal - Window 1
MRI-CTRL | MRI Control Room | D+W | WIFX-008.04 | WINDOW: fixed, internal, single glazed, sill at 1050H, radio frequency shielding | 1 | 1 | Internal - Window 1 | for observation of patient from control room
MRI-CTRL | MRI Control Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | 1
MRI-CTRL | MRI Control Room | FFE | FIDI-241 | DISPENSER: alcohol-based hand rub, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | FFE | FIHR-026 | HOOK: coat | 2 | 1 |  | optional
MRI-CTRL | MRI Control Room | FFE | FIHR-201 | BRACKET: central processing unit (CPU), under bench/desk | 2 | 1 |  | 1 to MRI control CPU, 1 to PACS CPU
MRI-CTRL | MRI Control Room | FFE | FIHR-221 | BRACKET: computer display screen, dual, articulated, bench/desk mounted | 1 | 1 |  | to display/processing workstation
MRI-CTRL | MRI Control Room | FFE | FIHR-341 | BRACKET: television/AV system display screen, single, articulated, wall mounted | 1 | 1 |  | optional - to CCTV display screen if provided
MRI-CTRL | MRI Control Room | FFE | FIHR-408 | BRACKET: patient monitor, articulated, wall mounted | 1 | 1 |  | to slave patient monitor
MRI-CTRL | MRI Control Room | FFE | FQBS-172 | CHAIR: office, ergonomic, high | 3 | 1
MRI-CTRL | MRI Control Room | FFE | FQSM-287 | CABINET: mobile, 1800H nominal, tambour unit | 1 | 1 |  | type and extent of storage to suit service requirements
MRI-CTRL | MRI Control Room | FFE | FQSN-341 | LOCKER: floor standing, 4 door | 2 | 1 |  | optional
MRI-CTRL | MRI Control Room | FFE | ITAV-102 | CONTROL PANEL: AV system, wall mounted | 1 | 1 |  | vendor supplied sound system
MRI-CTRL | MRI Control Room | FFE | ITAV-152 | SPEAKER: AV system, ceiling recessed | 1 | 1
MRI-CTRL | MRI Control Room | FFE | ITCL-381 | DISPLAY SCREEN: patient observation | 1 | 1 |  | optional - for physiological monitoring
MRI-CTRL | MRI Control Room | FFE | ITNE-071 | TELEPHONE: handset, desktop | 1 | 1
MRI-CTRL | MRI Control Room | FFE | ITNE-102 | COMPUTER: dual display screens, central processing unit (CPU), keyboard and mouse, desktop | 1 | 1
MRI-CTRL | MRI Control Room | FFE | ITNE-153 | COMPUTER: central processing unit (CPU), PACS | 1 | 1
MRI-CTRL | MRI Control Room | FFE | ITNE-172 | DISPLAY SCREEN: computer, PACS | 2 | 1 |  | display/processing workstation, on dual arm bracket
MRI-CTRL | MRI Control Room | FFE | ITNE-201 | INPUT DEVICE: keyboard and mouse | 1 | 1
MRI-CTRL | MRI Control Room | FFE | ITNE-311 | PRINTER: desktop | 1 | 1
MRI-CTRL | MRI Control Room | FFE | ITSE-111 | DISPLAY SCREEN: security, CCTV monitoring | 1 | 1 |  | optional
MRI-CTRL | MRI Control Room | FFE | JOBE-061 | BENCH: 750D, laminate | 1 | 1 |  | bench height to be confirmed; adjustable height desks may be considered depending on WHS policies and coordination with services and observation window
MRI-CTRL | MRI Control Room | FFE | JOBE-211 | BENCH SUPPORT: bracket | 2 | 1 |  | leg support is also acceptable
MRI-CTRL | MRI Control Room | FFE | JOBE-221 | BENCH SUPPORT: end panel | 1 | 1 |  | leg support is also acceptable
MRI-CTRL | MRI Control Room | FFE | JOBE-251 | CAP: cable access | 3 | 1
MRI-CTRL | MRI Control Room | FFE | MIGE-046 | CONTROL CONSOLE: contrast injector, touch screen | 1 | 1
MRI-CTRL | MRI Control Room | FFE | MIMR-011 | CONTROL CONSOLE: MRI, dual display screens, keyboard and mouse, control unit, microphone and speaker, desktop | 1 | 1
MRI-CTRL | MRI Control Room | FFE | MIMR-016 | COMPUTER: MRI, central processing unit (CPU) | 1 | 1
MRI-CTRL | MRI Control Room | FFE | MIMR-032 | BUTTON: quench, MRI | 1 | 1
MRI-CTRL | MRI Control Room | FFE | MMPM-051 | MONITOR: patient, slave | 1 | 1
MRI-CTRL | MRI Control Room | FFE | FQWS-051 | BIN: general waste, 20L | 1 | 1
MRI-CTRL | MRI Control Room | FIN | CLCN-041 | CORNICE: wall trim fixing, prefinished | 1 | 1 | Ceiling Cornice
MRI-CTRL | MRI Control Room | FIN | CLTI-021 | CEILING: drop-in tiles, prefinished | 1 | 1 | Ceiling Finish 1
MRI-CTRL | MRI Control Room | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Finish 1
MRI-CTRL | MRI Control Room | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
MRI-CTRL | MRI Control Room | FIN | MIMR-112 | ACCESS PANEL: penetration, waveguide port, MRI | 1 | 1 | Wall Finish 2
MRI-CTRL | MRI Control Room | FIN | WLFI-002 | WALL FINISH: paint, clinical areas | 1 | 1 | Wall Finish 1
MRI-CTRL | MRI Control Room | SER | ELGP-105 | GPO: single, wall mounted, cleaner | 1 | 1
MRI-CTRL | MRI Control Room | SER | ELGP-131 | GPO: single, emergency power, wall mounted | 1 | 1 |  | optional - to CCTV display screen if provided
MRI-CTRL | MRI Control Room | SER | ELGP-161 | GPO: single, UPS, wall mounted | 1 | 1 |  | to slave patient monitor
MRI-CTRL | MRI Control Room | SER | ELGP-161 | GPO: single, UPS, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | SER | ELGP-201 | GPO: double, wall mounted | 1 | 1 |  | to printer
MRI-CTRL | MRI Control Room | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 6 | 1
MRI-CTRL | MRI Control Room | SER | ELGP-261 | GPO: double, UPS, wall mounted | 2 | 1 |  | to control console, extent of UPS to be confirmed to suit service and equipment requirements
MRI-CTRL | MRI Control Room | SER | ELPR-031 | PANEL: UPS status indicator, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | SER | ELPR-071 | RCD: residual current device, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | SER | ELSW-021 | SWITCH: light, dimmer | 2 | 1 |  | to general room lighting, 1 to each door
MRI-CTRL | MRI Control Room | SER | ELSW-121 | BUTTON: emergency stop, wall mounted | 1 | 1 |  | main power emergency stop, refer to equipment manufacturer for safety and compatibility with equipment
MRI-CTRL | MRI Control Room | SER | ITCL-181 | BUTTON: nurse call, staff assist, with cancel, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | SER | ITCL-181 | BUTTON: nurse call, staff assist, with cancel, wall mounted | 1 | 1 |  | optional - for anaesthetic assist
MRI-CTRL | MRI Control Room | SER | ITCL-191 | BUTTON: nurse call, emergency, with cancel, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | SER | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 |  | to slave patient monitor
MRI-CTRL | MRI Control Room | SER | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 |  | optional - to CCTV display screen if provided
MRI-CTRL | MRI Control Room | SER | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | 1
MRI-CTRL | MRI Control Room | SER | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 3 | 1
MRI-CTRL | MRI Control Room | SER | ITSE-061 | BUTTON: security, duress, fixed, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | SER | ITSE-166 | INTERCOM: audio, slave, wall mounted | 1 | 1 |  | optional - inclusion dependent on service requirements and ICT policies, may be integral function of telephone, to control workstation area
MRI-CTRL | MRI Control Room | SER | MEGE-001 | CONTROL PANEL: HVAC | 1 | 1
MRI-CTRL | MRI Control Room | SER | MEGE-081 | ALARM: oxygen depletion | 1 | 1
MRI-CTRL | MRI Control Room | SER | MGFP-001 | ALARM: medical gases, panel, wall mounted | 1 | 1
MRI-CTRL | MRI Control Room | SER | HYGE-351 | CONTROL PANEL: water pump, cooling | 1 | 1
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## SC Set for Publish (hidden)
Code
Administration | BMFD-3
Administration | BMFD-7
Administration | MEET-15
Administration | MEET-20
Administration | MEET-30
Administration | MEET-55
Administration | MEET-9
Administration | OFF-1P-12
Administration | OFF-1P-9
Administration | OFF-2P
Administration | OFF-WS
Administration | STFS-10
Administration | STFS-20
Allied Health | ADLB
Allied Health | ADLD
Allied Health | ADLK
Allied Health | ADLL
Allied Health | AUD-CR
Allied Health | AUD-TR
Allied Health | GYAH-GP
Allied Health | GYAH-ID
Allied Health | HYDP
Allied Health | PLST
Allied Health | TRMT-HTS
Allied Health | TRMT-POD
Ambulatory Care | BHW
Ambulatory Care | BWC
Ambulatory Care | CONS
Ambulatory Care | CONS-UN
Ambulatory Care | INTV
Ambulatory Care | PBTR-MD
Ambulatory Care | PROC
Ambulatory Care | TRMT
Cardiac Care | 1BR-CCU
Cardiac Investigations | CLAB-EP
Cardiac Investigations | CLAB-I
Cardiac Investigations | CLCR-EP
Cardiac Investigations | CLCR-I
Cardiac Investigations | CLUP-7
Cardiac Investigations | ECHO-TOE
Cardiac Investigations | STRT
Clinical Support Set 1 - Medication Rooms and Stores | AHDR
Clinical Support Set 1 - Medication Rooms and Stores | BHWS-A
Clinical Support Set 1 - Medication Rooms and Stores | BHWS-B
Clinical Support Set 1 - Medication Rooms and Stores | CLN-10
Clinical Support Set 1 - Medication Rooms and Stores | CLN-MED-20
Clinical Support Set 1 - Medication Rooms and Stores | CLN-MED-S
Clinical Support Set 1 - Medication Rooms and Stores | MED-14
Clinical Support Set 1 - Medication Rooms and Stores | STEQ-14
Clinical Support Set 1 - Medication Rooms and Stores | STEQ-20
Clinical Support Set 1 - Medication Rooms and Stores | STGN
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | BMEQ
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | BRES
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | CORR
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | LNGE-10
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | LNGE-20
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | OFF-CLN
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | OFF-WI-5
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-10
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-14
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-20
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-10
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-12
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-14
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-S
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-10
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-12
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-14
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-8
Emergency | CONS-ENT-OP
Emergency | PBAT
Emergency | PBTR-A
Emergency | PBTR-AS
Emergency | PBTR-FT
Emergency | PBTR-NA
Emergency | PBTR-R
Emergency | REC-E
Emergency | TRIAGE-1
Emergency | TRIAGE-2
Inpatient Accommodation | 1BR-BA
Inpatient Accommodation | 1BR-IS-N1
Inpatient Accommodation | 1BR-IS-N2
Inpatient Accommodation | 1BR-SP-A1
Inpatient Accommodation | 1BR-SP-A2
Inpatient Accommodation | 1BR-ST-A1
Inpatient Accommodation | 1BR-ST-A2
Inpatient Accommodation | 1BR-ST-A3
Inpatient Accommodation | 1BR-ST-B1
Inpatient Accommodation | 1BR-ST-B2
Inpatient Accommodation | 1BR-ST-B3
Inpatient Accommodation | 1BR-ST-D
Inpatient Accommodation | 2BR-ST-A1
Inpatient Accommodation | 2BR-ST-A2
Inpatient Accommodation | 2BR-ST-B
Inpatient Accommodation | 4BR-ST
Inpatient Accommodation | ANRM
Inpatient Accommodation | ENS-ACC
Inpatient Accommodation | ENS-BA
Inpatient Accommodation | ENS-SH
Inpatient Accommodation | ENS-SP
Inpatient Accommodation | ENS-ST-A1
Inpatient Accommodation | ENS-ST-A2
Inpatient Accommodation | ENS-ST-A3
Inpatient Accommodation | ENS-ST-B
Inpatient Accommodation | ENS-ST-C
Intensive Care | 1BR-IC
Intensive Care | PBIC
Logistics / Support Services | BBEV
Logistics / Support Services | BLIN
Logistics / Support Services | BMT
Logistics / Support Services | CLRM
Logistics / Support Services | COMM
Logistics / Support Services | DINR
Logistics / Support Services | DISP-10
Logistics / Support Services | DISP-15
Logistics / Support Services | PTRY
Logistics / Support Services | SECR
Logistics / Support Services | STBK
Logistics / Support Services | STCL
Maternity | BIRM-A
Maternity | BIRM-B
Maternity | ENS-BR
Medical Imaging | CTIR
Medical Imaging | CTIR-CTRL
Medical Imaging | FLUO
Medical Imaging | FLUO-CTRL
Medical Imaging | MAMMO
Medical Imaging | MRI
Medical Imaging | MRI-CTRL
Medical Imaging | OPG
Medical Imaging | REPR
Medical Imaging | REPW
Medical Imaging | ULTR
Medical Imaging | ULTR-PR
Medical Imaging | XRAY
Medical Imaging | XRAY-OPG
Mental Health | 1BR-MH-A
Mental Health | 1BR-MH-C
Mental Health | DINBEV-25
Mental Health | ENS-MH-A
Mental Health | ENS-MH-B
Mental Health | INTF-MH
Mental Health | LAUN-MH
Mental Health | SECL
Mental Health | STPP
Mental Health | WAIT-SEC
Mortuary and Autopsy | MOR-AU
Mortuary and Autopsy | MOR-BH
Mortuary and Autopsy | MOR-BR
Mortuary and Autopsy | MOR-VR
Mortuary and Autopsy | MOR-W
Neonatal Care | LAUN-PT
Neonatal Care | NBIC-HD
Neonatal Care | NBLD
Nuclear Medicine / PET | PET-CT
Nuclear Medicine / PET | SPECT-CT
Nuclear Medicine / PET | SPECT-CTCR
Nuclear Medicine / PET | UTRM
Operating | ANAES
Operating | B-SCRUB
Operating | CLNUP
Operating | ENPR
Operating | EXIT
Operating | OR-GEN
Operating | STSS
Oral Health | DEN-MLB
Oral Health | DENSR-1
Oral Health | DENSR-2
Pathology | AHBF
Pathology | BBLD
Pathology | BES
Pathology | BPATH
Pathology | BPTS
Pathology | PBSC
Perioperative | BBW
Perioperative | CHPT
Perioperative | CHPT-AC
Perioperative | PT-HOLD-A
Perioperative | PT-HOLD-B
Perioperative | PT-RS1
Pharmacy | PHA-CO
Pharmacy | PHA-DB
Pharmacy | PHA-DS
Pharmacy | PHA-PR
Pharmacy | PHA-RE
Pharmacy | STAD
Public Areas | AIRLE-12
Public Areas | AIRLE-6
Public Areas | BVM
Public Areas | BWTR
Public Areas | PAR
Public Areas | PLAY
Public Areas | RECP-10
Public Areas | RECP-15
Public Areas | RECP-20
Public Areas | WAIT-10
Public Areas | WAIT-20
Public Areas | WAIT-30
Public Areas | WAIT-S
Radiation Oncology | CTPR
Radiation Oncology | LINAC
Radiation Oncology | LINAC-CR
Renal Dialysis | PBTR-RD-A
Renal Dialysis | PBTR-RD-B
Renal Dialysis | PLNT-WT
Sanitary Facilities | BATH
Sanitary Facilities | SHAC
Sanitary Facilities | SHPT
Sanitary Facilities | SHST
Sanitary Facilities | WCAC
Sanitary Facilities | WCPT
Sanitary Facilities | WCPU
Sanitary Facilities | WCST
Staff Amenities | BPROP
Staff Amenities | CHST-10
Staff Amenities | CHST-35
Staff Amenities | OVBR
Staff Amenities | OVES
Staff Amenities | SRM-15
Staff Amenities | SRM-35