## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
MRIR | MRI Room | 46 | Available 24 hours | 1 patient; Up to 3 staff |  | Acoustic privacy is required | Imaging Room for Magnetic Resonance Imaging (MRI) scanning procedures with an adjacent MRI Equipment Room and Control Room with a direct view of the patient. The Control Room and Equipment room will need to be located in close proximity due to acceptable cable lengths. General anaesthesia may be administered in this room. Additional Design Considerations: The MRI suite design to conform to 4 safety exclusion zones, access control is required. Removable external wall panel to outside wall preferred where possible in order to replace the magnet. Magnetic and radiofrequency shielding are required, specialist consultant advice should be sought. Cable ducts between Control Room and Equipment Room will be required. Imaging equipment to be installed to manufacturer's specifications. Each magnet will have a chiller provided by supplier requiring UPS power, location will be project specific. Structural assessment is required to ensure the floor and access routes to the room will accommodate the weight of the MRI unit. All equipment and fittings in the room to be MRI compatible and constructed of non-ferrous material. Fluorescent luminaires are not permitted in the MRI Room. Consideration should be given to room acoustics - materials that reduce the impact of magnet noise should be included where possible. Temperature and humidity control required.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
MRIR | MRI Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
MRIR | MRI Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
MRIR | MRI Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
MRIR | MRI Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
MRIR | MRI Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
MRIR | MRI Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | tiled ceiling also acceptable; ceiling treatment to enhance patient experience to be considered
MRIR | MRI Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
MRIR | MRI Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 9
MRIR | MRI Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | MRI compatible
MRIR | MRI Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | MRI compatible
MRIR | MRI Room | FFE | FIJO-116 | CUPBOARD: tall, double door, adj shelves | 3 | for coils
MRIR | MRI Room | FFE | FQBS-037 | FOOTSTOOL | 1 | MRI compatible
MRIR | MRI Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | MRI compatible
MRIR | MRI Room | FFE | MMGE-002 | ADAPTER: suction | 2 | MRI compatible
MRIR | MRI Room | FFE | MMGE-100 | SUCTION BOTTLE | 2 | MRI compatible
MRIR | MRI Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
MRIR | MRI Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1
MRIR | MRI Room | FFE | MMMI-175 | CABIN: radio frequency, faraday cage, MRI | 1 | 3 dimensions
MRIR | MRI Room | FFE | MMMI-180 | DEVICE: imaging, magnet, MRI | 1 | refer to manufacturer's requirements for quench pipes and other associated equipment
MRIR | MRI Room | FFE | MMMI-181 | DOOR: radio frequency, MRI | 1 | 1200mm clear opening; with door protection to 900mm AFFL
MRIR | MRI Room | FFE | MMMI-185 | PIPE: quench, MRI | 1 | installed to manufacturer's specifications with minimal bends
MRIR | MRI Room | FFE | MMMI-183 | INJECTOR: contrast, MRI | 1
MRIR | MRI Room | FFE | MMMI-188 | CAMERA: observation, MRI | 2 | with direct view of patient
MRIR | MRI Room | FFE | MMMI-191 | TABLE: patient, fixed, MRI | 1 | detachable table also acceptable; use of detachable table may also negate need for MRI compatible patient transport trolley
MRIR | MRI Room | FFE | MMMI-195 | WINDOW: radiofrequency, MRI | 1 | 1 to control, 1 to anaesthetic preparation
MRIR | MRI Room | FFE | MMMI-198 | MACHINE: anaesthetic, MRI compatible | 1
MRIR | MRI Room | FFE | MMMI-199 | PANEL: waveguide, MRI | 4
MRIR | MRI Room | FFE | MMMI-200 | TROLLEY: patient, transport, MRI compatible | 1
MRIR | MRI Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | to cupboards
MRIR | MRI Room | FFE | MMMI-368 | LASER: positioning lights, MRI | 1 | optional, for radiotherapy planning
MRIR | MRI Room | FFE | FIHR-082 | BRACKET: for observation camera, MRI compatible, wall mounted | 2
MRIR | MRI Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
MRIR | MRI Room | SE | ELPR-006 | RCD: residual current device | 2 | 1 for each UPS & emergency power circuits
MRIR | MRI Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | 1-laser positioning light if required
MRIR | MRI Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
MRIR | MRI Room | SE | ELBO-001 | BUTTON: emergency stop | 1
MRIR | MRI Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1
MRIR | MRI Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2
MRIR | MRI Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1
MRIR | MRI Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
MRIR | MRI Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2
MRIR | MRI Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
MRIR | MRI Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | 1-clock
MRIR | MRI Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 2
MRIR | MRI Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
MRIR | MRI Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
MRIR | MRI Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | MRI compatible
MRIR | MRI Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
MRIR | MRI Room | SE | MMMI-201 | PANEL: penetration, MRI | 1
MRIR | MRI Room | SE | ELGP-133 | GPO: UPS, single, on services panel | 2
MRIR | MRI Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4
MRIR | MRI Room | SE | ELPR-007 | UPS: status indicator | 1
MRIR | MRI Room | SE | MEGE-016 | SENSOR: oxygen depletion | 1