## Room Template Data
MRI Room (MRI) - Room Template Properties and Room Template Data
2025-05-15
No. of Standard Components = 1
Code | Template Name | Standard Area | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Briefing - Hours of Operation - Duration - Duration | Briefing - Hours of Operation - Duration - extended/limited hours | Briefing - Hours of Operation - Hours/days - Open time | Briefing - Hours of Operation - Hours/days - Close time | Briefing - Hours of Operation - Hours/days - Days | Briefing - Hours of Operation - Comment | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experiance system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statuatory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magentic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: sensitive | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
MRI | MRI Room | 46 | MRIR | 2 | 15.05.2025 | Up to 24 hours | FALSE |  |  |  |  | 1 patient; Up to 3 staff | The MRI Imaging Room is designed for undertaking Magnetic Resonance Imaging (MRI) scanning with direct overview from the adjoining MRI Control Room. The MRI Imaging Room must have the MRI Equipment Room immediately adjacent for accommodating supporting power and control equipment. | ~ The MRI Imaging Room is located within Zone IV as set out in the RANZCR MRI Safety Guidelines. ~ Magnetic and Radiofrequency shielding is to be as advised by specialist consultant. ~ The entrance of the MRI Imaging Room must be visible from the MRI Control Room. ~ Appropriate warning signs are required indicating the presence and permanency of a strong magnetic field and identifying items for which entry is prohibited. ~ There shall be a designated site outside the MRI Imaging Room, outside the 0.5m T line and away from the room entry door, for patients to be brought to if requiring emergency treatment. ~ Equipment is to be installed to manufacturer’s specifications with sufficient clearances for maintenance access. Installation and replacement access routes should also be considered. Additional building services connections and outlets are to be coordinated to suit equipment requirements. Cable ducts for services between imaging unit, the equipment room and the control room will be required and must be coordinated with the MME vendor. ~ Structural assessment is required to ensure the floor of the MRI Imaging Room and the access route for installation/replacement of equipment will accommodate the weight of the MRI unit. ~ A removable panel to the wall of the MRI Imaging Room is required to allow MRI Unit replacement. ~ Depending on service requirements, general anaesthesia may be administered in this room. This must be confirmed by project teams to ensure appropriate workflows are provided to support this model of care, to confirm the correct building services are provided to ensure patient safety, and to determine medical gas system requirements. ~ Provision of reticulated nitrous oxide is optional and must be conﬁrmed to suit local jurisdictional policies and operational service requirements. The use of nitrous oxide in operating theatres and procedural suites and emergency departments is declining due to a range of clinical and environmental concerns. Reticulated systems have been found to increase leakage of nitrous oxide (a potent greenhouse gas) to atmosphere, can increase facility operating costs and potentially expose staff to nitrous oxide. Reticulated nitrous oxide and associated scavenge outlets are not mandatory for any healthcare service and point of care cylinders can meet clinical requirements for the majority of healthcare facilities. Due consideration must be given to a range of operational considerations including monitoring and measurement of usage, management of leakage, WHS requirements relating to the use of cylinders, approach to the provision of scavenge where cylinders are used, appropriate storage for cylinders, and security of gas sources given it is used as a recreational drug. ~ Waveguide ports, also known as penetration panels, are specialised openings in the walls designed to allow the passage of various items like tubing and cables without compromising the magnetic and radiofrequency shielding. Waveguide ports are required from the MRI Control Room to the MRI Imaging Room. Where a service includes administration of anaesthetic, waveguide ports are required from the Anaesthetic Preparation Area to the MRI Imaging Room, along with an observation window with clear view of the patient. ~ All equipment and fittings in the room to be MRI compatible and constructed of non-ferrous material ~ Consideration for the use of multi-parameter MRI compatible monitors and infusion pumps to be confirmed at project level to suit local jurisdictional policies and operational service requirements. Where not provided, a cable port may need to be installed to run clinical leads from outside the imaging room to the patient for continuous monitoring and provision of care. ~ Florescent light fixtures are not permitted in the MRI Imaging Room ~ Temperature and Humidity control are required. ~ In planning, room location to be considered in relation to location and movement of metal object outside the MRI Room, such as beds, trolleys, lifts, motor vehicles, trains, trams, AGV’s, transformers and AC cables. Room location must also take into consideration static elements such as steel beams and reinforcement, especially beneath the MRI Unit. ~ Room location to consider building vibration limits that can impact image quality. This may include new building works and potential refurbishments nearby. ~ Smoke detection within the MRI room shall be multi-aspirated detection with the sampling unit located outside of the risk area. Fire sprinklers and pipework shall be of non-ferrous material and suitable for an MRI room. A performance solution shall be completed to address the minimum fire protection requirements of the MRI room. ~ The room will require sufficient clearance to allow access for a patient on an MRI compatible trolley. Trolley movement into, around and out of the room, and patient transfer should be considered during planning to reduce manual handling. ~ Patient transfer aids such as patient slides and hover mats are to be considered for assisted patient transfers to and from the MRI table. Final selection of transfer aids will be dependent on service requirements, the needs of the anticipated patient cohort (e.g. for ambulatory services, acute services, etc.), as well as WHS risk assessments/policies. All equipment to be used within the room must be non-ferrous/MRI compatible. ~ Acoustic treatment to be considered to manage noise from the imaging device and the adjacent equipment room. All finishes must be meet infection prevention and control requirements. ~ Dimmable light may be considered to support patient comfort. Similarly, feature artwork on the ceiling may be considered above the MRI table. ~ Time tracking requirements (i.e. provision of electric or synchronised clock) is to be confirmed to suit service requirements. ~ The type, extent and configuration of storage within the room for consumables, linen, equipment, etc. will be dependent on service requirements, operational policies for stock centralisation/decanting and infection prevention and control policies. ~ Hours of operation depend on the service in which this room is located. Ambulatory medical imaging services may only operate during business hours or extended business hours, some small hospital services may have an on-call, after-hours service. Large units are likely to provide a 24-hour service, especially where a satellite service is provided with an emergency department. ~ Size, type and quantity of waste bins is dependent on waste management policies (e.g. for waste separation, frequency of waste removal, etc.) and service requirements. ~ Requirement for fixed and/or mobile duress to be confirmed to suit service requirements. ~ An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. ~ Provision of an intercom for staff communication within the department is dependent on service requirements. This function may be provided as part of the telephone system. ~ The final services and configuration of the Medical Services Panel (MSP) - medical gases, nurse call, power and data - will be dependent on service requirements. ~ The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements and confirmed based on risk assessment considering the impact of a power outage on patient care/safety. | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Items In Room Template
Standard Components - Items in Room Template
2025-05-15
No. of Standard Components = 1
Code | Template Name | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
MRI | MRI Room | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | MRI compatible, to both sides of door
MRI | MRI Room | D+W | DOSP-141 | DOOR: hinged, 1 leaf, 1200 clear opening, radio frequency shielding, MRI imaging room | 1 | 1 | Internal - Door 1
MRI | MRI Room | D+W | WIFX-008.04 | WINDOW: fixed, internal, single glazed, sill at 1050H, radio frequency shielding | 1 | 1 | Internal - Window 1 | for observation from control room (see MRI-CTRL)
MRI | MRI Room | D+W | WIFX-008.04 | WINDOW: fixed, internal, single glazed, sill at 1050H, radio frequency shielding | 1 | 1 | Internal - Window 2 | for observation from anaesthetic preparation area
MRI | MRI Room | D+W | MIMR-121 | BARRIER: MRI, retractable belt, wall mounted | 1 | 1 | Internal - Door 1
MRI | MRI Room | FFE | FIHR-101 | BRACKET: suction bottle, wall mounted | 2 | 1 |  | MRI compatible
MRI | MRI Room | FFE | FIHR-341 | BRACKET: television/AV system display screen, single, articulated, wall mounted | 1 | 1 |  | optional - to functional MRI display screen if provided, MRI compatible
MRI | MRI Room | FFE | FIHR-408 | BRACKET: patient monitor, articulated, wall mounted | 1 | 1 |  | optional - to slave patient monitor if provided, MRI compatible
MRI | MRI Room | FFE | FQGE-161 | STOOL: step | 1 | 1 |  | MRI compatible
MRI | MRI Room | FFE | ITAV-151 | SPEAKER: AV system, ceiling mounted | 1 | 1 |  | MRI compatible
MRI | MRI Room | FFE | ITCL-051 | CLOCK: electric, analogue | 1 | 1 |  | MRI compatible
MRI | MRI Room | FFE | ITCL-351 | DISPLAY SCREEN: patient treatment information | 1 | 1 |  | optional - for functional MRI, MRI compatible
MRI | MRI Room | FFE | ITCL-437 | CAMERA: patient observation, MRI compatible, wall mounted | 2 | 1
MRI | MRI Room | FFE | JOCU-021 | CUPBOARD: tall, double door | 5 | 1 |  | all materials and hardware to be MRI compatible
MRI | MRI Room | FFE | JOGE-001 | BULKHEAD: joinery | 2 | 1
MRI | MRI Room | FFE | MIGE-052 | INJECTOR: contrast, MRI compatible, pedestal | 1 | 1
MRI | MRI Room | FFE | MIMR-001 | IMAGING DEVICE: MRI, gantry | 1 | 1 |  | refer to manufacturer's requirements for quench pipes and other associated equipment
MRI | MRI Room | FFE | MIMR-006 | TABLE: patient, MRI | 1 | 1 |  | mobile table with electronic docking capability, detachable table also acceptable, use of detachable table may negate need for additional MRI compatible patient transport trolley
MRI | MRI Room | FFE | MIMR-036 | PIPE: quench, MRI | 1 | 1 |  | installed to manufacturer's specifications with minimal bends
MRI | MRI Room | FFE | MIMR-041 | TROLLEY: patient, transport, MRI compatible | 1 | 1
MRI | MRI Room | FFE | MIMR-046 | MACHINE: anaesthetic, MRI compatible | 1 | 1
MRI | MRI Room | FFE | MMGE-191 | CANNISTER: suction bottle | 2 | 1 |  | MRI compatible
MRI | MRI Room | FFE | MMPM-051 | MONITOR: patient, slave | 1 | 1 |  | optional - depending on service requirements, MRI compatible
MRI | MRI Room | FFE | MMRA-042 | LASER: patient positioning, radiotherapy, overhead, MRI compatible | 1 | 1 |  | optional - as required for radiotherapy planning
MRI | MRI Room | FIN | CLCN-031 | CORNICE: square set | 1 | 1 | Ceiling Cornice
MRI | MRI Room | FIN | CLFI-002 | CEILING FINISH: paint, clinical areas | 1 | 1 | Ceiling Finish 2
MRI | MRI Room | FIN | CLFS-012 | CEILING: flush set, suspended, acoustic | 1 | 1 | Ceiling Finish 1 | ceiling tiles are also acceptable, grid must be MRI compatable
MRI | MRI Room | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Skirting
MRI | MRI Room | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Finish 1
MRI | MRI Room | FIN | MIMR-101 | CABIN: radio frequency shielding, faraday cage, MRI | 1 | 1 | Wall Finish 4
MRI | MRI Room | FIN | MIMR-112 | ACCESS PANEL: penetration, waveguide port, MRI | 1 | 1 | Wall Finish 3 | 2 to control room, 2 to anaesthetic preparation area
MRI | MRI Room | FIN | WLFI-002 | WALL FINISH: paint, clinical areas | 1 | 1 | Wall Finish 1
MRI | MRI Room | FIN | WLFI-011.06 | WALL FINISH: vinyl, to 2100 AFFL | 1 | 1 | Wall Finish 2
MRI | MRI Room | MSP | ELGP-146 | GPO: single, emergency power, on service panel | 4 | 1 | Panel 1
MRI | MRI Room | MSP | ELGP-176 | GPO: single, UPS, on service panel | 2 | 1 | Panel 1 | extent of UPS to be confirmed to suit service and equipment requirements
MRI | MRI Room | MSP | ELPR-033 | PANEL: UPS status indicator, on services panel | 1 | 1 | Panel 1
MRI | MRI Room | MSP | ELPR-073 | RCD: residual current device, on services panel | 2 | 1 | Panel 1 | 1 to emergency power circuit, 1 to UPS circuit
MRI | MRI Room | MSP | ITCL-183 | BUTTON: nurse call, staff assist, with cancel, on services panel | 1 | 1 | Panel 1 | MRI compatible
MRI | MRI Room | MSP | MGAS-022 | OUTLET: medical air, on services panel | 1 | 1 | Panel 1
MRI | MRI Room | MSP | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | 1 | Panel 1 | optional; provision of reticulated nitrous oxide to be confirmed to suit project and service requirements
MRI | MRI Room | MSP | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | 1 | Panel 1
MRI | MRI Room | MSP | MGAS-052 | OUTLET: scavenge, on services panel | 1 | 1 | Panel 1 | optional; provided if reticulated nitrous oxide provided - provision of reticulated nitrous oxide to be confirmed to suit project and service requirements
MRI | MRI Room | MSP | MGAS-062 | OUTLET: suction, on services panel | 2 | 1 | Panel 1
MRI | MRI Room | MSP | MGFP-021 | FLOWMETER: medical air | 1 | 1 | Panel 1 | MRI compatible
MRI | MRI Room | MSP | MGFP-041 | FLOWMETER: oxygen | 2 | 1 | Panel 1 | MRI compatible
MRI | MRI Room | MSP | MGFP-061 | ADAPTER: suction | 2 | 1 | Panel 1 | MRI compatible
MRI | MRI Room | MSP | MMSP-051 | MEDICAL SERVICES PANEL: wall mounted | 1 | 1 | Panel 1 | MRI compatible
MRI | MRI Room | SER | ELGP-101 | GPO: single, wall mounted | 1 | 1 |  | to clock
MRI | MRI Room | SER | ELGP-106 | GPO: single, ceiling mounted | 1 | 1 |  | optional - to laser positioning light if provided
MRI | MRI Room | SER | ELGP-131 | GPO: single, emergency power, wall mounted | 2 | 1 |  | to patient observation cameras
MRI | MRI Room | SER | ELGP-161 | GPO: single, UPS, wall mounted | 2 | 1 |  | optional - 1 to slave patient monitor, 1 to FMRI monitor
MRI | MRI Room | SER | ELPR-071 | RCD: residual current device, wall mounted | 1 | 1
MRI | MRI Room | SER | ELSW-021 | SWITCH: light, dimmer | 1 | 1 |  | to general room lighting
MRI | MRI Room | SER | ELSW-121 | BUTTON: emergency stop, wall mounted | 1 | 1 |  | main power emergency stop, refer to equipment manufacturer for safety and compatibility with equipment
MRI | MRI Room | SER | ITCL-191 | BUTTON: nurse call, emergency, with cancel, wall mounted | 1 | 1 |  | above medical services panel
MRI | MRI Room | SER | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 |  | to patient observation cameras
MRI | MRI Room | SER | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2 | 1 |  | optional - 1 to slave patient monitor, 1 to FMRI monitor
MRI | MRI Room | SER | ITSE-061 | BUTTON: security, duress, fixed, wall mounted | 1 | 1
MRI | MRI Room | SER | MEGE-061 | SENSOR: oxygen depletion | 1 | 1 |  | connection to BMS to be confirmed by project
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## SC Set for Publish (hidden)
Code
Administration | BMFD-3
Administration | BMFD-7
Administration | MEET-15
Administration | MEET-20
Administration | MEET-30
Administration | MEET-55
Administration | MEET-9
Administration | OFF-1P-12
Administration | OFF-1P-9
Administration | OFF-2P
Administration | OFF-WS
Administration | STFS-10
Administration | STFS-20
Allied Health | ADLB
Allied Health | ADLD
Allied Health | ADLK
Allied Health | ADLL
Allied Health | AUD-CR
Allied Health | AUD-TR
Allied Health | GYAH-GP
Allied Health | GYAH-ID
Allied Health | HYDP
Allied Health | PLST
Allied Health | TRMT-HTS
Allied Health | TRMT-POD
Ambulatory Care | BHW
Ambulatory Care | BWC
Ambulatory Care | CONS
Ambulatory Care | CONS-UN
Ambulatory Care | INTV
Ambulatory Care | PBTR-MD
Ambulatory Care | PROC
Ambulatory Care | TRMT
Cardiac Care | 1BR-CCU
Cardiac Investigations | CLAB-EP
Cardiac Investigations | CLAB-I
Cardiac Investigations | CLCR-EP
Cardiac Investigations | CLCR-I
Cardiac Investigations | CLUP-7
Cardiac Investigations | ECHO-TOE
Cardiac Investigations | STRT
Clinical Support Set 1 - Medication Rooms and Stores | AHDR
Clinical Support Set 1 - Medication Rooms and Stores | BHWS-A
Clinical Support Set 1 - Medication Rooms and Stores | BHWS-B
Clinical Support Set 1 - Medication Rooms and Stores | CLN-10
Clinical Support Set 1 - Medication Rooms and Stores | CLN-MED-20
Clinical Support Set 1 - Medication Rooms and Stores | CLN-MED-S
Clinical Support Set 1 - Medication Rooms and Stores | MED-14
Clinical Support Set 1 - Medication Rooms and Stores | STEQ-14
Clinical Support Set 1 - Medication Rooms and Stores | STEQ-20
Clinical Support Set 1 - Medication Rooms and Stores | STGN
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | BMEQ
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | BRES
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | CORR
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | LNGE-10
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | LNGE-20
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | OFF-CLN
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | OFF-WI-5
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-10
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-14
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-20
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-10
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-12
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-14
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-S
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-10
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-12
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-14
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-8
Emergency | CONS-ENT-OP
Emergency | PBAT
Emergency | PBTR-A
Emergency | PBTR-AS
Emergency | PBTR-FT
Emergency | PBTR-NA
Emergency | PBTR-R
Emergency | REC-E
Emergency | TRIAGE-1
Emergency | TRIAGE-2
Inpatient Accommodation | 1BR-BA
Inpatient Accommodation | 1BR-IS-N1
Inpatient Accommodation | 1BR-IS-N2
Inpatient Accommodation | 1BR-SP-A1
Inpatient Accommodation | 1BR-SP-A2
Inpatient Accommodation | 1BR-ST-A1
Inpatient Accommodation | 1BR-ST-A2
Inpatient Accommodation | 1BR-ST-A3
Inpatient Accommodation | 1BR-ST-B1
Inpatient Accommodation | 1BR-ST-B2
Inpatient Accommodation | 1BR-ST-B3
Inpatient Accommodation | 1BR-ST-D
Inpatient Accommodation | 2BR-ST-A1
Inpatient Accommodation | 2BR-ST-A2
Inpatient Accommodation | 2BR-ST-B
Inpatient Accommodation | 4BR-ST
Inpatient Accommodation | ANRM
Inpatient Accommodation | ENS-ACC
Inpatient Accommodation | ENS-BA
Inpatient Accommodation | ENS-SH
Inpatient Accommodation | ENS-SP
Inpatient Accommodation | ENS-ST-A1
Inpatient Accommodation | ENS-ST-A2
Inpatient Accommodation | ENS-ST-A3
Inpatient Accommodation | ENS-ST-B
Inpatient Accommodation | ENS-ST-C
Intensive Care | 1BR-IC
Intensive Care | PBIC
Logistics / Support Services | BBEV
Logistics / Support Services | BLIN
Logistics / Support Services | BMT
Logistics / Support Services | CLRM
Logistics / Support Services | COMM
Logistics / Support Services | DINR
Logistics / Support Services | DISP-10
Logistics / Support Services | DISP-15
Logistics / Support Services | PTRY
Logistics / Support Services | SECR
Logistics / Support Services | STBK
Logistics / Support Services | STCL
Maternity | BIRM-A
Maternity | BIRM-B
Maternity | ENS-BR
Medical Imaging | CTIR
Medical Imaging | CTIR-CTRL
Medical Imaging | FLUO
Medical Imaging | FLUO-CTRL
Medical Imaging | MAMMO
Medical Imaging | MRI
Medical Imaging | MRI-CTRL
Medical Imaging | OPG
Medical Imaging | REPR
Medical Imaging | REPW
Medical Imaging | ULTR
Medical Imaging | ULTR-PR
Medical Imaging | XRAY
Medical Imaging | XRAY-OPG
Mental Health | 1BR-MH-A
Mental Health | 1BR-MH-C
Mental Health | DINBEV-25
Mental Health | ENS-MH-A
Mental Health | ENS-MH-B
Mental Health | INTF-MH
Mental Health | LAUN-MH
Mental Health | SECL
Mental Health | STPP
Mental Health | WAIT-SEC
Mortuary and Autopsy | MOR-AU
Mortuary and Autopsy | MOR-BH
Mortuary and Autopsy | MOR-BR
Mortuary and Autopsy | MOR-VR
Mortuary and Autopsy | MOR-W
Neonatal Care | LAUN-PT
Neonatal Care | NBIC-HD
Neonatal Care | NBLD
Nuclear Medicine / PET | PET-CT
Nuclear Medicine / PET | SPECT-CT
Nuclear Medicine / PET | SPECT-CTCR
Nuclear Medicine / PET | UTRM
Operating | ANAES
Operating | B-SCRUB
Operating | CLNUP
Operating | ENPR
Operating | EXIT
Operating | OR-GEN
Operating | STSS
Oral Health | DEN-MLB
Oral Health | DENSR-1
Oral Health | DENSR-2
Pathology | AHBF
Pathology | BBLD
Pathology | BES
Pathology | BPATH
Pathology | BPTS
Pathology | PBSC
Perioperative | BBW
Perioperative | CHPT
Perioperative | CHPT-AC
Perioperative | PT-HOLD-A
Perioperative | PT-HOLD-B
Perioperative | PT-RS1
Pharmacy | PHA-CO
Pharmacy | PHA-DB
Pharmacy | PHA-DS
Pharmacy | PHA-PR
Pharmacy | PHA-RE
Pharmacy | STAD
Public Areas | AIRLE-12
Public Areas | AIRLE-6
Public Areas | BVM
Public Areas | BWTR
Public Areas | PAR
Public Areas | PLAY
Public Areas | RECP-10
Public Areas | RECP-15
Public Areas | RECP-20
Public Areas | WAIT-10
Public Areas | WAIT-20
Public Areas | WAIT-30
Public Areas | WAIT-S
Radiation Oncology | CTPR
Radiation Oncology | LINAC
Radiation Oncology | LINAC-CR
Renal Dialysis | PBTR-RD-A
Renal Dialysis | PBTR-RD-B
Renal Dialysis | PLNT-WT
Sanitary Facilities | BATH
Sanitary Facilities | SHAC
Sanitary Facilities | SHPT
Sanitary Facilities | SHST
Sanitary Facilities | WCAC
Sanitary Facilities | WCPT
Sanitary Facilities | WCPU
Sanitary Facilities | WCST
Staff Amenities | BPROP
Staff Amenities | CHST-10
Staff Amenities | CHST-35
Staff Amenities | OVBR
Staff Amenities | OVES
Staff Amenities | SRM-15
Staff Amenities | SRM-35