## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
OPG | OPG Room | 12 | Available 24 Hours | 1 patient; 1 staff |  | Acoustic privacy is required | OPG is a room for undertaking orthopantomogram procedures which involve panoramic x-ray of the jaw area in dental examinations. The OPG unit may optionally be located within a Dental/ Oral Health unit or Imaging Unit. The OPG Room is designed to accommodate OPG, Ceph and Cone Beam CT imaging as required. An additional computer will be required if cone beam CT is included. The OPG imaging unit will require independent communications cabling to the related workstation. Room design will need to be adjusted if patients on beds need to be imaged. Additional Design Requirements: Radiation shielding to walls as advised by Radiation Consultant. 'X-ray in Use' warning lights will be required.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
OPG | OPG Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
OPG | OPG Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
OPG | OPG Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
OPG | OPG Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 1
OPG | OPG Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 7
OPG | OPG Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin
OPG | OPG Room | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1
OPG | OPG Room | FA | CLCN-002 | CORNICE: aluminium, 24mm shadow angle, powdercoat | 1
OPG | OPG Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | safety glass
OPG | OPG Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
OPG | OPG Room | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
OPG | OPG Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
OPG | OPG Room | FFE | FIHR-026 | HOOK: coat | 2 | optional
OPG | OPG Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
OPG | OPG Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
OPG | OPG Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1
OPG | OPG Room | FFE | FIDI-030 | DISPENSER: soap | 1
OPG | OPG Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1
OPG | OPG Room | FFE | FIJO-142 | CUPBOARD: underbench, single, adj shelves, with single drawer | 1
OPG | OPG Room | FFE | FQSM-016 | TROLLEY | 1
OPG | OPG Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
OPG | OPG Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
OPG | OPG Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
OPG | OPG Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
OPG | OPG Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
OPG | OPG Room | FFE | MMMI-324 | DEVICE: imaging, dental imaging unit (OPG/CEPH/CONE BEAM CT) | 1 | may be mounted to pole or a wall plate
OPG | OPG Room | FFE | FIJO-069 | BENCH TOP: 750D, joinery under, laminate finish | 1
OPG | OPG Room | FFE | FIJO-099 | CAC: cable access cap | 2
OPG | OPG Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
OPG | OPG Room | FFE | MMMI-371 | BUTTON: dental imaging unit, remote activation | 1 | locate outside room
OPG | OPG Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 1
OPG | OPG Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
OPG | OPG Room | FFE | ITNE-069 | DISPLAY SCREEN: PACS, portrait | 1
OPG | OPG Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | for BUTTON:OPG, to connect to outlet adjacent to OPG
OPG | OPG Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | computer, PACS screen
OPG | OPG Room | SE | ELPR-006 | RCD: residual current device | 1
OPG | OPG Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
OPG | OPG Room | SE | ELBO-001 | BUTTON: emergency stop | 2 | main power emergency stop, consult equipment manufacturer for safety and compatibility with equipment.
OPG | OPG Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1 | outside room entry, auto on when modality is in use
OPG | OPG Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
OPG | OPG Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | clock
OPG | OPG Room | SE | ELGP-223.1 | GPO: emergency power, double, wall mounted (V) | 5
OPG | OPG Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
OPG | OPG Room | SE | WLSH-004 | SHIELDING: lead lining | 1
OPG | OPG Room | SE | ELBO-012 | SWITCH: light, dimmer | 2