## Room Template Data
OPG Room (OPG) - Room Template Properties and Room Template Data
2025-05-15
No. of Standard Components = 1
Code | Template Name | Standard Area | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Briefing - Hours of Operation - Duration - Duration | Briefing - Hours of Operation - Duration - extended/limited hours | Briefing - Hours of Operation - Hours/days - Open time | Briefing - Hours of Operation - Hours/days - Close time | Briefing - Hours of Operation - Hours/days - Days | Briefing - Hours of Operation - Comment | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experiance system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statuatory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magentic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: sensitive | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
OPG | OPG Room | 19 |  | 2 | 15.05.2025 | Up to 24 hours | FALSE |  |  |  |  | 1 patient; 1 staff | The OPG Room provides the space and equipment for diagnostic orthopantomogram x-ray imaging which produces panoramic x-rays of the jaw. This Standard Component may be located within a Dental/Oral Health Unit or Medical Imaging Unit. | ~ Equipment is to be installed to manufacturer’s specifications with sufficient clearances for maintenance access. Installation and replacement access routes should also be considered. Additional building services connections and outlets are to be coordinated to suit equipment requirements. Cable ducts for services between the imaging device and the control area will be required and must be coordinated with the MME vendor. ~ The OPG Room is designed to accommodate OPG, Lateral Cephalogram and Cone Beam CT imaging devices with minor adjustments to suit equipment and meet shielding requirements. ~ Radiation shielding requirements are to be assessed and advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative. Depending on the equipment selected, a door may be required between the imaging area and the control area to provide adequate shielding. ~ An ‘X-ray in Use’ light system is required with signs that are illuminated when the imaging device is on and turned off when it stops. Signs are to be at all doors entering the room. ~ Provision of height adjustable workstations or fixed joinery is to be in line with service requirements and local WHS policies. ~ The type, extent and configuration of storage within the room for consumables, linen, equipment, etc. will be dependent on service requirements, operational policies for stock centralisation/decanting and infection prevention and control policies. ~ Hours of operation depend on the service in which this room is located. Ambulatory medical imaging services may only operate during business hours or extended business hours, some small hospital services may have an on-call, after-hours service. Large units are likely to provide a 24-hour service, especially where a satellite service is provided with an emergency department. ~ Size, type and quantity of waste bins is dependent on waste management policies (e.g. for waste separation, frequency of waste removal, etc.) and service requirements. ~ Requirement for fixed and/or mobile duress to be confirmed to suit service requirements. ~ An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. ~ The final services and configuration of the Medical Services Panel (MSP) - medical gases, nurse call, power and data - will be dependent on service requirements. ~ The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements and confirmed based on risk assessment considering the impact of a power outage on patient care/safety. | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Items In Room Template
Standard Components - Items in Room Template
2025-05-15
No. of Standard Components = 1
Code | Template Name | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
OPG | OPG Room | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | to both sides of door
OPG | OPG Room | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 2 | to both sides of door
OPG | OPG Room | D+W | DOHI-003.02 | DOOR: hinged, 1 leaf, 900 clear opening, solid, shielded | 1 | 1 | Internal - Door 1 | to imaging area
OPG | OPG Room | D+W | DOHI-003.02 | DOOR: hinged, 1 leaf, 900 clear opening, solid, shielded | 1 | 1 | Internal - Door 2 | to control workstation area
OPG | OPG Room | D+W | WIFX-008.03 | WINDOW: fixed, internal, single glazed, sill at 1050H, shielded, lead glass | 1 | 1 | Internal - Window 1 | for observation from control room
OPG | OPG Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | 1
OPG | OPG Room | FFE | FIDI-106 | DISPENSER: PPE, disposable gloves, wall mounted | 1 | 1
OPG | OPG Room | FFE | FIDI-181 | DISPENSER: cleaning wipes, wall mounted | 1 | 1
OPG | OPG Room | FFE | FIDI-231 | DISPENSER: paper towel, wall mounted | 1 | 1 |  | to basin
OPG | OPG Room | FFE | FIDI-241 | DISPENSER: alcohol-based hand rub, wall mounted | 1 | 1
OPG | OPG Room | FFE | FIDI-256 | DISPENSER: soap, wall mounted | 1 | 1 |  | to basin
OPG | OPG Room | FFE | FIDI-261 | DISPENSER: soap, antiseptic, wall mounted | 1 | 1 |  | to basin
OPG | OPG Room | FFE | FIHR-026 | HOOK: coat | 2 | 1 |  | optional - for patient property, 1 mounted at 1200H if provided
OPG | OPG Room | FFE | FIHR-101 | BRACKET: suction bottle, wall mounted | 1 | 1
OPG | OPG Room | FFE | FIHR-126 | RACK: lead apron, multi, wall mounted | 1 | 1
OPG | OPG Room | FFE | FIHR-151 | PANEL: storage baskets, wall mounted | 1 | 1 |  | optional - for storage of consumables and PPE items to meet service requirements
OPG | OPG Room | FFE | FIHR-161 | BRACKET: storage basket, wall mounted | 1 | 1 |  | located close to medical services panel
OPG | OPG Room | FFE | FIHR-201 | BRACKET: central processing unit (CPU), under bench/desk | 1 | 1
OPG | OPG Room | FFE | FQBS-172 | CHAIR: office, ergonomic, high | 1 | 1
OPG | OPG Room | FFE | FQGE-112 | BASKET: storage, for wall mounted panel, small | 1 | 1
OPG | OPG Room | FFE | FQSN-204 | SHELVING UNIT: floor standing, adjustable shelves, 400D x 900W x 2100H nominal | 1 | 1 |  | type and extent of storage to suit service requirements, operational policies for stock centralisation/decanting and infection prevention and control policies
OPG | OPG Room | FFE | HYBA-121 | BASIN: type B, handwashing | 1 | 1
OPG | OPG Room | FFE | HYTP-067 | TAPWARE: basin, tap set, wall mounted, lever handles | 1 | 1
OPG | OPG Room | FFE | ITNE-071 | TELEPHONE: handset, desktop | 1 | 1
OPG | OPG Room | FFE | ITNE-101 | COMPUTER: single display screen, central processing unit (CPU), keyboard and mouse, desktop | 1 | 1
OPG | OPG Room | FFE | ITNE-153 | COMPUTER: central processing unit (CPU), PACS | 1 | 1
OPG | OPG Room | FFE | ITNE-172 | DISPLAY SCREEN: computer, PACS | 1 | 1
OPG | OPG Room | FFE | ITNE-201 | INPUT DEVICE: keyboard and mouse | 1 | 1
OPG | OPG Room | FFE | ITNE-341 | SCANNER: desktop, document, flatbed, desktop | 1 | 1
OPG | OPG Room | FFE | JOBE-061 | BENCH: 750D, laminate | 1 | 1
OPG | OPG Room | FFE | JOBE-251 | CAP: cable access | 2 | 1
OPG | OPG Room | FFE | MIOC-001 | IMAGING DEVICE: orthopantomogram (OPG) | 1 | 1
OPG | OPG Room | FFE | MIOC-011 | CONTROL PANEL: orthopantomogram (OPG), remote activation | 1 | 1
OPG | OPG Room | FFE | MMGE-191 | CANNISTER: suction bottle | 1 | 1
OPG | OPG Room | FFE | FQWS-051 | BIN: general waste, 20L | 1 | 1 |  | to basin
OPG | OPG Room | FFE | MMTR-001.02 | TROLLEY: stainless steel, 490W x 490D x 900H nom, 2 tier | 1 | 1
OPG | OPG Room | FIN | CLCN-041 | CORNICE: wall trim fixing, prefinished | 1 | 1 | Ceiling Cornice
OPG | OPG Room | FIN | CLTI-021 | CEILING: drop-in tiles, prefinished | 1 | 1 | Ceiling Finish 1
OPG | OPG Room | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Finish 1 | minimum slip rating R10 / pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
OPG | OPG Room | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
OPG | OPG Room | FIN | WLFI-002 | WALL FINISH: paint, clinical areas | 1 | 1 | Wall Finish 1
OPG | OPG Room | FIN | WLFI-101 | SPLASHBACK: vinyl | 1 | 1 | Wall Finish 2 | to basin
OPG | OPG Room | MSP | ELGP-146 | GPO: single, emergency power, on service panel | 1 | 1 | Panel 1
OPG | OPG Room | MSP | ELPR-073 | RCD: residual current device, on services panel | 2 | 1 | Panel 1
OPG | OPG Room | MSP | MGAS-022 | OUTLET: medical air, on services panel | 1 | 1 | Panel 1
OPG | OPG Room | MSP | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | 1 | Panel 1
OPG | OPG Room | MSP | MGAS-062 | OUTLET: suction, on services panel | 1 | 1 | Panel 1
OPG | OPG Room | MSP | MGFP-021 | FLOWMETER: medical air | 1 | 1 | Panel 1
OPG | OPG Room | MSP | MGFP-041 | FLOWMETER: oxygen | 1 | 1 | Panel 1
OPG | OPG Room | MSP | MGFP-061 | ADAPTER: suction | 1 | 1 | Panel 1
OPG | OPG Room | MSP | MMSP-051 | MEDICAL SERVICES PANEL: wall mounted | 1 | 1 | Panel 1
OPG | OPG Room | SER | ELGP-105 | GPO: single, wall mounted, cleaner | 1 | 1
OPG | OPG Room | SER | ELGP-201 | GPO: double, wall mounted | 1 | 1 |  | to workstation, over desk for portable devices
OPG | OPG Room | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 1 | 1 |  | to imaging equipment
OPG | OPG Room | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 4 | 1 |  | to control workstation
OPG | OPG Room | SER | ELPR-071 | RCD: residual current device, wall mounted | 1 | 1 |  | to imaging device
OPG | OPG Room | SER | ELSW-021 | SWITCH: light, dimmer | 1 | 1 |  | to general room lighting
OPG | OPG Room | SER | ELSW-043 | SWITCH: light, x-ray in use | 1 | 1
OPG | OPG Room | SER | ELSW-121 | BUTTON: emergency stop, wall mounted | 1 | 1 |  | main power emergency stop, refer to equipment manufacturer for safety and compatibility with equipment
OPG | OPG Room | SER | HYDR-201 | DIRECT CONNECTION: waste water | 1 | 1 |  | to basin
OPG | OPG Room | SER | HYTP-451 | DIRECT CONNECTION: water, cold | 1 | 1 |  | to basin
OPG | OPG Room | SER | HYTP-453 | DIRECT CONNECTION: water, warm | 1 | 1 |  | to basin
OPG | OPG Room | SER | ITCL-181 | BUTTON: staff assist, with cancel, wall mounted | 1 | 1
OPG | OPG Room | SER | ITCL-191 | BUTTON: emergency, with cancel, wall mounted | 1 | 1
OPG | OPG Room | SER | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 |  | as required for imaging device connection to remote activation panel
OPG | OPG Room | SER | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | 1 |  | to imaging device
OPG | OPG Room | SER | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | 1 |  | to control workstation
OPG | OPG Room | SER | ITSE-061 | BUTTON: security, duress, fixed, wall mounted | 2 | 1 |  | 1 close to imaging device, 1 to control workstation area
OPG | OPG Room | SER | LIFX-213 | LIGHT: illuminated sign, x-ray in use | 2 | 1 |  | outside room entries, to automatically turn on when modality is in use
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## SC Set for Publish (hidden)
Code
Administration | BMFD-3
Administration | BMFD-7
Administration | MEET-15
Administration | MEET-20
Administration | MEET-30
Administration | MEET-55
Administration | MEET-9
Administration | OFF-1P-12
Administration | OFF-1P-9
Administration | OFF-2P
Administration | OFF-WS
Administration | STFS-10
Administration | STFS-20
Allied Health | ADLB
Allied Health | ADLD
Allied Health | ADLK
Allied Health | ADLL
Allied Health | AUD-CR
Allied Health | AUD-TR
Allied Health | GYAH-GP
Allied Health | GYAH-ID
Allied Health | HYDP
Allied Health | PLST
Allied Health | TRMT-HTS
Allied Health | TRMT-POD
Ambulatory Care | BHW
Ambulatory Care | BWC
Ambulatory Care | CONS
Ambulatory Care | CONS-UN
Ambulatory Care | INTV
Ambulatory Care | PBTR-MD
Ambulatory Care | PROC
Ambulatory Care | TRMT
Cardiac Care | 1BR-CCU
Cardiac Investigations | CLAB-EP
Cardiac Investigations | CLAB-I
Cardiac Investigations | CLCR-EP
Cardiac Investigations | CLCR-I
Cardiac Investigations | CLUP-7
Cardiac Investigations | ECHO-TOE
Cardiac Investigations | STRT
Clinical Support Set 1 - Medication Rooms and Stores | AHDR
Clinical Support Set 1 - Medication Rooms and Stores | BHWS-A
Clinical Support Set 1 - Medication Rooms and Stores | BHWS-B
Clinical Support Set 1 - Medication Rooms and Stores | CLN-10
Clinical Support Set 1 - Medication Rooms and Stores | CLN-MED-20
Clinical Support Set 1 - Medication Rooms and Stores | CLN-MED-S
Clinical Support Set 1 - Medication Rooms and Stores | MED-14
Clinical Support Set 1 - Medication Rooms and Stores | STEQ-14
Clinical Support Set 1 - Medication Rooms and Stores | STEQ-20
Clinical Support Set 1 - Medication Rooms and Stores | STGN
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | BMEQ
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | BRES
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | CORR
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | LNGE-10
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | LNGE-20
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | OFF-CLN
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | OFF-WI-5
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-10
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-14
Clinical Support Set 2 - Corridor Bays, Staff Stations and Lounges | SSTN-20
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-10
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-12
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-14
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | DTUR-S
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-10
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-12
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-14
Clinical Support Set 3 - Dirty Utilities and Equipment Clean-up | ECL-8
Emergency | CONS-ENT-OP
Emergency | PBAT
Emergency | PBTR-A
Emergency | PBTR-AS
Emergency | PBTR-FT
Emergency | PBTR-NA
Emergency | PBTR-R
Emergency | REC-E
Emergency | TRIAGE-1
Emergency | TRIAGE-2
Inpatient Accommodation | 1BR-BA
Inpatient Accommodation | 1BR-IS-N1
Inpatient Accommodation | 1BR-IS-N2
Inpatient Accommodation | 1BR-SP-A1
Inpatient Accommodation | 1BR-SP-A2
Inpatient Accommodation | 1BR-ST-A1
Inpatient Accommodation | 1BR-ST-A2
Inpatient Accommodation | 1BR-ST-A3
Inpatient Accommodation | 1BR-ST-B1
Inpatient Accommodation | 1BR-ST-B2
Inpatient Accommodation | 1BR-ST-B3
Inpatient Accommodation | 1BR-ST-D
Inpatient Accommodation | 2BR-ST-A1
Inpatient Accommodation | 2BR-ST-A2
Inpatient Accommodation | 2BR-ST-B
Inpatient Accommodation | 4BR-ST
Inpatient Accommodation | ANRM
Inpatient Accommodation | ENS-ACC
Inpatient Accommodation | ENS-BA
Inpatient Accommodation | ENS-SH
Inpatient Accommodation | ENS-SP
Inpatient Accommodation | ENS-ST-A1
Inpatient Accommodation | ENS-ST-A2
Inpatient Accommodation | ENS-ST-A3
Inpatient Accommodation | ENS-ST-B
Inpatient Accommodation | ENS-ST-C
Intensive Care | 1BR-IC
Intensive Care | PBIC
Logistics / Support Services | BBEV
Logistics / Support Services | BLIN
Logistics / Support Services | BMT
Logistics / Support Services | CLRM
Logistics / Support Services | COMM
Logistics / Support Services | DINR
Logistics / Support Services | DISP-10
Logistics / Support Services | DISP-15
Logistics / Support Services | PTRY
Logistics / Support Services | SECR
Logistics / Support Services | STBK
Logistics / Support Services | STCL
Maternity | BIRM-A
Maternity | BIRM-B
Maternity | ENS-BR
Medical Imaging | CTCR
Medical Imaging | CTIR
Medical Imaging | FLUO
Medical Imaging | FLUOC
Medical Imaging | MAMMO
Medical Imaging | MRICR
Medical Imaging | MRIR
Medical Imaging | OPG
Medical Imaging | REPR
Medical Imaging | REPW
Medical Imaging | ULTR
Medical Imaging | ULTR-PR
Medical Imaging | XRAY
Medical Imaging | XRAY-OPG
Mental Health | 1BR-MH-A
Mental Health | 1BR-MH-C
Mental Health | DINBEV-25
Mental Health | ENS-MH-A
Mental Health | ENS-MH-B
Mental Health | INTF-MH
Mental Health | LAUN-MH
Mental Health | SECL
Mental Health | STPP
Mental Health | WAIT-SEC
Mortuary and Autopsy | MOR-AU
Mortuary and Autopsy | MOR-BH
Mortuary and Autopsy | MOR-BR
Mortuary and Autopsy | MOR-VR
Mortuary and Autopsy | MOR-W
Neonatal Care | LAUN-PT
Neonatal Care | NBIC-HD
Neonatal Care | NBLD
Nuclear Medicine / PET | PET-CT
Nuclear Medicine / PET | SPECT-CT
Nuclear Medicine / PET | SPECT-CTCR
Nuclear Medicine / PET | UTRM
Operating | ANAES
Operating | B-SCRUB
Operating | CLNUP
Operating | ENPR
Operating | EXIT
Operating | OR-GEN
Operating | STSS
Oral Health | DEN-MLB
Oral Health | DENSR-1
Oral Health | DENSR-2
Pathology | AHBF
Pathology | BBLD
Pathology | BES
Pathology | BPATH
Pathology | BPTS
Pathology | PBSC
Perioperative | BBW
Perioperative | CHPT
Perioperative | CHPT-AC
Perioperative | PT-HOLD-A
Perioperative | PT-HOLD-B
Perioperative | PT-RS1
Pharmacy | PHA-CO
Pharmacy | PHA-DB
Pharmacy | PHA-DS
Pharmacy | PHA-PR
Pharmacy | PHA-RE
Pharmacy | STAD
Public Areas | AIRLE-12
Public Areas | AIRLE-6
Public Areas | BVM
Public Areas | BWTR
Public Areas | PAR
Public Areas | PLAY
Public Areas | RECP-10
Public Areas | RECP-15
Public Areas | RECP-20
Public Areas | WAIT-10
Public Areas | WAIT-20
Public Areas | WAIT-30
Public Areas | WAIT-S
Radiation Oncology | CTPR
Radiation Oncology | LINAC
Radiation Oncology | LINAC-CR
Renal Dialysis | PBTR-RD-A
Renal Dialysis | PBTR-RD-B
Renal Dialysis | PLNT-WT
Sanitary Facilities | BATH
Sanitary Facilities | SHAC
Sanitary Facilities | SHPT
Sanitary Facilities | SHST
Sanitary Facilities | WCAC
Sanitary Facilities | WCPT
Sanitary Facilities | WCPU
Sanitary Facilities | WCST
Staff Amenities | BPROP
Staff Amenities | CHST-10
Staff Amenities | CHST-35
Staff Amenities | OVBR
Staff Amenities | OVES
Staff Amenities | SRM-15
Staff Amenities | SRM-35