## Room Template Data
Operating Room - General (OR-GEN) - Room Template Properties and Room Template Data
2026-08-06
No. of Standard Components = 1
Code | Template Name | Standard Area | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Briefing - Hours of Operation - Duration - Duration | Briefing - Hours of Operation - Duration - extended/limited hours | Briefing - Hours of Operation - Hours/days - Open time | Briefing - Hours of Operation - Hours/days - Close time | Briefing - Hours of Operation - Hours/days - Days | Briefing - Hours of Operation - Comment | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experiance system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statuatory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magentic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: sensitive | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
OR-GEN | Operating Room - General | 60 | ORGN | 7 | 06.05.2025 | Up to 24 hours | 0 |  |  |  |  | 1 patient; 5-12 staff | An Operating Room provides an aseptic environment in which to carry out surgical procedures on a patient under local, regional or general anaesthetic. The Operating Room - General is flexibly used for various surgical procedures that require complex and large equipment. This Standard Component may be adapted for Operating Rooms dedicated to different specialties, for example neurosurgery, orthopaedics, cardiac surgery, obstetrics and trauma. | ~ The minimum critical dimension centred on the sterile zone is 7400mm clear to ensure adequate space within the sterile zone (min 3400mm wide), circulation zone (min 1000mm wide) and equipment zone (min 1000mm wide). ~ Where the Operating Room is intended to be used for cardiothoracic procedures and for some forms of hybrid operating theatres, cardiac electrical protection is required. ~ Isolated power supply is required individually to each medical services pendant, together with its associated line isolation and overload monitors. ~ Location and quantity of RCD outlets to suit overall electrical service provision and is to be confirmed by the project to suit local engineering services guidelines/policies. ~ A wall mounted medical services panel has been indicated as optional and may be provided for theatres where, for example, obstetric procedures are intended to be undertaken and the additional panel provides access to services for an infant away from the operating table if required. ~ Provision of reticulated nitrous oxide is optional and must be conﬁrmed to suit local jurisdictional policies and operational service requirements. The use of nitrous oxide in operating theatres and procedural suites and emergency departments is declining due to a range of clinical and environmental concerns. Reticulated systems have been found to increase leakage of nitrous oxide (a potent greenhouse gas) to atmosphere, can increase facility operating costs and potentially expose staff to nitrous oxide. Reticulated nitrous oxide and associated scavenge outlets are not mandatory for any healthcare service and point of care cylinders can meet clinical requirements for the majority of healthcare facilities. Due consideration must be given to a range of operational considerations including monitoring and measurement of usage, management of leakage, WHS requirements relating to the use of cylinders, approach to the provision of scavenge where cylinders are used, appropriate storage for cylinders, and security of gas sources given it is used as a recreational drug. ~ If laser equipment is to be used, ‘laser in use’ warning lights are required outside all entry doors to the room. ~ If x-ray equipment is to be used, ‘x-ray in use’ warning lights are required outside all entry doors to the room. Radiation shielding will also be required and is to be assessed by a Radiation Consultant. ~ Automated doors with motion or elbow activation are provided to the Operating Room from the Anaesthetic Preparation Room and Exit Bay. Buttons/sensors should be appropriately located in relation to the door, including consideration for approach and movement of patients on beds and mobile equipment. 2-gang buttons that support the opening of either one or both leaves of the door as required are desirable. ~ Height of doors to be confirmed to suit the movement of mobile equipment such as mobile imaging devices. ~ If external windows are provided, they are to be double glazed with integral venetian blinds (with either manual or automated controls). ~ Room lights to be dimmable. Lighting control to room rights, warning signs, and operating lights to be coordinated within the room and an appropriate lighting control panel to be provided to functional requirements. ~ Configuration of services pendant (horizontal or vertical) to suit services requirements. Accessories (e.g. shelves, drawers, rails, etc.) are optional and can be selected to suit service and equipment requirements. Services allocations are indicative - type and number of services on the pendant to suit service requirements ~ Services pendant arms to have an adequate reach to be moved out of the way to allow for adequate room for movement of a patient trolley and mobile equipment. ~ If not dedicated for use in a particular Operating Room, microscope(s) should be located in a nearby equipment bay in close proximity for flexible use between Operating Rooms. | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE
This shape represents a table slicer. Table slicers are not supported in this version of Excel.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Items In Room Template
Standard Components - Items in Room Template
2026-08-06
No. of Standard Components = 1
Code | Template Name | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
OR-GEN | Operating Room - General | D+W | AHDEC-031 | SWITCH: door release, single gang, automatic door, wall mounted | 1 | 1 | Internal - Door 3
OR-GEN | Operating Room - General | D+W | AHDEC-036 | SWITCH: door release, double gang, automatic door, wall mounted | 1 | 1 | Internal - Door 1 | to allow for opening either one or both leaves of the door
OR-GEN | Operating Room - General | D+W | AHDEC-036 | SWITCH: door release, double gang, automatic door, wall mounted | 1 | 1 | Internal - Door 2 | to allow for opening either one or both leaves of the door
OR-GEN | Operating Room - General | D+W | AHDMO-001 | DOOR MOTORISATION: swing | 1 | 1 | Internal - Door 1 | delayed action closing
OR-GEN | Operating Room - General | D+W | AHDMO-001 | DOOR MOTORISATION: swing | 1 | 1 | Internal - Door 2 | delayed action closing
OR-GEN | Operating Room - General | D+W | AHDMO-001 | DOOR MOTORISATION: swing | 1 | 1 | Internal - Door 3
OR-GEN | Operating Room - General | D+W | AFDPR-006.05 | DOOR PROTECTION: plate, full height | 1 | 1 | Internal - Door 1 | to both sides
OR-GEN | Operating Room - General | D+W | AFDPR-006.05 | DOOR PROTECTION: plate, full height | 1 | 1 | Internal - Door 2 | to both sides
OR-GEN | Operating Room - General | D+W | AFDPR-006.05 | DOOR PROTECTION: plate, full height | 1 | 1 | Internal - Door 3 | to both sides
OR-GEN | Operating Room - General | D+W | DOHI-014.03 | DOOR: hinged, 1 leaf, 950 clear opening, solid, standard vision panel, with cover | 1 | 1 | Internal - Door 3 | to store - sterile stock, access provided for stock delivery where sterile stock is not directly adjacent to operating room
OR-GEN | Operating Room - General | D+W | DOHI-212.03 | DOOR: hinged, 2 leaves, 1600 clear opening, solid, standard vision panel, with cover | 1 | 1 | Internal - Door 1 | to anaesthetic preparation room (see ANAES), cover required to vision panel if laser is to be used
OR-GEN | Operating Room - General | D+W | DOHI-213.03 | DOOR: hinged, 2 leaves, 1700 clear opening, solid, standard vision panel, with cover | 1 | 1 | Internal - Door 2 | to exit bay (see EXIT), cover required to vision panel if laser is to be used
OR-GEN | Operating Room - General | D+W | WIFX-058.04 | WINDOW: fixed, internal, double glazed, sill at 1050H, with integral blind | 1 | 1 | Internal - Window 1 | optional, to store - sterile stock
OR-GEN | Operating Room - General | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 2 | 1
OR-GEN | Operating Room - General | FFE | FIDI-106 | DISPENSER: PPE, disposable gloves, wall mounted | 2 | 1
OR-GEN | Operating Room - General | FFE | FIDI-181 | DISPENSER: cleaning wipes, wall mounted | 1 | 1
OR-GEN | Operating Room - General | FFE | FIDI-241 | DISPENSER: alcohol-based hand rub, wall mounted | 2 | 1
OR-GEN | Operating Room - General | FFE | FIHR-101 | BRACKET: suction bottle, wall mounted | 2 | 1 |  | optional - to medical services panel if provided
OR-GEN | Operating Room - General | FFE | FIHR-212 | BRACKET: computer display screen, single, articulated, wall mounted | 1 | 1 |  | to workstation for digital operating room (DOR) computer
OR-GEN | Operating Room - General | FFE | FIHR-327 | BRACKET: keyboard and mouse, articulated, wall mounted | 1 | 1 |  | to workstation for digital operating room (DOR) computer
OR-GEN | Operating Room - General | FFE | FIHR-331 | BRACKET: television/AV system display screen, single, fixed, wall mounted | 1 | 1 |  | to digital operating room (DOR) display screen
OR-GEN | Operating Room - General | FFE | FQCL-211 | LINEN CARRIER: dirty, single skip | 1 | 1
OR-GEN | Operating Room - General | FFE | ITCL-069 | CLOCK: electric, master clock system, slave, digital, dual time/elapsed time | 1 | 1
OR-GEN | Operating Room - General | FFE | ITNE-076 | TELEPHONE: handset, wall mounted | 2 | 1
OR-GEN | Operating Room - General | FFE | ITNE-101 | COMPUTER: single display screen, central processing unit (CPU), keyboard and mouse, desktop | 1 | 1 |  | workstation on wheels (WOW) may be provided depending on workflows
OR-GEN | Operating Room - General | FFE | ITNE-111 | COMPUTER: workstation on wheels (WOW), medical grade, all-in-one computer, keyboard and mouse, on cart | 2 | 1
OR-GEN | Operating Room - General | FFE | ITNE-122 | COMPUTER: single display screen, all-in-one, medical grade | 1 | 1 |  | workstation for digital operating room (DOR) computer
OR-GEN | Operating Room - General | FFE | ITNE-204 | INPUT DEVICE: keyboard and mouse, wireless, medical grade, washable | 1 | 1 |  | workstation for digital operating room (DOR) computer
OR-GEN | Operating Room - General | FFE | ITNE-224 | INPUT DEVICE: bar code scanner, wireless, with charging dock, medical grade, washable | 1 | 1
OR-GEN | Operating Room - General | FFE | JOBE-002 | BENCH: 450D, stainless steel | 1 | 1
OR-GEN | Operating Room - General | FFE | JOBE-231 | BENCH SUPPORT: legs and frame, stainless steel | 1 | 1
OR-GEN | Operating Room - General | FFE | MMGE-045 | POLE: IV, mobile | 2 | 1
OR-GEN | Operating Room - General | FFE | MMGE-191 | CANNISTER: suction bottle | 2 | 1 |  | optional - to medical services panel if provided
OR-GEN | Operating Room - General | FFE | MMGE-291 | DEVICE: pump, infusion, single | 1 | 1
OR-GEN | Operating Room - General | FFE | MMGE-367 | DEVICE: patient warming, air blanket | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-106 | TABLE: operating, mobile, electric | 1 | 1 |  | type of operating table to suit service requirements
OR-GEN | Operating Room - General | FFE | MMOP-121 | STOOL: mobile, operating | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-122 | STOOL: mobile, anaesthetic | 2 | 1
OR-GEN | Operating Room - General | FFE | MMOP-162 | MACHINE: anaesthetic, with patient monitor, high acuity | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-166 | TROLLEY: anaesthetic | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-171 | MACHINE: electrosurgical, diathermy | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-201 | STAND: bowl, double | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-202 | STAND: bowl, single | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-211 | KICKBUCKET: with stand | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-302 | DISPLAY SCREEN: digital operating room (DOR) integration, wall mounted | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-311 | CAMERA: patient observation, digital operating room (DOR) integration, ceiling mounted | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-331 | SPEAKER: digital operating room (DOR) integration, ceiling recessed | 2 | 1
OR-GEN | Operating Room - General | FFE | MMOP-341 | CONTROL PANEL: digital operating room (DOR) integration, wall mounted | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-351 | PATCH PANEL: digital room operating (DOR) integration, wall mounted | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-371 | SYSTEM CABINET: digital integration in operating rooms (DOR) | 1 | 1 |  | located in ICT/AV cupboard (see EXIT)
OR-GEN | Operating Room - General | FFE | MMPH-002 | SAFE: dangerous drugs, wall mounted, medium | 1 | 1 |  | optional - drug safe may alternatively be located in the anaesthetic preparation room; with card reader, safe location and type of medication stored is dependent on the operational policies and medication management regulations in different jurisdictions
OR-GEN | Operating Room - General | FFE | FQWS-005 | BIN: sharps, clinical, extra large | 1 | 1 |  | on cart
OR-GEN | Operating Room - General | FFE | FQWS-032 | CART: sharps bin, clinical, foot pedal bin operation | 1 | 1
OR-GEN | Operating Room - General | FFE | FQWS-396 | TROLLEY: waste bag | 1 | 1
OR-GEN | Operating Room - General | FFE | MMTR-108 | TROLLEY: set-up, table 1250W x 800D x 900H nom | 1 | 1
OR-GEN | Operating Room - General | FFE | MMTR-123 | TROLLEY: gowns/scrubs, 1220W x 610D x 1830H nominal | 1 | 1
OR-GEN | Operating Room - General | FFE | MMTR-171 | TROLLEY: mayo table | 1 | 1
OR-GEN | Operating Room - General | FFE | MMTR-003.02 | TROLLEY: stainless steel, 750W x 490D x 900H nom, 2 tier | 1 | 1
OR-GEN | Operating Room - General | FFE | MMTR-004.02 | TROLLEY: stainless steel, 900W x 490D x 900H nom, 2 tier | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-123 | STOOL: mobile, surgeon | 1 | 1
OR-GEN | Operating Room - General | FFE | MMOP-173 | MACHINE: smoke plume evacuation system | 1 | 1
OR-GEN | Operating Room - General | FIN | CLFI-003 | CEILING FINISH: paint, sterile areas | 1 | 1 | Ceiling Finish 2
OR-GEN | Operating Room - General | FIN | CLFS-011 | CEILING: flush set, suspended | 1 | 1 | Ceiling Finish 1
OR-GEN | Operating Room - General | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Finish 1
OR-GEN | Operating Room - General | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Finish 2 | inset section of contrasting colour to sterile zone
OR-GEN | Operating Room - General | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
OR-GEN | Operating Room - General | FIN | WLFI-011.07 | WALL FINISH: vinyl, full height | 1 | 1 | Wall Finish 1 | alternative impermeable seamless materials may also be acceptable, confirm material selection to suit infection prevention and control requirements
OR-GEN | Operating Room - General | MSP | ELGP-146 | GPO: single, emergency power, on service panel | 2 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | ELGP-146 | GPO: single, emergency power, on service panel | 4 | 1 | Panel 2 | quantity and type of power outlets to suit service requirements
OR-GEN | Operating Room - General | MSP | ELGP-151 | GPO: single, emergency power, on services pendant | 12 | 1 | Pendant 1 | quantity and type of power outlets to suit service requirements
OR-GEN | Operating Room - General | MSP | ELGP-151 | GPO: single, emergency power, on services pendant | 12 | 1 | Pendant 2 | quantity and type of power outlets to suit service requirements
OR-GEN | Operating Room - General | MSP | ELGP-176 | GPO: single, UPS, on service panel | 4 | 1 | Panel 2 | quantity and type of power outlets to suit service requirements
OR-GEN | Operating Room - General | MSP | ELGP-181 | GPO: single, UPS, on services pendant | 4 | 1 | Pendant 1 | quantity and type of power outlets to suit service requirements
OR-GEN | Operating Room - General | MSP | ELGP-181 | GPO: single, UPS, on services pendant | 4 | 1 | Pendant 2 | quantity and type of power outlets to suit service requirements
OR-GEN | Operating Room - General | MSP | ELPR-033 | PANEL: UPS status indicator, on services panel | 1 | 1 | Panel 2
OR-GEN | Operating Room - General | MSP | ELPR-073 | RCD: residual current device, on services panel | 1 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | ELPR-073 | RCD: residual current device, on services panel | 2 | 1 | Panel 2
OR-GEN | Operating Room - General | MSP | ITCL-185 | BUTTON: staff assist, with cancel, on services pendant | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | ITCL-185 | BUTTON: staff assist, with cancel, on services pendant | 1 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | ITCL-195 | BUTTON: emergency, with cancel, on services pendant | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | ITCL-195 | BUTTON: emergency, with cancel, on services pendant | 1 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | ITCL-213 | BUTTON: orderly / porter, with cancel, on services pendant | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | ITCL-213 | BUTTON: orderly / porter, with cancel, on services pendant | 1 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | ITIN-014 | OUTLET: data, single RJ45, on services pendant | 4 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | ITIN-014 | OUTLET: data, single RJ45, on services pendant | 4 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | ITIN-025 | OUTLET: data, double RJ45, on services panel | 3 | 1 | Panel 2
OR-GEN | Operating Room - General | MSP | MGAS-001 | OUTLET: carbon dioxide (CO2), on services pendant | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MGAS-001 | OUTLET: carbon dioxide (CO2), on services pendant | 1 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | MGAS-021 | OUTLET: medical air, on services pendant | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MGAS-021 | OUTLET: medical air, on services pendant | 1 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | MGAS-022 | OUTLET: medical air, on services panel | 2 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | MGAS-031 | OUTLET: nitrous oxide (N2O), on services pendant | 1 | 1 | Pendant 1 | optional; provision of reticulated nitrous oxide to be confirmed to suit project and service requirements
OR-GEN | Operating Room - General | MSP | MGAS-031 | OUTLET: nitrous oxide (N2O), on services pendant | 1 | 1 | Pendant 2 | optional; provision of reticulated nitrous oxide to be confirmed to suit project and service requirements
OR-GEN | Operating Room - General | MSP | MGAS-041 | OUTLET: oxygen (O2), on services pendant | 2 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MGAS-041 | OUTLET: oxygen (O2), on services pendant | 2 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | MGAS-051 | OUTLET: scavenge, on services pendant | 1 | 1 | Pendant 1 | optional; provided if reticulated nitrous oxide provided - provision of reticulated nitrous oxide to be confirmed to suit project and service requirements
OR-GEN | Operating Room - General | MSP | MGAS-051 | OUTLET: scavenge, on services pendant | 1 | 1 | Pendant 2 | optional; provided if reticulated nitrous oxide provided - provision of reticulated nitrous oxide to be confirmed to suit project and service requirements
OR-GEN | Operating Room - General | MSP | MGAS-061 | OUTLET: suction, on services pendant | 3 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MGAS-061 | OUTLET: suction, on services pendant | 3 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | MGAS-062 | OUTLET: suction, on services panel | 2 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | MGAS-071 | OUTLET: tool air, integral control panel, on services pendant | 1 | 1 | Pendant 1 | optional - depending on medical equipment to be used, may also be wall mounted
OR-GEN | Operating Room - General | MSP | MGAS-071 | OUTLET: tool air, integral control panel, on services pendant | 1 | 1 | Pendant 2 | optional - depending on medical equipment to be used, may also be wall mounted
OR-GEN | Operating Room - General | MSP | MGFP-021 | FLOWMETER: medical air | 2 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | MGFP-021 | FLOWMETER: medical air | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MGFP-021 | FLOWMETER: medical air | 1 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | MGFP-041 | FLOWMETER: oxygen | 2 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | MGFP-041 | FLOWMETER: oxygen | 2 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MGFP-041 | FLOWMETER: oxygen | 2 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | MGFP-061 | ADAPTER: suction | 2 | 1 | Panel 1 | optional
OR-GEN | Operating Room - General | MSP | MGFP-061 | ADAPTER: suction | 3 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MGFP-061 | ADAPTER: suction | 3 | 1 | Pendant 2
OR-GEN | Operating Room - General | MSP | MMOP-303 | DISPLAY SCREEN: digital operating room (DOR) integration, pendant mounted | 1 | 1 | Pendant 3 - Arm 2
OR-GEN | Operating Room - General | MSP | MMOP-303 | DISPLAY SCREEN: digital operating room (DOR) integration, pendant mounted | 1 | 1 | Pendant 4 - Arm 2
OR-GEN | Operating Room - General | MSP | MMOP-316 | CAMERA: surgical field, digital operating room (DOR) integration, integral to operating light | 1 | 1 | Pendant 3 - Arm 1
OR-GEN | Operating Room - General | MSP | MMSP-051 | MEDICAL SERVICES PANEL: wall mounted | 1 | 1 | Panel 1 | optional - provided for theatres where obstetric procedures are undertaken or services are otherwise required separately to those in the medical services pendants
OR-GEN | Operating Room - General | MSP | MMSP-051 | MEDICAL SERVICES PANEL: wall mounted | 1 | 1 | Panel 2
OR-GEN | Operating Room - General | MSP | MMSP-101 | MEDICAL SERVICES PENDANT: single arm, horizontal services head | 1 | 1 | Pendant 1 | ergonomic control handles required; configuration (horizontal or vertical) to suit services requirements; services allocations are indicative - type and number of services on the pendant to suit service requirements
OR-GEN | Operating Room - General | MSP | MMSP-101 | MEDICAL SERVICES PENDANT: single arm, horizontal services head | 1 | 1 | Pendant 2 | ergonomic control handles required; configuration (horizontal or vertical) to suit services requirements; services allocations are indicative - type and number of services on the pendant to suit service requirements
OR-GEN | Operating Room - General | MSP | MMSP-224 | MEDICAL SERVICES PENDANT: double arm, 1 light head, operating / display screen | 1 | 1 | Pendant 3
OR-GEN | Operating Room - General | MSP | MMSP-224 | MEDICAL SERVICES PENDANT: double arm, 1 light head, operating / display screen | 1 | 1 | Pendant 4
OR-GEN | Operating Room - General | MSP | ELGP-318 | ISOLATOR: three phase, with socket outlet, on services pendant | 1 | 1 | Pendant 2 | optional - depending on equipment requirements, may be wall mounted
OR-GEN | Operating Room - General | MSP | ELPR-034 | PANEL: UPS status indicator, on services pendant | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MMSP-091 | LIGHT HEAD: operating, articulated arm, pendant mounted | 1 | 1 | Pendant 3 - Arm 1
OR-GEN | Operating Room - General | MSP | MMSP-091 | LIGHT HEAD: operating, articulated arm, pendant mounted | 1 | 1 | Pendant 4 - Arm 1
OR-GEN | Operating Room - General | MSP | MMOP-352 | PATCH PANEL: digital room operating (DOR) integration, on services pendant | 1 | 1 | Pendant 1
OR-GEN | Operating Room - General | MSP | MMOP-363 | OUTLET: digital operating room (DOR) integration, on services pendant | 1 | 1 | Pendant 1 | optional - for connection of image capture device to digital operating room (DOR) system
OR-GEN | Operating Room - General | MSP | MMOP-363 | OUTLET: digital operating room (DOR) integration, on services pendant | 1 | 1 | Pendant 2 | optional - for connection of image capture device to digital operating room (DOR) system
OR-GEN | Operating Room - General | SER | ELGP-105 | GPO: single, wall mounted, cleaner | 1 | 1
OR-GEN | Operating Room - General | SER | ELGP-131 | GPO: single, emergency power, wall mounted | 1 | 1 |  | for recharging equipment
OR-GEN | Operating Room - General | SER | ELGP-161 | GPO: single, UPS, wall mounted | 1 | 1
OR-GEN | Operating Room - General | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 3 | 1
OR-GEN | Operating Room - General | SER | ELPR-011 | MONITOR: line isolation, wall mounted | 1 | 1 |  | to UPS power outlets on medical services pendants
OR-GEN | Operating Room - General | SER | ELSW-042 | SWITCH: light, laser in use | 1 | 1 |  | optional - to be provided if laser to be used, may be integrated on lighting control panel
OR-GEN | Operating Room - General | SER | ELSW-043 | SWITCH: light, x-ray in use | 1 | 1 |  | optional - to be provided if x-ray to be used, may be integrated on lighting control panel
OR-GEN | Operating Room - General | SER | ELSW-051 | CONTROL PANEL: lighting | 2 | 1
OR-GEN | Operating Room - General | SER | ITCL-081 | CONTROL PANEL: master clock system | 1 | 1
OR-GEN | Operating Room - General | SER | ITCL-181 | BUTTON: staff assist, with cancel, wall mounted | 1 | 1
OR-GEN | Operating Room - General | SER | ITCL-191 | BUTTON: emergency, with cancel, wall mounted | 1 | 1
OR-GEN | Operating Room - General | SER | ITCL-211 | BUTTON: orderly / porter, with cancel, wall mounted | 1 | 1
OR-GEN | Operating Room - General | SER | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 |  | to wall mounted telephone
OR-GEN | Operating Room - General | SER | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | 1
OR-GEN | Operating Room - General | SER | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | 1
OR-GEN | Operating Room - General | SER | ITSE-261 | READER: security, access control, proximity card, wall mounted | 1 | 1 |  | optional - to drug safe if provided
OR-GEN | Operating Room - General | SER | LIFX-212 | LIGHT: illuminated sign, laser in use | 3 | 1 |  | located outside doors, optional - to be provided if laser to be used
OR-GEN | Operating Room - General | SER | LIFX-213 | LIGHT: illuminated sign, x-ray in use | 3 | 1 |  | located outside doors, optional - to be provided if x-ray to be used
OR-GEN | Operating Room - General | SER | MEEX-041 | GRILLE: exhaust, wall mounted, low level | 4 | 1 |  | for removal of anaesthetic gases, rectilinear or triangular ducts are also acceptable
OR-GEN | Operating Room - General | SER | MGFP-001 | ALARM: medical gases, panel, wall mounted | 1 | 1
OR-GEN | Operating Room - General | SER | MMOP-346 | CONTROL PANEL: operating lights, wall mounted | 1 | 1
OR-GEN | Operating Room - General | SER | MMOP-362 | OUTLET: digital operating room (DOR) integration, wall mounted | 1 | 1 |  | to wall mounted display screen
This shape represents a table slicer. Table slicers are not supported in this version of Excel.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## SC Set for Publish (hidden)
Code
1BR-IC
1BR-BA