## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
ORGN | Operating Room - General | 42 |  | 1 patient; 5 - 8 staff | 24 Hours |  | The Operating Room - General provides an aseptic environment in which to carry out surgical procedures under local, regional or general anaesthetic. The General Operating Room may be used for General surgery and specialties including ENT, Urology, Gynaecology, Ophthalmology, Plastic Surgery and any other procedures that do not require bulky equipment. The Operating Room will have direct access to a Scrub-up Room, exit bay and sterile stock/set-up area. Cardiac Protected electrical area Special Air Conditioning - HEPA filtration Additional Design Considerations: - Minimum critical dimension 6000mm. - &apos;Laser In Use&apos; warning lights may be required outside all exit doors to room if laser equipment is to be used - Lead shielding may be required depending on use of X-ray equipment - to be assessed by Radiation Consultant - External windows are desirable - double glazed with integral Venetian blinds

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
ORGN | Operating Room - General | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 3010 | Wall finish | 1 |  | Paint | Scrub resistant |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 3038 | Wall finish | 1 |  | Vinyl | Prefinished | Wall vinyl to 900 AFFL | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 5010 | Ceiling - Plasterboard | 1 |  | Plasterboard | Paint, washable | Flush set, suspended | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 6015 | Cornice | 1 |  | Plasterboard | Paint | Flush coved between wall & ceiling | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 7010 | Observation | 1 | with cover (sliding) | Glass | Clear | Through door/s | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 7035 | Observation | 1 | double glazing with integral venetian, optional | Alum/ Glass | Powdercoat | Internal window | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 8010 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | Protection plate to 900 AFFL | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 8040 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | To door frame | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 9004 | Door - 1100mm | 1 | 910mm clear opening to set-up/scrub rooms | Solid Core/ Glass | Paint | Single leaf, observation panel | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 9004 | Door - 1100mm | 1 | 1100mm clear opening, to clean-up room | Solid Core/ Glass | Paint | Single leaf, observation panel | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | Fabrics | 9225 | Door - 1500mm | 1 | 1500mm clear opening to patient entry/exit | Solid Core/ Glass | Paint | Double leaf, observation panel | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-11400 | Computer - PC & monitor | 1 | optional |  |  |  | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-16000 | Diathermy machine | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-2795 | Television (large flatscreen) | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-32600 | Light - operating, 2 arm, 950/550 diam nominal | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-35960 | Monitor - anaesthetic | 1 | with printer, on anaesth machine |  |  |  | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-47000 | Services pendant - ceiling | 2 | cable support arm optional |  |  |  | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-54520 | Telephone - handset, handsfree | 1 | to bench |  |  |  | No | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-55750 | Tourniquet - surgical | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-68900 | Xray viewer - 4 screen | 1 | recessed, 4x1 layout |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-700 | Anaesthetic machine | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-9750 | Clock - slave, analogue | 1 | recessed |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-9800 | Clock - time elapse | 1 | beside analogue, recessed |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-10200 | Trolley - instrument | 4 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-1102 | Bench - resinate, mobile | 1 | for computer, optional |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-11100 | IV stand - mobile | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-11500 | Kickbucket & stand | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-12300 | Linen carrier - dirty, single | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-1252 | Blind - Venetian integral | 1 | to internal window, if provided |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-13500 | Table - mayo | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-1440 | Bowl stand - single | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-1465 | Bracket - suction bottle | 4 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-150 | Air flowmeter | 2 | to services pendants |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-15100 | Bracket - monitor | 1 | optional, for TV monitor if required |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-1580 | Chair - anaesthetist | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-16000 | Operating table - manual | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-16250 | Oxygen flowmeter | 4 | to services pendants |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-21730 | Stool - operating, mobile | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-21950 | Suction adapter | 4 | to services pendants |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-22000 | Suction bottle | 4 | in mobile carrier |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-22050 | Suction bottle carrier | 1 | mobile |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-24405 | Whiteboard - magnetic | 1 | 900x900, optional |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-300 | Trolley - anaesthetic | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-6152 | Dispenser - disposable glove | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-8390 | Trolley - gowns | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-1006 | Voice / Data outlet: double | 1 | to bench; 1 outlet for telephone |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-1006 | Voice / Data outlet: double | 2 | 1 / each pendants |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-11000 | Wall panel | 4 | flush mounted |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-1120 | Monitor connection facilities and cabling | 1 | to wall TV monitor, optional |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-1416 | Light Control Panel | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-5005 | Air conditioning: special | 1 | HEPA filtered |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-5040 | Pressure: positive | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-6001 | General: colour corrected | 1 | variable lighting level, dimmable |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-6074 | Special: 'Laser In Use' lights | 4 | outside room entries, if laser to be used |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7000 | Oxygen (O2) | 4 | 2 / each pendant |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7010 | Nitrous Oxide (N2O) | 2 | 1 / each pendant |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7020 | Medical Air (MA) | 2 | 1 / each pendant |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7030 | Tool Air | 2 | 1 / each pendant, controls on wall panel |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7040 | Suction | 4 | 2 / each pendant |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7050 | Gas Alarm Panel | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7060 | Tourniquet air | 1 | and controls, on wall, optional |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-7070 | Scavenge | 2 | 1 / each pendant |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-8010 | Staff/ Nurse assist call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-8080 | Emergency call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-8090 | Orderly/ Porter call | 1 | optional |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9000 | GPO: Single | 4 | on walls |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9000 | GPO: Single | 8 | 4 per pendant |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9001 | GPO: Double | 1 | to bench, optional |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9010 | GPO: Emergency power, single (V) | 1 | to wall TV monitor, optional |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9010 | GPO: Emergency power, single (V) | 2 | on walls |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9010 | GPO: Emergency power, single (V) | 8 | 4 per pendant |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9015 | GPO: Cleaner | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9030 | Cardiac protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9050 | Direct connection | 3 | XRV, clocks, lights |  |  |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | SE | S-9060 | GPO: Special: for equipment | 1 | 20 amp, for laser, optional |  |  |  | No | No | No | No | No | No | No | No | No