## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
PET-CT | PET-CT Imaging Room | 50 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | Room for Positron Emission Tomography-Computed Tomography (PET/CT) scanning procedures with a Control Room located nearby and PET/CT Equipment Room. Control Room location will be determined in part by acceptable cable lengths between Control Room and PET/CT modality. Direct adjacency to the Control Room is not essential where patients are monitored via observation cameras. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Cable ducts between Control Room and Equipment Room will be required; consider cable lengths where control room is remote. 'X-ray in Use' lights are required outside each entry door. Dimmable downlights are recommended, placed to avoid glare to patients. Equipment should be installed to manufacturer's specifications with sufficient clearances for maintenance access. Consider providing a separate Equipment Room if space permits as this will enhance the acoustics in the room. Laser marker system for procedure planning may be required subject to model of care and equipment selected. Structural assessment is required to ensure the floor will be rigid with no vibration and will accommodate the weight of the PET/CT unit. UPS power is required to the PET scanning and control equipment. Room temperature and humidity control is required.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
PET-CT | PET-CT Imaging Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
PET-CT | PET-CT Imaging Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
PET-CT | PET-CT Imaging Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
PET-CT | PET-CT Imaging Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 4
PET-CT | PET-CT Imaging Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
PET-CT | PET-CT Imaging Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin, splashback material to comply with AusHFG Part D
PET-CT | PET-CT Imaging Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PET-CT | PET-CT Imaging Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 2
PET-CT | PET-CT Imaging Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1 | automated
PET-CT | PET-CT Imaging Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PET-CT | PET-CT Imaging Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 8
PET-CT | PET-CT Imaging Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | on services panel
PET-CT | PET-CT Imaging Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2
PET-CT | PET-CT Imaging Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 1
PET-CT | PET-CT Imaging Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door
PET-CT | PET-CT Imaging Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIHR-029 | HOOK: equipment | 2
PET-CT | PET-CT Imaging Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | on services panel
PET-CT | PET-CT Imaging Room | FFE | MMBE-042 | SLIDE: patient | 2
PET-CT | PET-CT Imaging Room | FFE | MMGE-142 | TROLLEY: sharps bin | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-091 | SPHYGMOMANOMETER: portable | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-002 | ADAPTER: suction | 2 | on services panel
PET-CT | PET-CT Imaging Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
PET-CT | PET-CT Imaging Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
PET-CT | PET-CT Imaging Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
PET-CT | PET-CT Imaging Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
PET-CT | PET-CT Imaging Room | FFE | MMBE-010 | BED: patient, electric with mattress | 1
PET-CT | PET-CT Imaging Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
PET-CT | PET-CT Imaging Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-205 | CAMERA: patient observation, PET/CT | 2
PET-CT | PET-CT Imaging Room | FFE | MMMI-207 | DEVICE: imaging, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-209 | PHANTOM: PET/CT | 1 | stored within shielded custom joinery
PET-CT | PET-CT Imaging Room | FFE | MMMI-217 | TABLE: patient, fixed, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-100 | INJECTOR: contrast, CT | 1 | optional
PET-CT | PET-CT Imaging Room | FFE | MMMI-218 | LASER: positioning lights, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-213 | PANEL: joinery carcass end | 3
PET-CT | PET-CT Imaging Room | FFE | MMMI-377 | WARMER: contrast | 1 | optional
PET-CT | PET-CT Imaging Room | FFE | FIJO-218 | CUPBOARD: tall, double door, adj shelves, lockable, heavy duty | 1
PET-CT | PET-CT Imaging Room | FFE | ITCL-035 | CAMERA: respiratory gating | 1
PET-CT | PET-CT Imaging Room | FFE | FIHR-082 | BRACKET: for observation camera, MRI compatible, wall mounted | 3
PET-CT | PET-CT Imaging Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 5
PET-CT | PET-CT Imaging Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | on wall near PET-CT unit
PET-CT | PET-CT Imaging Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room
PET-CT | PET-CT Imaging Room | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
PET-CT | PET-CT Imaging Room | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 2
PET-CT | PET-CT Imaging Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | 1-laser positioning light
PET-CT | PET-CT Imaging Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
PET-CT | PET-CT Imaging Room | SE | ELBO-001 | BUTTON: emergency stop | 1
PET-CT | PET-CT Imaging Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional, on services panel
PET-CT | PET-CT Imaging Room | SE | MEGE-007 | CONTROL PANEL: temperature | 1
PET-CT | PET-CT Imaging Room | SE | LIFX-003 | LIGHT: illuminated sign, "Laser in Use" | 2
PET-CT | PET-CT Imaging Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 3 | to outside of all doors to room, auto on when modality is in use
PET-CT | PET-CT Imaging Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional, on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
PET-CT | PET-CT Imaging Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel + indicator button & light
PET-CT | PET-CT Imaging Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | 1-clock
PET-CT | PET-CT Imaging Room | SE | ELGP-208 | GPO: double, wall mounted | 2
PET-CT | PET-CT Imaging Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | to walls
PET-CT | PET-CT Imaging Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 2
PET-CT | PET-CT Imaging Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
PET-CT | PET-CT Imaging Room | SE | ELGP-203 | GPO: double, in floorbox | 4
PET-CT | PET-CT Imaging Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
PET-CT | PET-CT Imaging Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | patient services
PET-CT | PET-CT Imaging Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to room doors, walls, (ceiling) as required by radiation consultant design and specification
PET-CT | PET-CT Imaging Room | SE | ELBO-012 | SWITCH: light, dimmer | 3 | number determined by number of doors
PET-CT | PET-CT Imaging Room | SE | ELGP-133 | GPO: UPS, single, on services panel | 2 | on service panel
PET-CT | PET-CT Imaging Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 6 | on service panel
PET-CT | PET-CT Imaging Room | SE | ELPR-007 | UPS: status indicator | 1 | on service panel