## Room Data
Standard Components - Room Data Sheet
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
PET-CT | PET-CT Imaging Room | 50 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | The PET-CT Imaging Room provides the area and facilities for Positron Emission Tomography-Computed Tomography (PET/CT) scanning procedures with a Control Room and Equipment Room (as required) located nearby. It is not essential that the control is directly adjacent to the Imaging Room where patients are monitored via observation cameras, however ready access is required to attend to patient needs. If the Control Room is adjacent an observation window may be provided and shielding requirements will need to be confirmed. Cable pathways for services between the Control Room, Equipment Room and Imaging Room must be established and coordinated prior to construction and developed based on vendor requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access and support structure as required. The PET-CT Imaging Room must be adequately shielded to ensure that radiation doses to surrounding areas do not exceed radiation dose limits and satisfy design/dose constraints. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. Additional Considerations: Configuration of medical gases is to be confirmed to meet clinical requirements such as the provision of anaesthetic services. Dimmable lights are recommended and are to be placed to avoid glare to patients. Back lit ceiling artwork may be installed for patient comfort. 'X-ray in Use' lights are required outside each entry door. Laser marker system for procedure planning may be required subject to model of care and equipment selected. Structural assessment is required to ensure the floor will be rigid with no vibration and will accommodate the weight of the PET-CT unit. Room temperature and humidity control is required. Temperature requirements for equipment are to be confirmed by the vendor, including requirement for circulated cooling water. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. UPS power is required to the PET-CT scanning and control equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. An optional laser system for patient positioning and procedure planning has been indicated and may be provided subject to service requirements. A mix of respiratory gating, surface guidance, motion management and/or patient position technologies may be considered depending on clinical/service requirements (e.g. for radiation therapy planning). The provision of these technologies is to be confirmed by the project and may have additional requirements including additional power/data, water cooling, heat load, possible increased set down, increased number of trenches/conduits, space requirements for computer cabinets in control room and additional penetrations through the suspended ceiling. Requirements to be confirmed with vendors. Consideration needs to be given to the clearances of equipment and replacement part delivery routes.

## Items
Standard Components - Items in room
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
PET-CT | PET-CT Imaging Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
PET-CT | PET-CT Imaging Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 10
PET-CT | PET-CT Imaging Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
PET-CT | PET-CT Imaging Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
PET-CT | PET-CT Imaging Room | FA | WLWA-007 | WALL FINISH: splashback | 1
PET-CT | PET-CT Imaging Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
PET-CT | PET-CT Imaging Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 6
PET-CT | PET-CT Imaging Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PET-CT | PET-CT Imaging Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1 | automated
PET-CT | PET-CT Imaging Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1 | automated
PET-CT | PET-CT Imaging Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PET-CT | PET-CT Imaging Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 1
PET-CT | PET-CT Imaging Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 | on services panel
PET-CT | PET-CT Imaging Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2
PET-CT | PET-CT Imaging Room | FFE | FQBS-030 | CHAIR: visitor, patient | 1
PET-CT | PET-CT Imaging Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-141 | CUPBOARD: underbench, single door, adj shelves, lockable | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 3
PET-CT | PET-CT Imaging Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door
PET-CT | PET-CT Imaging Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PET-CT | PET-CT Imaging Room | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-045 | IV STAND: mobile | 1 | may be ceiling mounted
PET-CT | PET-CT Imaging Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
PET-CT | PET-CT Imaging Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | on services panel
PET-CT | PET-CT Imaging Room | FFE | MMBE-042 | SLIDE: patient | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-142 | TROLLEY: sharps bin | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-091 | SPHYGMOMANOMETER: portable | 1
PET-CT | PET-CT Imaging Room | FFE | MMGE-002 | ADAPTER: suction | 2 | on services panel
PET-CT | PET-CT Imaging Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
PET-CT | PET-CT Imaging Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
PET-CT | PET-CT Imaging Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
PET-CT | PET-CT Imaging Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
PET-CT | PET-CT Imaging Room | FFE | ITCL-012 | CAMERA: observation | 3
PET-CT | PET-CT Imaging Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
PET-CT | PET-CT Imaging Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-207 | DEVICE: imaging, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-209 | PHANTOM: PET/CT | 1 | stored within shielded custom joinery
PET-CT | PET-CT Imaging Room | FFE | MMMI-217 | TABLE: patient, fixed, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-107 | INJECTOR: contrast, CT, ceiling mounted | 1 | also may be mobile or wall mounted
PET-CT | PET-CT Imaging Room | FFE | MMMI-218 | LASER: positioning lights, PET/CT | 1
PET-CT | PET-CT Imaging Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
PET-CT | PET-CT Imaging Room | FFE | MMMI-377 | WARMER: contrast | 1 | optional
PET-CT | PET-CT Imaging Room | FFE | FIJO-218 | CUPBOARD: tall, double door, adj shelves, lockable, heavy duty | 1
PET-CT | PET-CT Imaging Room | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 2
PET-CT | PET-CT Imaging Room | FFE | FIHR-078 | BRACKET: for observation camera, wall mounted | 1
PET-CT | PET-CT Imaging Room | FFE | FIHR-092 | HOOK: anchor, patient ladder | 1
PET-CT | PET-CT Imaging Room | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 2
PET-CT | PET-CT Imaging Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3
PET-CT | PET-CT Imaging Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 2 | on wall near PET-CT unit
PET-CT | PET-CT Imaging Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room
PET-CT | PET-CT Imaging Room | SE | ELPR-006 | RCD: residual current device | 1 | on services panel
PET-CT | PET-CT Imaging Room | SE | ITIN-143 | FLOORBOX: for power, data, AV, flush mounted | 2
PET-CT | PET-CT Imaging Room | SE | ELGP-101 | GPO: single, ceiling mounted | 3 | 1-laser positioning light
PET-CT | PET-CT Imaging Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
PET-CT | PET-CT Imaging Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional, on services panel
PET-CT | PET-CT Imaging Room | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | alarm for temperature range
PET-CT | PET-CT Imaging Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 2 | to outside of all doors to room, auto on when modality is in use, vendor to provide signal wiring to light
PET-CT | PET-CT Imaging Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional, on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
PET-CT | PET-CT Imaging Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
PET-CT | PET-CT Imaging Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel + indicator button & light
PET-CT | PET-CT Imaging Room | SE | ELGP-109 | GPO: single, wall mounted | 2 | 1-clock
PET-CT | PET-CT Imaging Room | SE | ELGP-208 | GPO: double, wall mounted | 3
PET-CT | PET-CT Imaging Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | to walls
PET-CT | PET-CT Imaging Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 1
PET-CT | PET-CT Imaging Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
PET-CT | PET-CT Imaging Room | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
PET-CT | PET-CT Imaging Room | SE | ELGP-203 | GPO: double, in floorbox | 4
PET-CT | PET-CT Imaging Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
PET-CT | PET-CT Imaging Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | patient services
PET-CT | PET-CT Imaging Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to room doors, walls, (ceiling) as required by radiation consultant design and specification
PET-CT | PET-CT Imaging Room | SE | ELBO-012 | SWITCH: light, dimmer | 1 | number determined by number of doors
PET-CT | PET-CT Imaging Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PET-CT | PET-CT Imaging Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4 | on service panel
PET-CT | PET-CT Imaging Room | SE | ELPR-007 | UPS: status indicator | 1 | on service panel
PET-CT | PET-CT Imaging Room | SE | LIFX-038 | LIGHT: backlit feature panels | 2