## Room Data
Standard Components - Room Data Sheet
Date of issue: 31.01.2024
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
PT-RS1 | Patient Bay - Recovery, Stage 1 | 9 | up to 24 Hours | 1 patient; 1-2 staff |  |  | A Patient Bay - Recovery, Stage 1 provides facilities for the recovery of patients from anaesthetic post procedure, and is located within the Operating Unit or Day Procedures Unit. A handwashing basin will be provided close by to be shared between up to 4 Patient Bays. Staff observation is required and a clear line of sight from a staff station should be provided. Additional Considerations: The configuration of the Medical Services Panel (MSP), including medical gases, power and data, will be dependent on service requirements. Data provision is to take into consideration the increasing demand and sophistication of information systems and the technological development of monitoring and diagnostic equipment. A patient monitor may be located on the left or right side of bedhead depending on monitor functionality and staff workflows. Appropriate wall supports are required for mounting the bracket where a monitor is provided. Provision of emergency/standby power or UPS for patient monitoring is to be confirmed to suit site and service requirements. Confirmation to be based on risk assessment considering the impact of a power outage on patient care/safety. Facilities to accommodate bariatric patients within a select number of bays, including accommodating bariatric beds, is to be considered to ensure sufficient amenity is provided. Provision will be dependent on service and patient cohort requirements. Patient transfer aids such as mobile hoists, patient slides and hover mats are to be considered for assisted patient transfers to and from beds in holding bays. A ceiling mounted hoist may also be considered for patient transfers. Selection of transfer aids will be dependent on service and patient cohort requirements as well as WHS risk assessment. Appropriate ceiling support is required for the provision of a ceiling mounted hoist. Depending on service requirements, enclosed recovery bays may be provided. Where a Patient Bay is to be enclosed, additional area is required to accommodate a dedicated handwashing basin and circulation at the door. Observation windows may be required and the extent of glazing to be provided is to suit requirements for staff supervision. Additional area may be required in a percentage of Recovery Stage 1 bays to suit various services requirements, e.g. to support maternity patients or paediatric patients, for accommodating bariatric patients, for care of complex patients, or for extended recovery models. The percentage of bays that may be provided with increased area will be dependent on the anticipated case mix. Electronic Medical Record (EMR) management may be performed on a wirelessly connected workstation on wheels (WOW) that is shared between a number of bays. Ensure circulation space is sufficient for bedside and bed end data entry. Parking for WOWs and facilities for battery charging required to support this model of data entry. Alternative configurations, such as wall mounted monitors in bays, may be considered depending on local ICT policies. Access to natural light is required. Where windows are provided, integral blinds are required for infection prevention and control.

## Items
Standard Components - Items in room
Date of issue: 31.01.2024
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | WLPR-001 | WALL PROTECTION: bed protection wall panel | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MGFP-002 | FLOWMETER: medical air | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-019 | BRACKET: for suction bottle | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-029 | HOOK: equipment | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-119 | TROLLEY: equipment | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-045 | IV STAND: mobile | 1 | bed mounted IV pole or ceiling mounted IV track are also acceptable
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-033 | PANEL: for storage bins, wall mount | 1 | type and extent of storage dependant on operational policies for stock centralisation/decanting and infection prevention and control policies
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MGFP-004 | FLOWMETER: oxygen | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-142 | TROLLEY: sharps bin | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FQSN-006 | BIN: storage bins, for wall mounted panel | 4 | type and size of bins dependant on items to be stored
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-002 | ADAPTER: suction | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-100 | SUCTION BOTTLE | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMBE-007 | BED: patient, day procedure, electric | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMHA-008 | MONITOR: patient, cardiac with vital signs | 1 | may be located on left or right side of bedhead depending on monitor functionality and staff workflows; appropriate wall support required for mounting bracket
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | shared between multiple bays
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1 | may be located on left or right side of bedhead depending on monitor functionality and staff workflows; appropriate wall support required
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIHR-068 | BRACKET: for alcohol based hand rub, attached to bed | 1 | may be wall mounted depending on bedhead arrangment
PT-RS1 | Patient Bay - Recovery, Stage 1 | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to patient monitor
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | LIFX-032 | LIGHT: downlight, examination, ceiling mounted | 1 | articulated examination light is also acceptable
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel, with indicator button and light
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELBO-015 | SWITCH: light, single | 1 | on services panel, for examination downlight
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MGAS-062 | OUTLET: suction, on services panel | 2
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on services panel, with indicator button and light
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on services panel
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | with indicator button and light
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELGP-208 | GPO: double, wall mounted | 1 | low, for patient bed and pressure mattress
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 1 | to patient monitor
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | quantity to suit service requirements
PT-RS1 | Patient Bay - Recovery, Stage 1 | SE | ELGP-122 | GPO: emergency power, single, on services panel | 7 | type and quantity to suit service requirements