## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
SCRN | Screening Room (Fluoroscopy) | 36 |  | 1 patient; 1 - 2 staff | Available 24 Hours |  | Screening Room (Fluoroscopy) is a room for undertaking radiographic procedures involving contrast media. Patient Change, Toilets and Shower should be located in close proximity. The room may include a preparation area with sink. Additional Design Requirements: - Radiation shielding to walls and glazing as advised by Radiation Consultant - Unistrut type ceiling may be required, depending on equipment to be installed - Cable ducting between the X-ray table, vertical bucky and control area may be required, depending on equipment type and layout - &apos;X-ray in Use&apos; warning lights will be required outside each entry door Body Protected electrical area

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
SCRN | Screening Room (Fluoroscopy) | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 3060 | Wall finish | 1 | to basin & sink, tiles also acceptable | Vinyl | Prefinished | Splashback | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 5030 | Ceiling - Plasterboard | 1 | at 3000 mm AFFL | Plasterboard | Prefinished | Drop-in tiles, 600x1200 | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 6030 | Cornice | 1 |  | Aluminium | Powdercoat | 24mm shadow angle | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 7025 | Observation | 1 | to control booth | Lead glass | Clear | Internal window | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 8010 | Door protection | 1 | Equal to Acrovyn | Composite | Prefinished PVC | Protection plate to 900 AFFL | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 8040 | Door protection | 1 | Equal to Acrovyn | Composite | Prefinished PVC | To door frame | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 9001 | Door - 1100mm | 1 | to Patient Change/ Toilet | Solid Core | Paint | Single leaf | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | Fabrics | 9130 | Door - 1600mm | 1 | Patient entry/exit 1400mm clear opening | Solid Core | Paint | 1 1/2 leaf, with lead shielding | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-11415 | Computer - PACS | 1 | flat screen |  |  |  | Yes | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-2800 | Basin - handwash, Type B | 1 | refer to BHWS-B (sim) |  |  |  | No | No | Yes | No | Yes | Yes | Yes | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-4045 | TAPWARE - ELBOW MIXER: TAP MIXER, LONG LEVER | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-47800 | Sink, SS - single in bench | 1 | warm water |  |  |  | No | No | Yes | Yes | Yes | Yes | Yes | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-54520 | Telephone - handset, handsfree | 1 |  |  |  |  | No | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-61600 | IMG - bucky, chest | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-62152 | IMG - console, control (integrated stand) | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-62400 | IMG - control booth | 1 | with lead glazing |  |  |  | Yes | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-66000 | IMG - X-ray table, floating top | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-67000 | IMG - Screening (Fluoroscopy) imaging unit | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FE | FE-68400 | IMG: generator | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-1010 | Bench - stainless steel | 1 | with wet edge |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-1090 | Bench - laminate, free standing | 1 | workstation |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-11460 | Keyboard shelf - slide out | 1 | for PACS computer screen |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-11905 | Rack - lead apron | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-1480 | Bracket - ECG monitor | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-150 | Air flowmeter | 1 | to services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-16250 | Oxygen flowmeter | 1 | to services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-1650 | Chair - client | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-21950 | Suction adapter | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-22000 | Suction bottle | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-4410 | Cupboard - under bench | 1 | adjustable shelf |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-4420 | Cupboard - under bench, lockable | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-4470 | Cupboard - wall mounted, lockable | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-5824 | Waste bin - general | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-6152 | Dispenser - disposable glove | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-6154 | Dispenser - paper/ hand towel | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-6156 | Dispenser - soap | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FF | FF-6158 | Dispenser - soap antiseptic | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-1006 | Voice / Data outlet: double | 3 | 1 outlet for telephone |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-1040 | Intercom | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-11000 | Wall panel | 1 | flush mounted |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-12030 | Lead lining | 1 | to walls and doors as specified |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-5000 | Airconditioning | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-6001 | General: colour corrected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-6110 | Downlights: dimmable | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-6200 | Light switches | 5 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-7000 | Oxygen (O2) | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-7020 | Medical Air (MA) | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-7040 | Suction | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-8010 | Staff/ Nurse assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-8080 | Emergency call | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-9000 | GPO: Single | 2 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-9001 | GPO: Double | 5 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-9010 | GPO: Emergency power, single (V) | 2 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-9030 | Cardiac protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-9040 | Body protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | S-9060 | GPO: Special: for equipment | 2 | Xray Fluoroscopy Equipment, chest bucky |  |  |  | No | No | No | No | No | No | No | No | No