## Room Data
Standard Components - Room Data Sheet
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
SPECT-CTCR | SPECT-CT Control Room | 14 | Available 24 hours | Up to 3 staff |  | Acoustic privacy required | The SPECT-CT Control room is located adjacent to the SPECT-CT Imaging Room, with direct visibility to the patient, for remote control of equipment with audio and visual contact with the patient via cameras, speakers and microphones. Each Imaging Room requires a Control Room. Cable pathways for services between the Control Room, Equipment Room and Imaging Room must be established and coordinated prior to construction and developed based on vendor requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. Additional Considerations: The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. UPS power is required to the SPECT-CT scanning and control equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. Some models of SPECT/CT require three phase power to be supplied. Cable management and the potential for computers to be consolidated into a rack (either within the room or in a dedicated, collocated air conditioned room/cupboard) to be determined at project level. Provision of a height adjustable workstation may be considered depending on local policies and staff accessibility requirements.

## Items
Standard Components - Items in room
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
SPECT-CTCR | SPECT-CT Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
SPECT-CTCR | SPECT-CT Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
SPECT-CTCR | SPECT-CT Control Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
SPECT-CTCR | SPECT-CT Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
SPECT-CTCR | SPECT-CT Control Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
SPECT-CTCR | SPECT-CT Control Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | safety glass
SPECT-CTCR | SPECT-CT Control Room | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1
SPECT-CTCR | SPECT-CT Control Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
SPECT-CTCR | SPECT-CT Control Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 2
SPECT-CTCR | SPECT-CT Control Room | FFE | DWWD-004 | BLIND: roller | 1 | to observation window for patient privacy
SPECT-CTCR | SPECT-CT Control Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-044 | RACK: lead apron | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FQSM-007 | PEDESTAL: mobile, drawer unit, 660H | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | MMGE-032 | DISPLAY SCREEN: slave, for patient monitor | 1 | optional
SPECT-CTCR | SPECT-CT Control Room | FFE | FQGE-056 | PRINTER: desktop, labels | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1 | communication with patient, on bench required if integrated intercom system not provided by equipment manufacturer
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 1 | communication with patient, on bench required if integrated intercom system not provided by equipment manufacturer
SPECT-CTCR | SPECT-CT Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 2
SPECT-CTCR | SPECT-CT Control Room | FFE | MMMI-269 | COMPUTER: tower, acquisition, SPECT/CT | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | MMMI-274 | CONSOLE: control, SPECT/CT | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-066 | BRACKET: for display screen, wall mounted, articulated, adj height | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | MMMI-369 | DISPLAY SCREEN: contrast injector | 1 | optional
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 2
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 1
SPECT-CTCR | SPECT-CT Control Room | FFE | ITNE-070 | DISPLAY SCREEN: PACS | 2
SPECT-CTCR | SPECT-CT Control Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3 | 1 outlet for display screen, patient monitor slave; one outlet for telephone
SPECT-CTCR | SPECT-CT Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 4
SPECT-CTCR | SPECT-CT Control Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1 | mains power, consult manufacturer for safety and compatibility with equipment.
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-135 | GPO: UPS, single, wall mounted | 5
SPECT-CTCR | SPECT-CT Control Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1 | auto on when modality is in use
SPECT-CTCR | SPECT-CT Control Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-208 | GPO: double, wall mounted | 1 | one for printer
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 4
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
SPECT-CTCR | SPECT-CT Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 2