## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
SPECT-CT | SPECT-CT Imaging Room | 48 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | A Room for SPECT/CT scanning procedures with an adjacent Control Room and Equipment Room. The Control Room and Equipment room will need to be located in close proximity due to acceptable cable lengths. The Control Room may be shared between two SPECT/CT Scanning Rooms. The SPECT/CT room will be located with ready access to patient dosing areas. General anaesthesia may be administered in this room. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Cable ducts between Control Room, Equipment room and imaging room will be required. Heat and noise generated by equipment to be considered for patient comfort may require equipment room. Some models of SPECT/CT require three phase power to be supplied. Observation window required where patient observation cameras are not in use. 'X-ray in Use' lights are required outside each entry door. Dimmable downlights are recommended, placed to avoid glare to patients. Equipment should be installed to manufacturer's specifications with sufficient clearance for equipment maintenance. UPS power is required to the SPECT camera and control equipment. Room temperature and humidity control is required. Consideration to be given to equipment delivery route clearances.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
SPECT-CT | SPECT-CT Imaging Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
SPECT-CT | SPECT-CT Imaging Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
SPECT-CT | SPECT-CT Imaging Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 4 | optional
SPECT-CT | SPECT-CT Imaging Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 2 | to basin, splashback material to comply with AusHFG Part D
SPECT-CT | SPECT-CT Imaging Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
SPECT-CT | SPECT-CT Imaging Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | optional, patient observation by camera may be considered, glazing to be AS1428 compliant safety glass
SPECT-CT | SPECT-CT Imaging Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 7
SPECT-CT | SPECT-CT Imaging Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMOP-015 | CYLINDER: gas, argon | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-029 | HOOK: equipment | 1 | slide board
SPECT-CT | SPECT-CT Imaging Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMBE-042 | SLIDE: patient | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-142 | TROLLEY: sharps bin | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-002 | ADAPTER: suction | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-262 | DEVICE: nuclear medicine, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMBE-010 | BED: patient, electric with mattress | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | ITCL-012 | CAMERA: observation | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-264 | CART: collimator, SPECT/CT | 3
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-267 | TABLE: patient handling system, pivot, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-265 | BOARD: distribution, mains, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-263 | BOX: line connection, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-100 | INJECTOR: contrast, CT | 1 | optional
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-272 | BOARD: disconnect, mains, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-273 | PHANTOM: SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-377 | WARMER: contrast | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-071 | BRACKET: for E2/G2 cylinder, 1 cylinder | 1 | for argon
SPECT-CT | SPECT-CT Imaging Room | FFE | FQSM-044 | CUPBOARD: mobile, double door, adjustable shelves | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-078 | BRACKET: for observation camera, wall mounted | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMHA-032 | MONITOR: patient, cardiac, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 3
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | on wall near SPECT/CT unit
SPECT-CT | SPECT-CT Imaging Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
SPECT-CT | SPECT-CT Imaging Room | SE | ELBO-001 | BUTTON: emergency stop | 1
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional, on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MEEX-008 | EXHAUST: low level | 1 | optional
SPECT-CT | SPECT-CT Imaging Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 2 | auto on when modality is in use
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional, on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel + indicator button & light
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | clock
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | to walls
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 2
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ELGT-007 | UPS: uninterruptable power supply unit | 1 | optional
SPECT-CT | SPECT-CT Imaging Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
SPECT-CT | SPECT-CT Imaging Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | patient services
SPECT-CT | SPECT-CT Imaging Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors extent as specified by radiation consultant
SPECT-CT | SPECT-CT Imaging Room | SE | ELBO-012 | SWITCH: light, dimmer | 2 | number determined by number of doors
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-133 | GPO: UPS, single, on services panel | 2
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 6
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-007 | UPS: status indicator | 1