## Room Data
Standard Components - Room Data Sheet
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
SPECT-CT | SPECT-CT Imaging Room | 50 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | The SPECT-CT Imaging Room provides the area and facilities for Single Photon Emission Computed Tomography (SPECT) scanning procedures with an adjacent Control Room and Equipment Room. The SPECT-CT Imaging Room will be located with ready access to patient dosing areas. Cable pathways for services between the Control Room, Equipment Room and Imaging Room must be established and coordinated prior to construction and developed based on vendor requirements. Equipment is to be installed to manufacturer's specifications with sufficient clearances for maintenance access. The SPECT-CT Imaging Room must be adequately shielded to ensure that radiation doses to surrounding areas do not exceed radiation dose limits and satisfy design/dose constraints. Where shielding significantly increases the weight of the doors, automatic doors may be considered. Radiation shielding is to be as advised by a qualified Radiation Consultant to suit the practice, workload and equipment procured, and all shielding on the Standard Component is indicative only. Additional Considerations: Configuration of medical gases is to be confirmed to meet clinical requirements such as the provision of anaesthetic services. Dimmable lights are recommended and are to be placed to avoid glare to patients. Back lit ceiling artwork may be installed for patient comfort. 'X-ray in Use' lights are required outside each entry door. Room temperature and humidity control is required. Temperature requirements for equipment are to be confirmed by the vendor, including requirement for circulated cooling water. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. UPS power is required to the SPECT-CT scanning and control equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. Some models of SPECT/CT require three phase power to be supplied. Structural assessment is required to ensure the floor will be rigid with no vibration and will accommodate the weight of the SPECT-CT unit. Consideration needs to be given to the clearances of equipment and replacement part delivery routes. Where a service is may be performing minor interventional procedures, a Type A handwashing basin will be required instead of Type B.

## Items
Standard Components - Items in room
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
SPECT-CT | SPECT-CT Imaging Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
SPECT-CT | SPECT-CT Imaging Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 11
SPECT-CT | SPECT-CT Imaging Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-007 | WALL FINISH: splashback | 2
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
SPECT-CT | SPECT-CT Imaging Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 7
SPECT-CT | SPECT-CT Imaging Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
SPECT-CT | SPECT-CT Imaging Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | optional, patient observation by camera may be considered, glazing to be AS1428 compliant safety glass
SPECT-CT | SPECT-CT Imaging Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-117 | CUPBOARD: tall, double door, adj shelves, lockable | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-045 | IV STAND: mobile | 1 | optional
SPECT-CT | SPECT-CT Imaging Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMBE-042 | SLIDE: patient | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-139 | TROLLEY: procedure | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-002 | ADAPTER: suction | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-262 | DEVICE: nuclear medicine, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | ITCL-012 | CAMERA: observation | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-264 | CART: collimator, SPECT/CT | 4 | required quantity dependent on patient mix
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-265 | BOARD: distribution, mains, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-263 | BOX: line connection, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-100 | INJECTOR: contrast, CT | 1 | optional, floor or ceiling mounted to be confirmed
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-272 | BOARD: disconnect, mains, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-273 | PHANTOM: SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | MMMI-377 | WARMER: contrast | 1
SPECT-CT | SPECT-CT Imaging Room | FFE | FIHR-078 | BRACKET: for observation camera, wall mounted | 2
SPECT-CT | SPECT-CT Imaging Room | FFE | MMGE-179 | MACHINE: electrocardiography (ECG), c/w trolley and printer shelf | 1 | procured with SPECT/CT system vendor
SPECT-CT | SPECT-CT Imaging Room | FFE | MMHA-032 | MONITOR: patient, cardiac, SPECT/CT | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | on wall near SPECT/CT unit
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-006 | RCD: residual current device | 1 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | optional, on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1 | auto on when modality is in use
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | optional, on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
SPECT-CT | SPECT-CT Imaging Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel + indicator button & light
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-109 | GPO: single, wall mounted | 2 | clock
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 5 | to walls
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-123.1 | GPO: emergency power, single, wall mounted (V) | 2
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ELGT-007 | UPS: uninterruptable power supply unit | 1 | optional, dependent on vendor requirements
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
SPECT-CT | SPECT-CT Imaging Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | patient services
SPECT-CT | SPECT-CT Imaging Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors extent as specified by radiation consultant
SPECT-CT | SPECT-CT Imaging Room | SE | ELBO-012 | SWITCH: light, dimmer | 1 | number determined by number of doors
SPECT-CT | SPECT-CT Imaging Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
SPECT-CT | SPECT-CT Imaging Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4
SPECT-CT | SPECT-CT Imaging Room | SE | ELPR-007 | UPS: status indicator | 1
SPECT-CT | SPECT-CT Imaging Room | SE | LIFX-038 | LIGHT: backlit feature panels | 2