## Room Template Data
Room Data Sheet - File for Import
2026-10-29
No. of Standard Components = 1
Code | Template Name | Standard Component Set | Standard Area | Modeled Ceiling Height | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Panorama - URL for Project Use | Briefing - Hours of Operation | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experience system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statutory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magnetic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: low | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium | Performance Requirements - Acoustics - NOISE SENSITIVITY: high | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | 10 | 2700 |  | 1 | 29.10.2026 |  | 8am to 6pm, Mon - Fri (service may operate with extended hours, on weekends and/or for emergency and urgent care situations) | 1-2 staff intermittently | The Packing and Sterilizing Room - Satellite provides the space and facilities for the inspection, assembly, packing, labelling and sterilization of reusable medical devices (RMDs) within a clinical unit. | Planning Context ~ Room must be configured to meet the requirements of AS/NZ 5369, providing clear circulation and support the unidirectional flow of reusable medical devices (RMDs) from receipt following disinfection, through inspection, assembly, packing and labelling, sterilization, cooling and transfer to the Sterile Store. ~ The Packing and Sterilizing Room may be physically separated from, or open to, the Decontamination Room. The configuration should be determined at project level based on the reprocessing workflow, equipment configuration and applicable requirements. Where the rooms are combined, dirty and clean zones and associated PPE requirements must be clearly identified. Where the rooms are enclosed and separated and the washer disinfector is not pass-through, an openable transfer hatch/window should be provided to enable racks to be transferred to bench space within the Packing and Sterilizing Room. ~ Consider staff circulation between the Decontamination and Sterilizing Rooms, including requirements for hand hygiene and changing PPE before moving from dirty to clean activities.~ Provision of automated, hands-free door operation for delivery and movement of instruments may be considered at project level where this supports the selected workflow and local IPC requirements. ~ Direct access or immediate adjacency to the Sterile Store is recommended to support transfer of sterilised RMDs following cooling. Furniture, Fixtures and Equipment (FF&E) Selection ~ Requirements for drying manually cleaned instruments are to be determined at project level based on equipment selection, and service requirements. ~ Confirm requirements for a handpiece oiling unit where reusable medical devices require oiling or lubrication following disinfection, based on the RMDs to be processed and operational workflow. ~ Confirm the method of printing labels, including whether printing is integral to the sterilizer or provided through a standalone printer, based on the selected equipment, instrument tracking process and operational requirements. Information and Communication Technology (ICT) ~ Confirm any in room requirements for instrument tracking, including location, process and configuration, in line with AS/NZ 5369, local policies and service requirements. Waste Management ~ Size, type and quantity of waste bins provided within the room will be dependent on services requirements, operational models for waste management (including separation of waste streams and frequency of waste removal) and local infection prevention and control (IPC) policies. Security ~ Where there is potential for unauthorised access to staff-only areas, consider access control and staff safety requirements, including swipe card access and/or a fixed duress button. Engineering Services Engineering Services~ Confirm requirements for RO water or water of suitable quality to decontamination sinks and washer disinfector in line with AS/NZ 5369, local policies and service requirements. Lighting ~ Confirm general and task lighting requirements based on the activities undertaken within the room, including inspection, assembly, packing and labelling of RMDs. | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE

## Items In Room Template
Room Data Sheet - File for Import
2026-10-29
Code | Template Name | Standard Component Set | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | to room side
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | to corridor side, extent dependent on movement of beds and mobile equipment in adjacent area
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | D+W | AFDPR-056.01 | DOOR FRAME PROTECTION: full wrap, to 900H | 1 | 3 | Internal - Door 1 | optional, provision and extent dependent on frame material/finish and movement of beds and mobile equipment in adjacent area
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FIN | CLCN-031 | CORNICE: square set | 1 | 1 | Ceiling Cornice
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FIN | CLFI-003 | CEILING FINISH: paint, sterile areas | 1 | 1 | Ceiling Finish 2
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FIN | CLFS-011 | CEILING: flush set, suspended | 1 | 1 | Ceiling Finish 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | D+W | DOHI-043.01 | DOOR: hinged, 1 leaf, 900 clear opening, solid, wide vision panel | 1 | 1 | Internal - Door 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | SER | ELGP-101 | GPO: single, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | SER | ELGP-201 | GPO: double, wall mounted | 3 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | SER | ELSW-001 | SWITCH: light | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | FIDI-241 | DISPENSER: alcohol-based hand rub, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | FIHR-161 | BRACKET: storage basket, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | FIHR-294 | BRACKET: combined single screen all-in-one computer, keyboard and mouse, articulated, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FIN | FLVY-111 | FLOOR FINISH: vinyl, seamless, non-slip | 1 | 1 | Floor Finish 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | FQGE-113 | BASKET: storage, for wall mounted panel, medium | 2 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | SER | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | ITNE-122 | COMPUTER: single display screen, all-in-one, medical grade | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | ITNE-204 | INPUT DEVICE: keyboard and mouse, wireless, medical grade, washable | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | ITNE-223 | INPUT DEVICE: bar code scanner, medical grade, washable | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | ITNE-316 | PRINTER: labels, small, desktop | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | JOBE-033 | BENCH: 600D, stainless steel lipped | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | JOBE-231 | BENCH SUPPORT: legs and frame, stainless steel | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | MMIC-031 | DEVICE: heat sealer, crimper | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | MMIC-142 | STERILIZER: dental, benchtop | 2 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | MMIC-404.01 | WORKSTATION: instrument inspection, straight, adjustable height, 900H nom, 750D x 1500W | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | MMIC-461 | LIGHT: task, magnifying glass, instrument inspection, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | MMIC-482 | DISPENSER: sheet material, roll, sterilizing wrap, wall mounted | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FFE | MMIC-496 | TROLLEY: transfer, instruments, dirty | 1 | 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | D+W | WIOP-017 | WINDOW: operable, hung, internal, single glazed, sill at 900H | 1 | 1 | Internal - Window 1 | for transfer of washed instruments from decontamination room
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FIN | WLFI-003 | WALL FINISH: paint, sterile areas | 1 | 1 | Wall Finish 1
STERI-SAT | Packing and Sterilizing Room - Satellite | Oral Health | FIN | WLFI-011.06 | WALL FINISH: vinyl, to 2100 AFFL | 1 | 1 | Wall Finish 2