## Room Template Data
Store - Sterile Stock (STSS) - Room Template Properties and Room Template Data
2026-08-05
No. of Standard Components = 1
Code | Template Name | Standard Area | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Briefing - Hours of Operation - Duration - Duration | Briefing - Hours of Operation - Duration - extended/limited hours | Briefing - Hours of Operation - Hours/days - Open time | Briefing - Hours of Operation - Hours/days - Close time | Briefing - Hours of Operation - Hours/days - Days | Briefing - Hours of Operation - Comment | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experiance system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statuatory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magentic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: sensitive | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
STSS | Store - Sterile Stock | 40 | STSS-20 | 6 | 06.05.2025 | 24 hours | 0 |  |  |  |  | Up to 3 intermittently | A room with an aseptic environment for storage and holding of sterile stock for use in the Operating Suite or other acute treatment areas. | ~ This Standard Component assumes direct access to Operating Rooms and that it is multiplied as a module to serve Operating Rooms that are planned in a row – with the end of the room to be enclosed once modules are combined. Where sterile stock storage is centralised and transported through corridors to Operating Rooms. This room is to be enclosed (size, FF&E, and services increased according to number of Operating Rooms to be served), and door arrangements are to be adjusted to suit. ~ Sterile stock is to be stored in shelving compliant with AS 5369, this may be provided as either static units, single mobile units and/or as a compactus. If a compactus is to be used, depending on the product selected and the materials it is made from, structural engineering advice may need to be sought to ensure adequate load-bearing capacity and potential deﬂection eﬀect on walls, doors, etc. ~ Storage solutions must support consumable items to be stored separately (i.e. in different shelving units) to reusable medical devices (RMDs) as per AS 5369. ~ The size, quantity and type of storage solution will vary to suit service requirements. Additional shelving units require an approximate area of 1m2 per unit (assumes 900W x 400D x 2100H unit with space in front to access). ~ Storage height to be limited to 520mm min below ceiling to meet ﬁre sprinkler system requirements within AS 2118.1. ~ A clearance of 1100mm is recommended to be provided for circulation spaces within Store – Sterile Stock to facilitate movement of trolleys and staff. ~ Conﬁguration of computer workstations (bench mounted, wall mounted on articulated brackets, and/or workstations on wheels) is to be conﬁrmed to suit service requirements and local ICT policies. ~ Provision of automated doors to access this store from the corridor may be considered to suit service requirements and operational policies. ~ Light ﬁxtures to be sealed. | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
This shape represents a table slicer. Table slicers are not supported in this version of Excel.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Items In Room Template
Standard Components - Items in Room Template
2026-08-05
No. of Standard Components = 1
Code | Template Name | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
STSS | Store - Sterile Stock | D+W | AFDPR-006.05 | DOOR PROTECTION: plate, full height | 1 | 1 | Internal - Door 1
STSS | Store - Sterile Stock | D+W | AFDPR-006.05 | DOOR PROTECTION: plate, full height | 1 | 1 | Internal - Door 2
STSS | Store - Sterile Stock | D+W | DOHI-014.03 | DOOR: hinged, 1 leaf, 950 clear opening, solid, standard vision panel, with cover | 1 | 1 | Internal - Door 2 | to operating room
STSS | Store - Sterile Stock | D+W | DOHI-114.02 | DOOR: hinged, 1 1/2 leaves, 1400 clear opening, solid, standard vision panel, with integral blind | 1 | 1 | Internal - Door 1 | to corridor, large leaf to open 170 degrees
STSS | Store - Sterile Stock | FFE | FIHR-151 | PANEL: storage baskets, wall mounted | 3 | 1 |  | type and configuration of storage solutions to suit project requirements
STSS | Store - Sterile Stock | FFE | FIHR-212 | BRACKET: computer display screen, single, articulated, wall mounted | 1 | 1
STSS | Store - Sterile Stock | FFE | FQBS-203 | STOOL: clinical areas, mobile, high | 1 | 1
STSS | Store - Sterile Stock | FFE | FQSM-211 | SHELVING UNIT: mobile, adjustable shelves, 600D x 400W x 2100H nominal | 8 | 1 |  | type and configuration of storage solutions to suit project requirements
STSS | Store - Sterile Stock | FFE | FQSN-141 | COMPACTUS: rail mounted | 1 | 1 |  | type and configuration of storage solutions to suit project requirements
STSS | Store - Sterile Stock | FFE | ITNE-071 | TELEPHONE: handset, desktop | 1 | 1 |  | optional, to bench or wall mounted
STSS | Store - Sterile Stock | FFE | ITNE-122 | COMPUTER: single display screen, all-in-one, medical grade | 1 | 1
STSS | Store - Sterile Stock | FFE | ITNE-203 | INPUT DEVICE: keyboard and mouse, medical grade, washable | 1 | 1
STSS | Store - Sterile Stock | FFE | ITNE-223 | INPUT DEVICE: bar code scanner, medical grade, washable | 2 | 1 |  | 1 for data for tray tracking, 1 for H tracking
STSS | Store - Sterile Stock | FFE | JOBE-033 | BENCH: 600D, stainless steel lipped | 1 | 1
STSS | Store - Sterile Stock | FFE | JOBE-211 | BENCH SUPPORT: bracket | 2 | 1
STSS | Store - Sterile Stock | FFE | MMTR-104 | TROLLEY: set-up, 900W x 490D x 900H nom | 1 | 1
STSS | Store - Sterile Stock | FIN | CLCN-031 | CORNICE: square set | 1 | 1 | Ceiling Cornice
STSS | Store - Sterile Stock | FIN | CLFI-003 | CEILING FINISH: paint, sterile areas | 1 | 1 | Ceiling Finish 1
STSS | Store - Sterile Stock | FIN | CLFS-011 | CEILING: flush set, suspended | 1 | 1 | Ceiling Finish 1
STSS | Store - Sterile Stock | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Finish 1
STSS | Store - Sterile Stock | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
STSS | Store - Sterile Stock | FIN | WLFI-003 | WALL FINISH: paint, sterile areas | 1 | 1 | Wall Finish 1
STSS | Store - Sterile Stock | FIN | WLFI-011.06 | WALL FINISH: vinyl, to 2100 AFFL | 1 | 1 | Wall Finish 1
STSS | Store - Sterile Stock | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 1 | 1 |  | to computer workstation
STSS | Store - Sterile Stock | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 1 | 1 |  | to WOW workstation zone
STSS | Store - Sterile Stock | SER | ELSW-001 | SWITCH: light | 3 | 1 |  | to general room lighting, controlled from 3 entry points
STSS | Store - Sterile Stock | SER | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | 1
STSS | Store - Sterile Stock | SER | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | 1
STSS | Store - Sterile Stock | SER | LIFX-212 | LIGHT: illuminated sign, laser in use | 1 | 1 |  | controlled from operating room
STSS | Store - Sterile Stock | SER | LIFX-213 | LIGHT: illuminated sign, x-ray in use | 1 | 1 |  | controlled from operating room
This shape represents a table slicer. Table slicers are not supported in this version of Excel.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## SC Set for Publish (hidden)
Code
Operating | ANAES
Operating | B-SCRUB
Operating | CLNUP
Operating | ENPR
Operating | EXIT
Operating | OR-GEN
Operating | STSS