## Room Template Data
Triage - Emergency, Assessment Room (TRIAGE-RM) - Room Template Properties and Room Template Data
2025-08-18
No. of Standard Components = 1
Code | Template Name | Standard Area | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Briefing - Hours of Operation - Duration - Duration | Briefing - Hours of Operation - Duration - extended/limited hours | Briefing - Hours of Operation - Hours/days - Open time | Briefing - Hours of Operation - Hours/days - Close time | Briefing - Hours of Operation - Hours/days - Days | Briefing - Hours of Operation - Comment | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experiance system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statuatory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magentic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: sensitive | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
TRIAGE-RM | Triage - Emergency, Assessment Room | 13 | TRIAGE-2 | 2 | 18.08.2025 | 24 hours | FALSE |  |  |  |  | 1 patient; 1 visitor; 1-2 staff | The Triage Assessment Room - Emergency provides the space and facilities for initial assessment, physical examination and commencement of first aid for patients presenting to the Emergency Department. This room will be located adjacent or in close proximity to the Reception - Emergency. | ~ Clear line of sight to the waiting room is recommended for observation when the triage nurse is not seeing a patient. The extent of glazing is to be determined at project level to suit the overall department layout and requirements for waiting room observation. Provision of integral blinds or extension of the privacy curtain track over the internal window is to be determined to suit requirements for observation and patient privacy. ~ Triage - Emergency, Assessment Rooms should be located in close proximity to the Reception - Emergency. ~ Provision of a window for visual access from Reception - Emergency (when these rooms are collocated) may be considered to ensure staff are not isolated in the front of house area. ~ A handwash basin is shown within this Standard Component but may instead be provided nearby and shared between multiple Triage Assessment Rooms. Removal of the handwash basin is subject to consultation with Emergency Department and Infection Prevention and Control (IPC) staff. Refer to Part D: Infection Prevention and Control for more information. ~ Triage journey board/bed flow board should be located nearby to Triage Assessment Rooms. ~ Ready access to weighing facilities, including baby weighing, is recommended. ~ Ready access to medications, including refrigerated medications, is required. Depending on model of care, access to a separate refrigerator for icepacks, oral rehydration solutions, etc. may also be required to support waiting room medicine. ~ A wall clock should be visible from all clinical areas. The location and type of clocks provided, including any required connection to a master clock system, is to be determined at a project level. ~ Electronic medical record (EMR) input may be performed on a wall mounted computer workstation, a desk/bench with computer, or a wirelessly connected workstation on wheels (WOW). Final configuration is dependent on clinical service requirements, operational models for medical records and local ICT policies. ~ The location of the patient monitor is to be determined to suit clinical workflows and equipment functionality. Appropriate in-wall support is required for mounting the monitor bracket. ~ The content and configuration of the medical services panel (MSP), including medical gases, power and data, will be dependent on clinical service requirements and is to be confirmed at project level. ~ The weight capacity of recliners selected must be considered to suit a wide range of patient requirements. Recliner must have the ability to lay flat. ~ The clear area required for a patient trolley and patient transfer aids such as mobile hoists, patient slides and hover mats, and the staff required to assist/complete the patient transfer, is to be considered. Work health and Safety (WHS) policies and patient transfer devices that are intended to be used in the patient bay should be confirmed at project level. ~ Mobile duress coverage is to be assessed and planned at a department level and coverage of all patient bays is to suit local security and WHS policies and operational service requirements. ~ The type, extent and configuration of storage within the room for consumables will be dependent on clinical service requirements, operational policies for stock centralisation/decanting and IPC policies. ~ Size, type and quantity of waste bins will be dependent on clinical service requirements, operational models for waste management and local infection prevention and control (IPC) policies. Bins may be located in an area shared between multiple triage rooms. ~ Provision of emergency/standby power or uninterruptible power supply (UPS) is to be confirmed to suit site and service requirements. Confirmation to be based on risk assessment considering the impact of a power outage on patient care/safety. ~ Provision of cleaner’s power outlets is to be rationalised across the department and spaced in accordance with AS/NZS 3003. | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Items In Room Template
Standard Components - Items in Room Template
2025-08-18
No. of Standard Components = 1
Code | Template Name | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
TRIAGE-RM | Triage - Emergency, Assessment Room | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | to both sides
TRIAGE-RM | Triage - Emergency, Assessment Room | D+W | AFDPR-056.01 | DOOR FRAME PROTECTION: full wrap, to 900H | 1 | 3 | Internal - Door 1 | optional; provision dependent on frame material and finish
TRIAGE-RM | Triage - Emergency, Assessment Room | D+W | DOHI-043.01 | DOOR: hinged, 1 leaf, 900 clear opening, solid, wide vision panel | 1 | 1 | Internal - Door 1 | entry from waiting room; extent of glazing is to be determined at project level to suit the overall department layout and requirements for waiting room observation
TRIAGE-RM | Triage - Emergency, Assessment Room | D+W | WIFX-007.01 | WINDOW: fixed, internal, single glazed, sill at 900H | 1 | 1 | Internal - Window 1 | extent of glazing is to be determined at project level to suit the overall department layout and requirements for waiting room observation; provision of integral blinds or provision of privacy curtain track is to be determined to suit requirements for observation and patient privacy
TRIAGE-RM | Triage - Emergency, Assessment Room | D+W | ITSE-261 | READER: security, access control, proximity card, wall mounted | 1 | 1 | Internal - Door 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIHR-408 | BRACKET: patient monitor, articulated, wall mounted | 1 | 1 |  | with cable and accessory management items as required
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIHR-297 | BRACKET: combined single screen all-in-one computer, keyboard and mouse, height adjustable, wall mounted | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | ITNE-122 | COMPUTER: single display screen, all-in-one, medical grade | 1 | 1 |  | write-up function may be provided by a wall mounted computer workstation, a desk/bench with computer, or a wirelessly connected workstation on wheels (WOW). Final configuration is dependent on clinical service requirements, operational models for medical records and local ICT policies
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIDI-106 | DISPENSER: PPE, disposable gloves, wall mounted | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIDI-101 | DISPENSER: PPE, disposable masks, wall mounted | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIDI-231 | DISPENSER: paper towel, wall mounted | 1 | 1 |  | to basin
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIDI-241 | DISPENSER: alcohol-based hand rub, wall mounted | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIDI-256 | DISPENSER: soap, wall mounted | 1 | 1 |  | to basin
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIDI-261 | DISPENSER: soap, antiseptic, wall mounted | 1 | 1 |  | to basin
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIHR-101 | BRACKET: suction bottle, wall mounted | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIHR-110 | BRACKET: sharps bin, wall mounted | 1 | 1 |  | sharps bin may alternatively be provided on a dedicated cart; solution to suit clinical service requirements and local IPC policies
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQBS-202 | STOOL: clinical areas, mobile | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQBS-101 | CHAIR: visitor, clinical areas | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQGE-101 | CURTAIN: privacy screen | 2 | 1 |  | 1 to entry door from waiting room; 1 to treatment areas within the unit
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQWS-003 | BIN: sharps, clinical, medium | 1 | 1 |  | size to suit clinical requirements and operational models for frequency of disposal; may alternatively be provided on a dedicated cart; solution to suit clinical service requirements and local IPC policies
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQWS-051 | BIN: general waste, 20L | 1 | 1 |  | size, type, quantity and distribution of waste bins will be dependent on clinical service requirements, operational models for waste management and local IPC policies
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQWS-231 | BIN: clinical waste, 20L | 1 | 1 |  | size, type, quantity and distribution of waste bins will be dependent on clinical service requirements, operational models for waste management and local IPC policies
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | MMBE-091 | CHAIR: patient, recliner, clinical, electric | 1 | 1 |  | must have ability to lay flat
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | MMGE-191 | CANNISTER: suction bottle | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | MMGE-238 | DIAGNOSTIC SET: otoscope, ophthalmoscope and thermometer, wall mounted | 1 | 1 |  | configuration of equipment to be confirmed to suit clinical services requirements
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | MMTR-062 | TROLLEY: dressing, 600W x 490D x 900H nom | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | MMPM-011 | MONITOR: patient, medium acuity | 1 | 1 |  | configuration of monitoring to be confirmed to suit clinical services requirements
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQGE-101 | CURTAIN: privacy screen | 1 | 3 |  | to window facing waiting room; provision of integral blinds or provision of privacy curtain track is to be determined to suit requirements for observation and patient privacy
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIDI-181 | DISPENSER: cleaning wipes, wall mounted | 1 | 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | ITNE-203 | INPUT DEVICE: keyboard and mouse, medical grade, washable | 1 | 1 |  | to all-in-one computer and display screen
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIRT-051 | TRACK: curtain, privacy | 2 | 1 |  | 1 to entry door from waiting room; 1 to treatment areas within the unit
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIRT-051 | TRACK: curtain, privacy | 1 | 1 |  | to window facing waiting room; provision of integral blinds or provision of privacy curtain track is to be determined to suit requirements for observation and patient privacy
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FIHR-161 | BRACKET: storage basket, wall mounted | 1 | 1 |  | type and extent of storage for consumables to suit clinical service requirements, operational models for stock centralisation/decanting and local IPC policies
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | FQGE-113 | BASKET: storage, for wall mounted panel, medium | 1 | 1 |  | type and extent of storage for consumables to suit clinical service requirements, operational models for stock centralisation/decanting and local IPC policies
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | HYBA-121 | BASIN: type B, handwashing | 1 | 1 |  | location of thermostatic mixing valve to be determined based on ability to share with nearby basins
TRIAGE-RM | Triage - Emergency, Assessment Room | FFE | HYTP-067 | TAPWARE: basin, tap set, wall mounted, lever handles | 1 | 1 |  | selection of basin and tapware should be coordinated to ensure compatibility
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Finish 1 | floor finish must be minimum slip rating R10/Pendulum P3 or agreed equivalent to extend under basin or an inset section of non-slip vinyl is required
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | WLFI-002 | WALL FINISH: paint, clinical areas | 1 | 1 | Wall Finish 1
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | CLCN-011 | CORNICE: shadow angle, prefinished | 1 | 1 | Ceiling Cornice | to be square set if flush set ceiling is provided
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | WLFI-011.04 | WALL FINISH: vinyl, to 1500 AFFL | 1 | 1 | Wall Finish 2
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | WLFI-101 | SPLASHBACK: vinyl | 1 | 1 | Wall Finish 3 | to basin
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | WLPR-006.06 | WALL PROTECTION: corner guards, to 2100 AFFL | 1 | 1 | Wall Finish 4
TRIAGE-RM | Triage - Emergency, Assessment Room | FIN | CLTI-024.02 | CEILING: drop-in tiles, anti-microbial, acoustic, prefinished, 600 x 1200 | 1 | 1 | Ceiling Finish 1 | flush set ceiling is also acceptable
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MMSP-051 | MEDICAL SERVICES PANEL: wall mounted | 1 | 1 | Panel 1 | to patient's right; content and configuration to be confirmed based on clinical service requirements and medical equipment selection
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MGAS-062 | OUTLET: suction, on services panel | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MGAS-022 | OUTLET: medical air, on services panel | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MGFP-021 | FLOWMETER: medical air | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MGFP-041 | FLOWMETER: oxygen | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MGFP-061 | ADAPTER: suction | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | MMSP-051 | MEDICAL SERVICES PANEL: wall mounted | 1 | 1 | Panel 2 | to patient's left; content and configuration to be confirmed based on clinical service requirements and medical equipment selection
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | ITCL-133 | BUTTON: nurse call, patient to staff call, with cancel, with handset connection point, on services panel | 1 | 1 | Panel 2 | may be combined on a single face plate with staff assist call button; placement to ensure quick access by staff
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | ELGP-146 | GPO: single, emergency power, on services panel | 2 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | ELGP-116 | GPO: single, on services panel | 2 | 1 | Panel 2
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | ELSW-001 | SWITCH: light | 1 | 1 | Panel 1 | optional; required if recessed examination light provided
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | ELPR-083 | RCD: residual current device, emergency power, on services panel | 1 | 1 | Panel 1
TRIAGE-RM | Triage - Emergency, Assessment Room | MSP | ELPR-083 | RCD: residual current device, emergency power, on services panel | 1 | 1 | Panel 2
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ELGP-105 | GPO: single, wall mounted, cleaner | 1 | 1 |  | as required for distribution across the unit in accordance with AS/NZS 3003
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ELGP-131 | GPO: single, emergency power, wall mounted | 1 | 1 |  | to all-in-one computer and display screen
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | 1 |  | to wall mounted telephone
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ITCL-191 | BUTTON: nurse call, emergency, with cancel, wall mounted | 1 | 1 |  | placement to support quick access by staff
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ITCL-251 | HANDSET: nurse call, patient to staff call | 1 | 1 |  | with holder and cable management items as required
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ELGP-101 | GPO: single, wall mounted | 1 | 1 |  | to diagnostic set
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | 1 |  | to patient monitor
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | 1 |  | to all-in-one computer and display screen
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ITCL-111 | LIGHT: nurse call indicator, ceiling mounted | 1 | 1 |  | outside entry from treatment areas; placement to ensure visibility along paths of travel
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ITCL-181 | BUTTON: nurse call, staff assist, with cancel, wall mounted | 1 | 1 |  | may be combined on a single face plate with patient to staff call button; placement to support quick access by staff
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ELGP-131 | GPO: single, emergency power, wall mounted | 1 | 1 |  | to patient monitor
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ELGP-131 | GPO: single, emergency power, wall mounted | 1 | 1 |  | low height, for patient recliner
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | LIFX-111 | LIGHT: examination, ceiling recessed | 1 | 1 |  | optional; provision dependent on clinical service requirements; light switch located on MSP; light may alternatively be provided as an articulated examination light with fixture mounted controls
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ELPR-071 | RCD: residual current device, wall mounted | 1 | 1 |  | quantity and type to be confirmed to suit final power provisions
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | ELPR-081 | RCD: residual current device, emergency power, wall mounted | 1 | 1 |  | quantity and type to be confirmed to suit final power provisions
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | HYTP-451 | DIRECT CONNECTION: water, cold | 1 | 1 |  | to basin
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | HYTP-453 | DIRECT CONNECTION: water, warm | 1 | 1 |  | to basin
TRIAGE-RM | Triage - Emergency, Assessment Room | SER | HYDR-201.01 | DIRECT CONNECTION: wastewater, general | 1 | 1 |  | to basin
This shape represents a table slicer. Table slicers are not supported in this version of Excel. If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## SC Set for Publish (hidden)
Code
Emergency | CONS-ENT-OP
Emergency | PTB-E-ACU
Emergency | PTB-E-ACU-E
Emergency | PTB-E-AMB
Emergency | PTB-E-AMT
Emergency | PTB-E-NAC
Emergency | PTB-E-RES
Emergency | RECP-E
Emergency | SSTN-E
Emergency | TRIAGE-CO
Emergency | TRIAGE-RM