## Room Data
Standard Components - Room Data Sheet
Date of issue: 26.04.2021
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | 28 | up to 24 Hours | 2 patients, 2 -3 staff |  | Sound attenuation level - high; acoustic privacy is required | The Treatment Room - Hand Therapy/Splinting provides a controlled environment, privacy and facilities for carrying out treatment and rehabilitation performed by an occupational therapist or physiotherapist with patients that suffer from conditions affecting the hands and upper limbs, such as wounds, scars, burns, fractures or amputations. Additional Considerations: A second door may be required for staff safety and security; the location of the second door will be dependent on the availability an appropriate area to egress. Depending on operational model, the hand therapy treatment room may be an open area adjoining a space with other allied health treatment bays. If open, floor boxes may be required for power to height adjustable tables. Printing facilities to be shared within the overall department. Refrigeration requirements for some materials in warmer climates may need to be considered.

## Items
Standard Components - Items in room
Date of issue: 26.04.2021
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; pendulum P3 or agreed equivalent to extend under basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 8
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | WLWA-007 | WALL FINISH: splashback | 2 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to bench
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | optional, second door for egress
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FA | CLCN-009 | CORNICE: wall trim fixing | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQBS-030 | CHAIR: visitor, patient | 6
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-116 | CUPBOARD: tall, double door, adj shelves | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-140 | CUPBOARD: underbench, single door, adj shelves | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQGE-010 | CURTAIN: privacy screen | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-104 | CUPBOARD: drawer unit, 3 equal | 1 | for thermaplastic storage, width to be confirmed to suit size of material
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-152 | CUPBOARD: wall mount, shelf unit, adj shelves | 4
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYBA-010 | BASIN: handwash, Type B | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYBA-069 | SINK: single, with drainer, 930mmL | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMAH-012 | DEVICE: continuous passive motion (CPM), hand and fingers | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMAH-021 | PAN: splint, molding | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | LIFX-026 | LIGHT: examination, wall mounted | 2 | to therapy tables, ceiling mounted is also acceptable
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYDR-008 | TRAP: plaster | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | HYTP-040 | TAPWARE: mixer, sink, accessible | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQGE-043 | MACHINE: sewing | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMLA-014 | DEVICE: wax bath | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | MMGE-150 | VIEWER: X-ray, 1 screen | 1 | optional
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQCW-010 | BIN: waste, clinical, 80L | 1 | type, size and quantity of bins to be determined based on service requirements and waste management policies
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIJO-212 | BULKHEAD: joinery | 2 | plasterboard bulkhead also acceptable
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 2 | to therapy tables, dependent on ICT strategy
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | ITNE-066 | COMPUTER: all-in-one, wall mounted | 2 | to therapy tables, dependent on ICT strategy
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQTA-084 | TABLE: adjustable height, hand therapy | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FQDW-049 | DESK: adj height, 1200W x 600D x 750H nom. | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIDI-053 | DISPENSER: velcro | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | FFE | FIRT-047 | TRACK: curtain, privacy, U shaped | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELPR-006 | RCD: residual current device | 2
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 2 x examination lights
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELGP-109 | GPO: single, wall mounted | 3 | 1 x optional to x-ray viewing box
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELGP-208 | GPO: double, wall mounted | 9 | extent of emergency power to be determined based on service requirements
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | SE | ITSE-006 | BUTTON: duress, fixed | 2