## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Single line edit) | (Single line edit) | (Single line edit) | (Single line edit)
TRMT-RD | Treatment Bay - Renal Dialysis | 9 |  | 1 patient, 1 support person, 1 staff assisting | 0800 Hours to 1800 Hours, daily |  | The Treatment Bay - Renal Dialysis is used for day treatment of patients undergoing renal dialysis treatments. Treatments for each patient may take from 3 to 6 hours, for 3 days per week. Natural light is desirable. Furniture and furnishings should provide a high level of patient comfort and amenity. A staff handwashing basin will be located in close proximity. Body Protected electrical area

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Yes/No) | (Yes/No) | (Yes/No) | (Yes/No) | (Yes/No) | (Yes/No) | (Yes/No) | (Yes/No) | (Yes/No)
TRMT-RD | Treatment Bay - Renal Dialysis | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | Fabrics | 3092 | Wall protection | 1 |  | MDF | Laminate | Bed protection wall panel | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | Fabrics | 5060 | Ceiling - Acoustic | 1 | Minimum height 2700mm, flush set plasterboard ceiling also acceptable | Acoustic | Prefinished | Drop-in tiles, 600x1200 | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | Fabrics | 6030 | Cornice | 1 |  | Aluminium | Powdercoat | 24mm shadow angle | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | Fabrics | 7001 | Windows | 1 | Desirable |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FE | FE-15980 | Dialysis machine | 1 | if space used for dialysis |  |  |  | Yes | No | Yes | No | No | No | Yes | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FE | FE-15985 | Dialysis station | 1 | if space used for dialysis; RO reticulated water and drain |  |  |  | No | No | Yes | No | Yes | Yes | Yes | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FE | FE-54510 | Telephone - handset, standard | 1 |  |  |  |  | No | Yes | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FE | FE-5857 | Light - procedures small, ceiling | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-1250 | Blind - Venetian | 1 | to external windows, if provided |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-1465 | Bracket - suction bottle | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-150 | Air flowmeter | 1 | optional |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-16200 | Table - overbed | 1 | may use a fixed table attached to a recliner chair |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-16250 | Oxygen flowmeter | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-21950 | Suction adapter | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-2200 | Chair - recliner | 1 | or patient bed |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-22000 | Suction bottle | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-2650 | Chair - visitor, patient | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-4600 | Curtain track - bed screen | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-5000 | Curtain - bed screen | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FE | FF-5547 | Patient entertainment | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-5825 | Bracket - patient entertainment, ceiling | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | FF | FF-7402 | Trolley - equipment | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-1006 | Voice / Data outlet: double | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-1090 | TV: MATV | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-11000 | Wall panel | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-4037 | Tundish drain | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-4045 | Reverse osmosis (RO) water outlet | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-5000 | Airconditioning | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-6001 | General: colour corrected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-6050 | Special: reading | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-6200 | Light switches | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-7000 | Oxygen (O2) | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-7020 | Medical Air (MA) | 1 | optional |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-7040 | Suction | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-8000 | Patient/ Staff call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-8005 | Patient call handset | 1 | call button on the handset |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-8010 | Staff/ Nurse assist call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-8080 | Emergency call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-9001 | GPO: Double | 3 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-9010 | GPO: Emergency power, single (V) | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-9015 | GPO: Cleaner | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-9040 | Body protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
TRMT-RD | Treatment Bay - Renal Dialysis | SE | S-9060 | GPO: Special: for equipment | 2 | TV, exam light |  |  |  | No | No | No | No | No | No | No | No | No