## Room Template Data
Ultrasound Room (ULTR) - Room Template Properties and Room Template Data
2026-08-05
No. of Standard Components = 1
Code | Template Name | Standard Area | Previous Code | RDS Rev: Name | RDS Rev Date: Name | Briefing - Hours of Operation - Duration - Duration | Briefing - Hours of Operation - Duration - extended/limited hours | Briefing - Hours of Operation - Hours/days - Open time | Briefing - Hours of Operation - Hours/days - Close time | Briefing - Hours of Operation - Hours/days - Days | Briefing - Hours of Operation - Comment | Briefing - Occupancy | Briefing - Description | Briefing - Additional Considerations | Performance Requirements - Electrical - PROTECTION: body protected | Performance Requirements - Electrical - PROTECTION: cardiac protected | Performance Requirements - Lighting - LIGHTING: general | Performance Requirements - Lighting - LIGHTING: colour corrected | Performance Requirements - Lighting - LIGHTING: dimmable | Performance Requirements - Lighting - LIGHTING: indirect | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: buttons / handset | Performance Requirements - Nurse Call and Duress - NURSE CALL SYSTEM: annunciator | Performance Requirements - Nurse Call and Duress - DURESS: fixed | Performance Requirements - Nurse Call and Duress - DURESS: wireless coverage | Performance Requirements - Security - ACCESS CONTROL: to door | Performance Requirements - Security - ACCESS CONTROL: to item / joinery | Performance Requirements - Security - INTERCOM: service communications | Performance Requirements - Security - INTERCOM: security and access control | Performance Requirements - Security - CCTV: camera coverage within room | Performance Requirements - Security - INTRUSION DETECTION: door monitoring | Performance Requirements - Security - INTRUSION DETECTION: spatial monitoring | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: patient entertainment system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: visitor experiance system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: virtual collaboration system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: clinical support system | Performance Requirements - ICT and Audio Visual - AUDIO VISUAL: digital operating room system | Performance Requirements - Accessibility - AUDIO: hearing augmentation | Performance Requirements - Accessibility - VISUAL: luminance contrast | Performance Requirements - Accessibility - SIGNAGE: accessible, statuatory | Performance Requirements - HVAC - AIRCONDITIONING: general | Performance Requirements - HVAC - AIRCONDITIONING: HEPA filtered | Performance Requirements - HVAC - AIRCONDITIONING: positive pressure | Performance Requirements - HVAC - AIRCONDITIONING: negative pressure | Performance Requirements - HVAC - VENTILATION: exhaust | Performance Requirements - HVAC - VENTILATION: supply | Performance Requirements - HVAC - VENTILATION: natural | Performance Requirements - Medical Gas - MEDICAL GAS: general anaesthesia | Performance Requirements - Medical Gas - MEDICAL GAS: special care | Performance Requirements - Medical Gas - MEDICAL GAS: special care, neonatal ventilation | Performance Requirements - Medical Gas - MEDICAL GAS: birthing | Performance Requirements - Hydraulic - WATER: drinking | Performance Requirements - Hydraulic - WATER: specialty | Performance Requirements - Hydraulic - DRAINAGE: sanitary | Performance Requirements - Hydraulic - DRAINAGE: specialty | Performance Requirements - Fire - DETECTION: smoke | Performance Requirements - Fire - DETECTION: heat | Performance Requirements - Shielding - SHIELDING: ionising radiation | Performance Requirements - Shielding - SHIELDING: magentic and radio frequency | Performance Requirements - Acoustics - SPEECH PRIVACY: not private | Performance Requirements - Acoustics - SPEECH PRIVACY: moderate | Performance Requirements - Acoustics - SPEECH PRIVACY: private | Performance Requirements - Acoustics - SPEECH PRIVACY: confidential | Performance Requirements - Acoustics - NOISE SENSITIVITY: not sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: medium sensitive | Performance Requirements - Acoustics - NOISE SENSITIVITY: sensitive | Performance Requirements - Acoustics - NOISE GENERATION: low | Performance Requirements - Acoustics - NOISE GENERATION: moderate | Performance Requirements - Acoustics - NOISE GENERATION: high | Performance Requirements - Acoustics - NOISE GENERATION: very high
ULTR | Ultrasound Room | 16 |  | 6 | 15.05.2025 | Up to 24 hours | 0 |  |  |  |  | 1 patient; 1 visitor/carer; 1-2 staff | The Ultrasound Room provides the space and amenities for sonographers to perform diagnostic ultrasound imaging on patients. | ~ The room will require sufficient clearance for the ultrasound machine and sonographer to be positioned to either side of the examination trolley and may need to occasionally accommodate a patient on a bed to either be transferred to the examination trolley or have the ultrasound performed while they are in the bed. ~ Close access to a patient toilet and change room is required, with direct access from the room being preferred to support patients that have changed into gowns and require ultrasounds to be performed with a full bladder and then with an empty bladder. ~ A combination of dimmable general room lighting and dimmable indirect lighting (wall and/ or ceiling fixtures) is preferred to obtain appropriate lighting levels and avoid glare on the screen of the ultrasound machine. ~ Provision of height adjustable workstations or fixed joinery is to be in line with service requirements and local WHS policies. ~ Hours of operation depend on the service in which this room is located. Ambulatory medical imaging services may only operate during business hours or extended business hours, some small hospital services may have an on-call, after-hours service. Large units are likely to provide a 24-hour service, especially where a satellite service is provided with an emergency department. ~ Requirement for fixed and/or mobile duress to be confirmed to suit service requirements. ~ An operational model for responding to calls from staff/patients will be needed to support the inclusion of a nurse call system and is to be determined based on service requirements. ~ Provision of an intercom for staff communication within the department is dependent on service requirements. This function may be provided as part of the telephone system. ~ The final services and configuration of the Medical Services Panel (MSP) - medical gases, nurse call, power and data - will be dependent on service requirements. ~ The need for and extent of emergency power must be reviewed to suit site specific requirements and confirmed based on risk assessment considering the impact of a power outage on patient care/safety. | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE
This shape represents a table slicer. Table slicers are not supported in this version of Excel.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## Items In Room Template
Standard Components - Items in Room Template
2026-08-05
No. of Standard Components = 1
Code | Template Name | Item List: Name | Item Number | Name | Quantity | Priority | Category: Name | Comment
ULTR | Ultrasound Room | D+W | AFDPR-006.01 | DOOR PROTECTION: plate, to 900H | 1 | 1 | Internal - Door 1 | to both sides of door
ULTR | Ultrasound Room | D+W | DOHI-104.01 | DOOR: hinged, 1 1/2 leaves, 1400 clear opening, solid | 1 | 1 | Internal - Door 1
ULTR | Ultrasound Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | 1
ULTR | Ultrasound Room | FFE | FIDI-106 | DISPENSER: PPE, disposable gloves, wall mounted | 1 | 1
ULTR | Ultrasound Room | FFE | FIDI-231 | DISPENSER: paper towel, wall mounted | 1 | 1 |  | to basin
ULTR | Ultrasound Room | FFE | FIDI-241 | DISPENSER: alcohol-based hand rub, wall mounted | 1 | 1
ULTR | Ultrasound Room | FFE | FIDI-256 | DISPENSER: soap, wall mounted | 1 | 1 |  | to basin
ULTR | Ultrasound Room | FFE | FIDI-261 | DISPENSER: soap, antiseptic, wall mounted | 1 | 1 |  | to basin
ULTR | Ultrasound Room | FFE | FIHR-026 | HOOK: coat | 2 | 1 |  | optional - for patient property, 1 mounted at 1200H if provided
ULTR | Ultrasound Room | FFE | FIHR-101 | BRACKET: suction bottle, wall mounted | 1 | 1
ULTR | Ultrasound Room | FFE | FIHR-333 | BRACKET: television/AV system display screen, single, fixed, ceiling mounted | 1 | 1 |  | to slave display screen
ULTR | Ultrasound Room | FFE | FIHR-481 | BRACKET: sling, arm support, ceiling mounted | 1 | 1 |  | optional
ULTR | Ultrasound Room | FFE | FIRT-051 | TRACK: curtain, privacy | 1 | 1 |  | to entry door
ULTR | Ultrasound Room | FFE | FQBS-101 | CHAIR: visitor, clinical areas | 1 | 1
ULTR | Ultrasound Room | FFE | FQBS-171 | CHAIR: office, ergonomic | 1 | 1
ULTR | Ultrasound Room | FFE | FQBS-202 | STOOL: clinical areas, mobile | 1 | 1
ULTR | Ultrasound Room | FFE | FQCL-211 | LINEN CARRIER: dirty, single skip | 1 | 1
ULTR | Ultrasound Room | FFE | FQGE-101 | CURTAIN: privacy screen | 1 | 1 |  | to entry door
ULTR | Ultrasound Room | FFE | FQGE-121 | CLOCK: analogue, battery operated | 1 | 1
ULTR | Ultrasound Room | FFE | FQGE-161 | STOOL: step | 1 | 1 |  | optional - for patient access to examination trolley, provision dependent on adjustable height range of trolley
ULTR | Ultrasound Room | FFE | HYBA-121 | BASIN: type B, handwashing | 1 | 1
ULTR | Ultrasound Room | FFE | HYTP-067 | TAPWARE: basin, tap set, wall mounted, lever handles | 1 | 1
ULTR | Ultrasound Room | FFE | ITNE-071 | TELEPHONE: handset, desktop | 1 | 1
ULTR | Ultrasound Room | FFE | ITNE-102 | COMPUTER: dual display screens, central processing unit (CPU), keyboard and mouse, desktop | 1 | 1
ULTR | Ultrasound Room | FFE | JOCU-001 | CUPBOARD: tall, single door | 1 | 1 |  | type and extent of storage to suit service requirements, operational policies for stock centralisation/decanting and infection prevention and control policies
ULTR | Ultrasound Room | FFE | JOGE-001 | BULKHEAD: joinery | 1 | 1
ULTR | Ultrasound Room | FFE | MIUS-001 | IMAGING DEVICE: ultrasound | 1 | 1
ULTR | Ultrasound Room | FFE | MIUS-011 | RACK: ultrasound probes, wall mounted | 1 | 1
ULTR | Ultrasound Room | FFE | MIUS-021 | DEVICE: warmer, ultrasound gel, benchtop | 1 | 1 |  | may be integral to ultrasound machine
ULTR | Ultrasound Room | FFE | MMBE-156 | TROLLEY: patient, examination | 1 | 1 |  | selection to suit service requirements, may require cut out section for cardiac scans
ULTR | Ultrasound Room | FFE | MMGE-191 | CANNISTER: suction bottle | 1 | 1
ULTR | Ultrasound Room | FFE | FQDW-164.02 | WORKSTATION: straight, adjustable height, electric, 720H nom, 750D x 1200W | 1 | 1
ULTR | Ultrasound Room | FFE | FQWS-053 | BIN: general waste, 50L | 1 | 1
ULTR | Ultrasound Room | FFE | FQWS-231 | BIN: clinical waste, 20L | 1 | 1 |  | type, size and quantity of bins to be determined based on service requirements and waste management policies
ULTR | Ultrasound Room | FFE | MMTR-001.02 | TROLLEY: stainless steel, 490W x 490D x 900H nom, 2 tier | 1 | 1
ULTR | Ultrasound Room | FFE | ITCL-362 | DISPLAY SCREEN: medical equipment, slave | 1 | 1 |  | for viewing ultrasound imaging by patient, may also be wall mounted
ULTR | Ultrasound Room | FFE | MMGE-056 | POLE: patient positioning support handle, bed mounted | 1 | 1 |  | optional
ULTR | Ultrasound Room | FIN | CLCN-031 | CORNICE: square set | 1 | 1 | Ceiling Cornice | or wall trim fixing for ceiling tiles
ULTR | Ultrasound Room | FIN | CLFI-002 | CEILING FINISH: paint, clinical areas | 1 | 1 | Ceiling Finish 2
ULTR | Ultrasound Room | FIN | CLFS-011 | CEILING: flush set, suspended | 1 | 1 | Ceiling Finish 1 | ceiling tiles are also acceptable
ULTR | Ultrasound Room | FIN | FLVY-101 | FLOOR FINISH: vinyl, seamless, standard slip resistance | 1 | 1 | Floor Finish 1 | minimum slip rating R10 / pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
ULTR | Ultrasound Room | FIN | FLSK-021 | SKIRTING: vinyl, integral with floor vinyl, coved | 1 | 1 | Floor Skirting
ULTR | Ultrasound Room | FIN | WLFI-002 | WALL FINISH: paint, clinical areas | 1 | 1 | Wall Finish 1
ULTR | Ultrasound Room | FIN | WLFI-101 | SPLASHBACK: vinyl | 1 | 1 | Wall Finish 2
ULTR | Ultrasound Room | FIN | WLPR-002 | WALL PROTECTION: bedhead wall panel, low height | 1 | 1 | Wall Finish 3
ULTR | Ultrasound Room | MSP | ELGP-146 | GPO: single, emergency power, on service panel | 8 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | ELPR-073 | RCD: residual current device, on services panel | 2 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | ITCL-133 | BUTTON: patient to staff call, with cancel, with handset connection point, on services panel | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | ITCL-183 | BUTTON: staff assist, with cancel, on services panel | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | ITIN-025 | OUTLET: data, double RJ45, on services panel | 2 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | ITIN-065 | OUTLET: data, HDMI, on services panel | 2 | 1 | Panel 1 | for connecting ultrasound machine to slave display screen on ceiling, type of AV outlet to be confirmed to suit equipment requirements
ULTR | Ultrasound Room | MSP | MGAS-022 | OUTLET: medical air, on services panel | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | MGAS-062 | OUTLET: suction, on services panel | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | MGFP-021 | FLOWMETER: medical air | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | MGFP-041 | FLOWMETER: oxygen | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | MGFP-061 | ADAPTER: suction | 1 | 1 | Panel 1
ULTR | Ultrasound Room | MSP | MMSP-051 | MEDICAL SERVICES PANEL: wall mounted | 1 | 1 | Panel 1
ULTR | Ultrasound Room | SER | ELGP-105 | GPO: single, wall mounted, cleaner | 1 | 1
ULTR | Ultrasound Room | SER | ELGP-106 | GPO: single, ceiling mounted | 1 | 1 |  | to slave display screen
ULTR | Ultrasound Room | SER | ELGP-201 | GPO: double, wall mounted | 1 | 1 |  | to workstation, over desk for portable devices
ULTR | Ultrasound Room | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 1 | 1 |  | low height, for examination trolley
ULTR | Ultrasound Room | SER | ELGP-231 | GPO: double, emergency power, wall mounted | 2 | 1 |  | to workstation
ULTR | Ultrasound Room | SER | ELSW-022 | SWITCH: light, dimmer, 2-gang | 1 | 1 |  | to general room lighting and indirect lighting
ULTR | Ultrasound Room | SER | HYDR-201 | DIRECT CONNECTION: waste water | 1 | 1 |  | to basin
ULTR | Ultrasound Room | SER | HYTP-451 | DIRECT CONNECTION: water, cold | 1 | 1 |  | to basin
ULTR | Ultrasound Room | SER | HYTP-453 | DIRECT CONNECTION: water, warm | 1 | 1 |  | to basin
ULTR | Ultrasound Room | SER | ITCL-191 | BUTTON: emergency, with cancel, wall mounted | 1 | 1
ULTR | Ultrasound Room | SER | ITCL-251 | HANDSET: patient to staff call | 1 | 1
ULTR | Ultrasound Room | SER | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 1 | 1 |  | to workstation
ULTR | Ultrasound Room | SER | ITIN-061 | OUTLET: data, HDMI, ceiling mounted | 1 | 1 |  | to slave display screen
ULTR | Ultrasound Room | SER | ITSE-061 | BUTTON: security, duress, fixed, wall mounted | 1 | 1
ULTR | Ultrasound Room | SER | ITSE-166 | INTERCOM: audio, slave, wall mounted | 1 | 1 |  | optional - inclusion dependent on service requirements and ICT policies
This shape represents a table slicer. Table slicers are not supported in this version of Excel.
If the shape was modified in an earlier version of Excel, or if the workbook was saved in Excel 2007 or earlier, the slicer can't be used.

## SC Set for Publish (hidden)
Code
1BR-IC
1BR-BA