## Room Data
Standard Components - Room Data Sheet
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
UTRM | Uptake Room | 10 | 8am to 6pm, daily | 1 patient, 1 staff |  |  | The Uptake Room is used to inject patient with a radiopharmaceutical and allow them to rest until 'uptake' has occurred before transfer to the scanning room. Patients need to be relaxed in a warm, quiet and dimmed environment to promote uptake. Uptake duration is typically 45 to 60 minutes. Additional Considerations: Visual and audio observation of the patient during this time is typically provided from the Control Room of the PET-CT Imaging Room and intercoms, cameras and display screens are to be provided to support patient observation. Injections of radiopharmaceuticals can be undertaken within the room or from outside the room via an automated dose delivery system, which will be accommodated outside each uptake room with a suitable lead-lined injection port in the wall that will allow the coil to be connected to both the machine and the patient (details should be confirmed with vendor) Radiation shielding is required to the room and is to be in accordance with a radiation shielding consultants advice. This will be based on the type of radiopharmaceutical, the number of patients seen on a daily basis and proximity to other services.

## Items
Standard Components - Items in room
Date of issue: 24.06.2022
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
UTRM | Uptake Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
UTRM | Uptake Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
UTRM | Uptake Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 5
UTRM | Uptake Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
UTRM | Uptake Room | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1
UTRM | Uptake Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1 | automated, hold open
UTRM | Uptake Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
UTRM | Uptake Room | FFE | FIHR-019 | BRACKET: for suction bottle | 1
UTRM | Uptake Room | FFE | FIHR-016 | BRACKET: for sharps bin | 1
UTRM | Uptake Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
UTRM | Uptake Room | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1
UTRM | Uptake Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1
UTRM | Uptake Room | FFE | FIHR-029 | HOOK: equipment | 2
UTRM | Uptake Room | FFE | MGFP-004 | FLOWMETER: oxygen | 1
UTRM | Uptake Room | FFE | FQCW-044 | BIN: waste, sharps, wall mounted | 1
UTRM | Uptake Room | FFE | MMGE-002 | ADAPTER: suction | 1
UTRM | Uptake Room | FFE | MMGE-100 | SUCTION BOTTLE | 1
UTRM | Uptake Room | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1
UTRM | Uptake Room | FFE | ITCL-012 | CAMERA: observation | 1
UTRM | Uptake Room | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | optional
UTRM | Uptake Room | FFE | FIHR-077 | BRACKET: for observation camera, ceiling mounted | 1
UTRM | Uptake Room | SE | ITIN-011 | OUTLET: data, single RJ45, ceiling mounted | 1
UTRM | Uptake Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room
UTRM | Uptake Room | SE | ITIN-070 | OUTLET: MATV | 1 | to patient entertainment screen, may be replaced with data outlet in accordance with ICT policy
UTRM | Uptake Room | SE | ELGP-101 | GPO: single, ceiling mounted | 3
UTRM | Uptake Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
UTRM | Uptake Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1
UTRM | Uptake Room | SE | MGAS-062 | OUTLET: suction, on services panel | 1
UTRM | Uptake Room | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1
UTRM | Uptake Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
UTRM | Uptake Room | SE | ELGP-109 | GPO: single, wall mounted | 1
UTRM | Uptake Room | SE | ELGP-208 | GPO: double, wall mounted | 1
UTRM | Uptake Room | SE | ELPR-001 | BODY PROTECTED ELECTRICAL AREA | 1
UTRM | Uptake Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
UTRM | Uptake Room | SE | ELBO-012 | SWITCH: light, dimmer | 2 | inside and outside room, to be reviewed to suit operational requirements
UTRM | Uptake Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2