## Level 2
Level 2
AusHFG Room Code | Room / Space | SC / SC-D | Qty | m2 | Remarks
ENTRY / RECEPTION / CLERICAL
WAIT-20 | Waiting, 20m2 | Yes | Shared |  | 1.2 m2 per seat, 1.5m2 for w/chair. Respectively 15, 25 & 35 seats/ wheelchairs
BWD-1 | Bay - Water Dispenser | Yes
BVM-3 | Bay - Vending | Yes
PLAP-10 | Play Area - Paediatric, 10m2 | Yes |  |  | 4-5 places for children
WCAC | Toilet – Accessible, 6m2 | Yes | Shared |  | Unless available nearby
WCPU-3 | Toilet - Public, 3m2 | Yes | Shared |  | Unless available nearby
RECL-10 | Reception / Clerical, 10m2 | Yes | Shared |  | 1, 2 & 2 staff
OFF-3P | Office -Shared | Yes | Shared |  | Clerical staff. May need larger area if they also combine booking system for the facility/AHS
| Film Storage - Current |  | Shared |  | Check need according to use of PACS and film management policy
CONS | Consult Room | Yes | Shared |  | Number dependent on volumes and patient requirements
BMEQ-4 | Bay - Mobile Equipment, 4m2 | Yes | Shared |  | For mobile units depending on policy re decentralising
| Trolley / Wheelchair Park |  | Shared |  | Smaller area required if porter service provided by central Transport Unit
| Transport Staff Workbase |  | Shared |  | Not required if porter service provided by central Transport Unit.
Discounted Circulation
SUPPORT AREAS
CLRM-5 | Cleaner's Room, 5m2 | Yes | Shared |  | A separate cleaners room is scheduled for MRI if no collocated with Medical Imaging
| Laser Printing / Digitiser Room |  |  |  | This may be included as work space in the radiographers
PACS Server Room
STGN-9 | Store - General, 9m2 | Yes |  |  | Bulk stock e.g. contrast media, CT syringes, bulk items related to patient care
STSS_12 | Store – Sterile Stock, 12m2 | Yes |  |  | Bulk clinical sterile stock
STGN-9 | Store – General, 9m2 | Yes |  |  | Dependent on service/site storage management of hard film
STGN-9 | Store – General, 9m2 | Yes |  |  | For optical discs
OFF-S9 | Office – Single Person, 9m2 | Yes |  |  | Optional. PACS Operations/Management Team
DISP-8 | Disposal Room, 8m2 | Yes | Shared
Discounted Circulation
GENERAL X-RAY AND FLUOROSCOPY
GENXR | General X-Ray | Yes | 1 | 35 | Includes Control console space. Number of rooms adjusted to suit service plan
SCRN | Screening Room (Fluoroscopy) | Yes |  |  | Includes Control & area for contrast media prep. Number of rooms adjusted to suit service plan
ENS-ACC | Ensuite - Accessible, 7m2 | Yes |  |  | Dual access from fluoroscopy room and corridor
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2 | 1 cubicle per imaging room. Less required if centralised
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 1 | 4 | 1 cubicle per imaging room. Less required if centralised
WCPT | Toilet - Patient, 4m2 | Yes
WAIT-SUB | Waiting – Sub, 5m2 | Yes |  |  | Required if Change Cubicles are centralised. Separate male & female patients.
| Patient Locker Bay |  |  |  | Operational process may be for patients to take clothes/possessions with them. For centralised changed waiting.
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  |  | Recovery bays. For Level 5 & 6 refer to combined general nursing holding/recovery unit.
BLIN | Bay - Linen | Yes
BPPE | Bay - PPE | Yes | 1 | 0.5 | Storage outside of modality room for Lead Apron
| Computed Radiography (CR) Processing Workroom |  |  |  | Adjacent to control rooms.
| Digital Radiography (DR) Processing Workroom |  |  |  | For DR system – 8m2 per Imaging Room for equipment & workstations. Adjacent to control rooms
DTUR-10 | Dirty Utility, 10m2 | Yes |  |  | Disposal of contrast media, dirty linen storage. May be shared with other modalities
BRES | Bay - Resuscitation | Yes
| Discounted Circulation |  | 35%
ULTRASOUND AND MAMMOGRAPHY
WAIT-SUB | Waiting - Sub, 5m2 | Yes |  |  | Size according to activity data. Assumed patients are waiting and are changed
| Patient Locker Bay |  |  |  | Operational process may be for patients to take clothes / possessions with them. For Centralised Changed Waiting.
ULTR | Ultrasound | Yes |  |  | Number of rooms adjusted to suit service plan.
| Ultrasound - Procedures |  |  |  | For interventional procedures. Recovery from US procedures is recommended to be in the combined holding / nursing recovery area for all interventional imaging procedures.
CHPT | Change Cubicle - Patient, 2m2 | Yes |  |  | Total 1 cubicle per US room (including accessible).
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes |  |  | Total 1 cubicle per US room.
DTUR-10 | Sonographer Work Room |  |  |  | Per person. Benching. Minimum size 8m.
XRRR | X-Ray Viewing & Reporting Room |  |  |  | 1 workstation per 2 rooms. Adjust size accordingly. Altrasound and Mammography.
MAMMO | Mammography | Yes |  |  | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT | Change Cubicle - Patient, 2m2 | Yes |  |  | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes |  |  | Total 2 cubicles per mammography room.
| Ultrasound / Mammography Prep Room / Lab |  |  |  | Direct access to mammography and reporting rooms.
Discounted Circulation
ANGIOGRAPHY / DSA
ANIN | Anaesthetic Induction Room | Yes |  |  | More common alternative is anaesthetic induction in the angiography room.
SCRB-6 | Scrub-Up / Gowning, 6m2 | Yes |  |  | 2 scrub stations. May be shared between Angiography Rooms if design permits.
ANSS | Angiography Sterile Store / Set-Up Room | Yes
ANPR | Angiography Room | Yes |  |  | Number of rooms adjusted to suit service plan. This is consistent with HPU 170 Cardiac Investigations Unit Angiography Room.
ANCRT | Angiography Control / Reporting Room | Yes
| Computer Equipment / Machine Room |  |  |  | 1 per Angiography Room. Refer to specifications for equipment type.
BPPE | Bay – PPE | Yes |  |  | No dependent on number of Angiography / DSA rooms. Accommodates lead aprons.
SDTUR-S | Dirty Utility – Sub, 8m2 | Yes |  |  | May be shared with other modalities in smaller units. Not in level 5 & 6 services where instruments require cleaning prior to sterilising.
XRRR | X-Ray Viewing & Reporting Room | Yes |  |  | Adjust as necessary. Angiography
BRES | Bay - Resuscitation | Yes |  |  | May be shared with recovery areas recovery trolley.
BLIN | Bay – Linen | Yes |  |  | May access allocated linen trolleys.
STGN-8 | Store - General, 8m2 |  |  |  | If a stand-alone service only for films/ CDs/ videos – this can be shared space.
Discounted Circulation
CT SCANNING
CTPR | CT Scanning | Yes |  |  | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan.
CTCR | CT Control Room | Yes
CTCR | CT Control Room - Shared | Yes
ANPR | CT Computer Room |  |  |  | Dependent on equipment selected.
XRRR | X-Ray Viewing & Reporting Room | Yes |  |  | One workstation per two CT Scanning Rooms.
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes |  |  | 1 per CT room.
WAIT-SUB | Waiting - Sub, 5m2 | Yes |  |  | Optional. Shared for MRI and CT and provides male/female separation.
WCPT | Toilet - Patient, 4m2 | Yes
PBTR-H-6 | Patient Bay - Holding, 6m2 | Yes |  |  | 1 outside each CT scanning room
| Preparation / Chair Bay |  |  |  | Space for a chair and storage with a curtain for privacy for inserting cannula if IV contrast is required.
| Patient Locker Bay |  |  |  | Shared between MRI and CT.
CLUR-8 | Clean Utility – Sub, 8m2 | Yes |  |  | May be shared by MRI and CT and other modalities.
DTUR-8 | Dirty Utility – Sub, 8m2 | Yes |  |  | May be shared with other modalities.
SCRB-6 | Scrub-up/ Gowning, 6m2 | Yes |  |  | 2 stations. May be shared between rooms.
BHWS-B | Bay – Handwashing, Type B | Yes |  |  | Part of preparation area.
Discounted Circulation
MRI
ANIN | Anaesthetic Induction Room | Yes |  |  | Include small staff base/clean utility and 1 patient trolley/bed bay per room. Also used for recovery
BHWS-B | Bay – Handwashing, Type B | Yes |  |  | Part of preparation area.
| MRI Scanning Room |  |  |  | Size will depend on equipment selected.
| MRI Control Room |  |  |  | Must oversee and control entry into magnet room.
MRI Computer Room
MRI Reporting Room
CHPT | Change Cubicle - Patient, 2m2 | Yes
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes
BLIN | Bay - Linen | Yes |  |  | Optional if stand-alone unit.
| Patient Locker Bay | Yes
WAIT-SUB | Waiting - Sub, 5m2 | Yes
WCPT | Toilet - Patient, 4m2 | Yes
Bay – Wheelchair Park/ Trolley Park
BRES | Bay - Resuscitation | Yes |  |  | Non-ferrous construction
| Store - Dewar Tank |  |  |  | Must provide easy access into MRI room for top-up using Dewar.
Discounted Circulation
PATIENT HOLDING / RECOVERY
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  |  | At least two patient holding / recovery bays per interventional room are required (three bed spaces for CT), generally configured as a single area to maximise staff efficiencies. MRI patient holding / recovery bays however are excluded and are instead located in the MRI suite – MRI Induction / Preparation / Recovery Room. Bays should have direct observation from Staff Station. They may be used for preparation
SSTN-10 | Staff Station, 10m2 | Yes
CLUR-12 | Clean Utility / Medication Room | Yes
BLIN | Bay - Linen | Yes
BRES | Bay - Resuscitation | Yes
DTUR-S | Dirty Utility - Sub, 8m2 | Yes
STEQ-14 | Store - Equipment | Yes
Discounted Circulation
STAFF OFFICES AND REPORTING
OFF-S12 | Office - Single Person, 12m2 | Yes |  |  | Director
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 | Chief Radiologist
OFF-S9 | Office - Single Person, 9m2 | Yes |  |  | Chief Radiographer
| PACS Reporting Workstation |  |  |  | Number of spaces as per staff establishment.
| Office - Workstation, 5.5m2 |  |  |  | Number of spaces as per staff establishment.
OFF-S9 | Office - Single Person, 9m2 | Yes |  |  | NUM
| Office - Workstation, 5.5m2 |  |  |  | Number of spaces as per staff establishment for Transcription
| Office - Workstation, 5.5m2 |  |  |  | Number of spaces as per staff establishment for IT & Administration
| Discounted Circulation |  | 35%
STAFF AREAS
| Film Library / Study |  |  |  | 5 places
SSPS-8 | Store - Photocopy / Stationery, 8m2 | Yes
MEET-L-20 | Meeting Room, 20m2 | Yes
MEET-L-30 | Meeting Room, 30m2 | Yes |  |  | Access to larger meeting room.
SRM-15 | Staff Room, 15m2 | Yes |  |  | Depending on Staff Establishment. Shared staff room for small Units.
PROP-2 | Property Bay - Staff | Yes |  |  | Numbers dependent on staff establishment.
SHST | Shower - Staff, 3m2 | Yes | 1 | 3 (o) | Numbers dependent on staff establishment.
WCST | Toilet - Staff, 3m2 | Yes | 2 | 3 | No. dependent on peak staffing numbers. Easily accessible if main change rooms remote.
CHST-10 | Change Room - Staff (Male / Female), 10m2 | Yes |  |  | 1 Shower, 2 WCs plus lockers. Required for facilities with high interventional workload.
| Discounted Circulation % |  | 35%

## Level 3
Level 3
AusHFG Room Code | Room / Space | SC / SC-D | Qty | m2 | Remarks
ENTRY / RECEPTION / CLERICAL
WAIT-20 | Waiting, 20m2 | Yes | Shared |  | 1.2 m2 per seat, 1.5m2 for w/chair. Respectively 15, 25 & 35 seats/ wheelchairs
BWD-1 | Bay - Water Dispenser | Yes
BVM-3 | Bay - Vending | Yes
PLAP-10 | Play Area - Paediatric, 10m2 | Yes |  |  | 4-5 places for children
WCAC | Toilet – Accessible, 6m2 | Yes | Shared |  | Unless available nearby
WCPU-3 | Toilet - Public, 3m2 | Yes | 1 | 3 | Unless available nearby
RECL-10 | Reception / Clerical, 10m2 | Yes | Shared |  | 1, 2 & 2 staff
OFF-3P | Office -Shared | Yes | 1 | 9 | Clerical staff. May need larger area if they also combine booking system for the facility/AHS
| Film Storage - Current |  | 1 | 30 (o) | Check need according to use of PACS and film management policy
CONS | Consult Room | Yes | 1 | 12 | Number dependent on volumes and patient requirements
BMEQ-4 | Bay - Mobile Equipment, 4m2 | Yes | 1 | 4 | For mobile units depending on policy re decentralising
| Trolley / Wheelchair Park |  | 1 | 5 | Smaller area required if porter service provided by central Transport Unit
| Transport Staff Workbase |  | Shared |  | Not required if porter service provided by central Transport Unit.
| Discounted Circulation |  | 35%
SUPPORT AREAS
CLRM-5 | Cleaner's Room, 5m2 | Yes | Shared |  | A separate cleaners room is scheduled for MRI if no collocated with Medical Imaging
| Laser Printing / Digitiser Room |  |  |  | This may be included as work space in the radiographers
| PACS Server Room |  | 1 | 10
STGN-9 | Store - General, 9m2 | Yes | 1 | 9 | Bulk stock e.g. contrast media, CT syringes, bulk items related to patient care
STSS_12 | Store – Sterile Stock, 12m2 | Yes | 1 | 12 | Bulk clinical sterile stock
STGN-9 | Store – General, 9m2 | Yes | 1 | 9 (o) | Dependent on service/site storage management of hard film
STGN-9 | Store – General, 9m2 | Yes | 1 | 9 (o) | For optical discs
OFF-S9 | Office – Single Person, 9m2 | Yes | 1 | 9 (o) | Optional. PACS Operations/Management Team
DISP-8 | Disposal Room, 8m2 | Yes |  | 8
| Discounted Circulation |  | 35%
GENERAL X-RAY AND FLUOROSCOPY
GENXR | General X-Ray | Yes | 1 | 35 | Includes Control console space. Number of rooms adjusted to suit service plan
SCRN | Screening Room (Fluoroscopy) | Yes | 1 | 36 (o) | Includes Control & area for contrast media prep. Number of rooms adjusted to suit service plan
ENS-ACC | Ensuite - Accessible, 7m2 | Yes | 1 | 7 | Dual access from fluoroscopy room and corridor
CHPT | Change Cubicle - Patient, 2m2 | Yes | 2 | 2 | 1 cubicle per imaging room. Less required if centralised
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 2 | 4 | 1 cubicle per imaging room. Less required if centralised
WCPT | Toilet - Patient, 4m2 | Yes
WAIT-SUB | Waiting – Sub, 5m2 | Yes |  |  | Required if Change Cubicles are centralised. Separate male & female patients.
| Patient Locker Bay |  |  |  | Operational process may be for patients to take clothes/possessions with them. For centralised changed waiting.
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes | 1 | 9 | Recovery bays. For Level 5 & 6 refer to combined general nursing holding/recovery unit.
BLIN | Bay - Linen | Yes | 1 | 2
BPPE | Bay - PPE | Yes | 2 | 0.5 | Storage outside of modality room for Lead Apron
| Computed Radiography (CR) Processing Workroom |  | 1 | 20 | Adjacent to control rooms.
| Digital Radiography (DR) Processing Workroom |  | 1 | 16 | For DR system – 8m2 per Imaging Room for equipment & workstations. Adjacent to control rooms
DTUR-10 | Dirty Utility, 10m2 | Yes | 1 | 10 (o) | Disposal of contrast media, dirty linen storage. May be shared with other modalities
BRES | Bay - Resuscitation | Yes
| Discounted Circulation |  | 35%
ULTRASOUND AND MAMMOGRAPHY
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 1 | 5 | Size according to activity data. Assumed patients are waiting and are changed
| Patient Locker Bay |  |  |  | Operational process may be for patients to take clothes / possessions with them. For Centralised Changed Waiting.
ULTR | Ultrasound | Yes | 1 | 14 (o) | Number of rooms adjusted to suit service plan.
| Ultrasound - Procedures |  |  |  | For interventional procedures. Recovery from US procedures is recommended to be in the combined holding / nursing recovery area for all interventional imaging procedures.
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2 (o) | Total 1 cubicle per US room (including accessible).
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4 (o) | Total 1 cubicle per US room.
DTUR-10 | Sonographer Work Room |  |  | 5.5 | Per person. Benching. Minimum size 8m.
XRRR | X-Ray Viewing & Reporting Room |  | Use PACS reporting workstation |  | 1 workstation per 2 rooms. Adjust size accordingly. Altrasound and Mammography.
MAMMO | Mammography | Yes |  |  | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT | Change Cubicle - Patient, 2m2 | Yes |  |  | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes |  |  | Total 2 cubicles per mammography room.
| Ultrasound / Mammography Prep Room / Lab |  |  |  | Direct access to mammography and reporting rooms.
| Discounted Circulation |  | 35%
ANGIOGRAPHY / DSA
ANIN | Anaesthetic Induction Room | Yes |  |  | More common alternative is anaesthetic induction in the angiography room.
SCRB-6 | Scrub-Up / Gowning, 6m2 | Yes |  |  | 2 scrub stations. May be shared between Angiography Rooms if design permits.
ANSS | Angiography Sterile Store / Set-Up Room | Yes
ANPR | Angiography Room | Yes |  |  | Number of rooms adjusted to suit service plan. This is consistent with HPU 170 Cardiac Investigations Unit Angiography Room.
ANCRT | Angiography Control / Reporting Room | Yes
| Computer Equipment / Machine Room |  |  |  | 1 per Angiography Room. Refer to specifications for equipment type.
BPPE | Bay – PPE | Yes |  |  | No dependent on number of Angiography / DSA rooms. Accommodates lead aprons.
SDTUR-S | Dirty Utility – Sub, 8m2 | Yes |  |  | May be shared with other modalities in smaller units. Not in level 5 & 6 services where instruments require cleaning prior to sterilising.
XRRR | X-Ray Viewing & Reporting Room | Yes |  |  | Adjust as necessary. Angiography
BRES | Bay - Resuscitation | Yes |  |  | May be shared with recovery areas recovery trolley.
BLIN | Bay – Linen | Yes |  |  | May access allocated linen trolleys.
STGN-8 | Store - General, 8m2 |  |  |  | If a stand-alone service only for films/ CDs/ videos – this can be shared space.
Discounted Circulation
CT SCANNING
CTPR | CT Scanning | Yes | 1 | 45 (o) | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan.
CTCR | CT Control Room | Yes | 1 | 12
CTCR | CT Control Room - Shared | Yes
ANPR | CT Computer Room |  | 1 | 10 (o) | Dependent on equipment selected.
XRRR | X-Ray Viewing & Reporting Room | Yes | 1 | 9 | One workstation per two CT Scanning Rooms.
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 1 | 4 | 1 per CT room.
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 2 | 4 (6 seats) | Optional. Shared for MRI and CT and provides male/female separation.
WCPT | Toilet - Patient, 4m2 | Yes
PBTR-H-6 | Patient Bay - Holding, 6m2 | Yes | 1 | 6 | 1 outside each CT scanning room
| Preparation / Chair Bay |  |  |  | Space for a chair and storage with a curtain for privacy for inserting cannula if IV contrast is required.
| Patient Locker Bay |  | 1 | 1 | Shared between MRI and CT.
CLUR-8 | Clean Utility – Sub, 8m2 | Yes |  |  | May be shared by MRI and CT and other modalities.
DTUR-8 | Dirty Utility – Sub, 8m2 | Yes |  |  | May be shared with other modalities.
SCRB-6 | Scrub-up/ Gowning, 6m2 | Yes |  |  | 2 stations. May be shared between rooms.
BHWS-B | Bay – Handwashing, Type B | Yes |  |  | Part of preparation area.
| Discounted Circulation |  | 35%
MRI
ANIN | Anaesthetic Induction Room | Yes |  |  | Include small staff base/clean utility and 1 patient trolley/bed bay per room. Also used for recovery
BHWS-B | Bay – Handwashing, Type B | Yes |  |  | Part of preparation area.
| MRI Scanning Room |  |  |  | Size will depend on equipment selected.
| MRI Control Room |  |  |  | Must oversee and control entry into magnet room.
MRI Computer Room
MRI Reporting Room
CHPT | Change Cubicle - Patient, 2m2 | Yes
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes
BLIN | Bay - Linen | Yes |  |  | Optional if stand-alone unit.
| Patient Locker Bay | Yes
WAIT-SUB | Waiting - Sub, 5m2 | Yes
WCPT | Toilet - Patient, 4m2 | Yes
Bay – Wheelchair Park/ Trolley Park
BRES | Bay - Resuscitation | Yes |  |  | Non-ferrous construction
| Store - Dewar Tank |  |  |  | Must provide easy access into MRI room for top-up using Dewar.
Discounted Circulation
PATIENT HOLDING / RECOVERY
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  |  | At least two patient holding / recovery bays per interventional room are required (three bed spaces for CT), generally configured as a single area to maximise staff efficiencies. MRI patient holding / recovery bays however are excluded and are instead located in the MRI suite – MRI Induction / Preparation / Recovery Room. Bays should have direct observation from Staff Station. They may be used for preparation
SSTN-10 | Staff Station, 10m2 | Yes
CLUR-12 | Clean Utility / Medication Room | Yes
BLIN | Bay - Linen | Yes
BRES | Bay - Resuscitation | Yes
DTUR-S | Dirty Utility - Sub, 8m2 | Yes
STEQ-14 | Store - Equipment | Yes
Discounted Circulation
STAFF OFFICES AND REPORTING
OFF-S12 | Office - Single Person, 12m2 | Yes |  |  | Director
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 | Chief Radiologist
OFF-S9 | Office - Single Person, 9m2 | Yes |  |  | Chief Radiographer
| PACS Reporting Workstation |  |  | 5.5 | Number of spaces as per staff establishment.
| Office - Workstation, 5.5m2 |  |  |  | Number of spaces as per staff establishment.
OFF-S9 | Office - Single Person, 9m2 | Yes |  |  | NUM
| Office - Workstation, 5.5m2 |  |  |  | Number of spaces as per staff establishment for Transcription
| Office - Workstation, 5.5m2 |  |  |  | Number of spaces as per staff establishment for IT & Administration
| Discounted Circulation |  | 35%
STAFF AREAS
| Film Library / Study |  |  |  | 5 places
SSPS-8 | Store - Photocopy / Stationery, 8m2 | Yes
MEET-L-20 | Meeting Room, 20m2 | Yes
MEET-L-30 | Meeting Room, 30m2 | Yes |  |  | Access to larger meeting room.
SRM-15 | Staff Room, 15m2 | Yes |  |  | Depending on Staff Establishment. Shared staff room for small Units.
PROP-2 | Property Bay - Staff | Yes |  |  | Numbers dependent on staff establishment.
SHST | Shower - Staff, 3m2 | Yes | 1 | 3 (o) | Numbers dependent on staff establishment.
WCST | Toilet - Staff, 3m2 | Yes | 2 | 3 | No. dependent on peak staffing numbers. Easily accessible if main change rooms remote.
CHST-10 | Change Room - Staff (Male / Female), 10m2 | Yes |  |  | 1 Shower, 2 WCs plus lockers. Required for facilities with high interventional workload.
| Discounted Circulation % |  | 35%

## Level 4
Level 4
AusHFG Room Code | Room / Space | SC / SC-D | Qty | m2 | Remarks
ENTRY / RECEPTION / CLERICAL
WAIT-20 | Waiting, 20m2 | Yes | 1 | 20 | 1.2 m2 per seat, 1.5m2 for w/chair. Respectively 15, 25 & 35 seats/ wheelchairs
BWD-1 | Bay - Water Dispenser | Yes | 1 | 1
BVM-3 | Bay - Vending | Yes | 1 | 3 (o)
PLAP-10 | Play Area - Paediatric, 10m2 | Yes | 1 | 10 (o) | 4-5 places for children
WCAC | Toilet – Accessible, 6m2 | Yes | 1 | 5 | Unless available nearby
WCPU-3 | Toilet - Public, 3m2 | Yes | 2 | 3 | Unless available nearby
RECL-10 | Reception / Clerical, 10m2 | Yes | 1 | 10 | 1, 2 & 2 staff
OFF-3P | Office -Shared | Yes | 1 | 12 | Clerical staff. May need larger area if they also combine booking system for the facility/AHS
| Film Storage - Current |  | 1 | 50 (o) | Check need according to use of PACS and film management policy
CONS | Consult Room | Yes | 1 | 12 | Number dependent on volumes and patient requirements
BMEQ-4 | Bay - Mobile Equipment, 4m2 | Yes | 2 | 4 | For mobile units depending on policy re decentralising
| Trolley / Wheelchair Park |  | 1 | 6 | Smaller area required if porter service provided by central Transport Unit
| Transport Staff Workbase |  | 1 | 4 (o) | Not required if porter service provided by central Transport Unit.
| Discounted Circulation |  | 35%
SUPPORT AREAS
CLRM-5 | Cleaner's Room, 5m2 | Yes | 1 | 5 | A separate cleaners room is scheduled for MRI if no collocated with Medical Imaging
| Laser Printing / Digitiser Room |  | 1 | 6 | This may be included as work space in the radiographers
| PACS Server Room |  | 1 | 12
STGN-9 | Store - General, 9m2 | Yes | 1 | 9 | Bulk stock e.g. contrast media, CT syringes, bulk items related to patient care
STSS_12 | Store – Sterile Stock, 12m2 | Yes | 1 | 12 | Bulk clinical sterile stock
STGN-9 | Store – General, 9m2 | Yes | 1 | 9 (o) | Dependent on service/site storage management of hard film
STGN-9 | Store – General, 9m2 | Yes | 1 | 9 (o) | For optical discs
OFF-S9 | Office – Single Person, 9m2 | Yes | 1 | 9 (o) | Optional. PACS Operations/Management Team
DISP-8 | Disposal Room, 8m2 | Yes |  | 8
| Discounted Circulation |  | 35%
GENERAL X-RAY AND FLUOROSCOPY
GENXR | General X-Ray | Yes | 2 | 35 | Includes Control console space. Number of rooms adjusted to suit service plan
SCRN | Screening Room (Fluoroscopy) | Yes | 1 | 36 (o) | Includes Control & area for contrast media prep. Number of rooms adjusted to suit service plan
ENS-ACC | Ensuite - Accessible, 7m2 | Yes | 1 | 7 | Dual access from fluoroscopy room and corridor
CHPT | Change Cubicle - Patient, 2m2 | Yes | 3 | 2 | 1 cubicle per imaging room. Less required if centralised
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 3 | 4 | 1 cubicle per imaging room. Less required if centralised
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
WAIT-SUB | Waiting – Sub, 5m2 | Yes | 2 | 5 (8 seats) | Required if Change Cubicles are centralised. Separate male & female patients.
| Patient Locker Bay |  | 1 | 1 (o) | Operational process may be for patients to take clothes/possessions with them. For centralised changed waiting.
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes | 2 | 9 | Recovery bays. For Level 5 & 6 refer to combined general nursing holding/recovery unit.
BLIN | Bay - Linen | Yes | 1 | 2
BPPE | Bay - PPE | Yes | 3 | 0.5 | Storage outside of modality room for Lead Apron
| Computed Radiography (CR) Processing Workroom |  | 1 | 20 | Adjacent to control rooms.
| Digital Radiography (DR) Processing Workroom |  | 1 | 24 | For DR system – 8m2 per Imaging Room for equipment & workstations. Adjacent to control rooms
DTUR-10 | Dirty Utility, 10m2 | Yes | 1 | 10 (o) | Disposal of contrast media, dirty linen storage. May be shared with other modalities
BRES | Bay - Resuscitation | Yes | 1 | 1.5 (o)
| Discounted Circulation |  | 35%
ULTRASOUND AND MAMMOGRAPHY
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 1 | 8 | Size according to activity data. Assumed patients are waiting and are changed
| Patient Locker Bay |  | 1 | 1 (o) | Operational process may be for patients to take clothes / possessions with them. For Centralised Changed Waiting.
ULTR | Ultrasound | Yes | 2 | 14 | Number of rooms adjusted to suit service plan.
| Ultrasound - Procedures |  |  |  | For interventional procedures. Recovery from US procedures is recommended to be in the combined holding / nursing recovery area for all interventional imaging procedures.
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2 | Total 1 cubicle per US room (including accessible).
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4 | Total 1 cubicle per US room.
DTUR-10 | Sonographer Work Room |  |  | 5.5 | Per person. Benching. Minimum size 8m.
XRRR | X-Ray Viewing & Reporting Room |  | 1 | 8 | 1 workstation per 2 rooms. Adjust size accordingly. Altrasound and Mammography.
MAMMO | Mammography | Yes | 1 | 16 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4 | Total 2 cubicles per mammography room.
| Ultrasound / Mammography Prep Room / Lab |  |  |  | Direct access to mammography and reporting rooms.
| Discounted Circulation |  | 35%
ANGIOGRAPHY / DSA
ANIN | Anaesthetic Induction Room | Yes |  | 15 (o) | More common alternative is anaesthetic induction in the angiography room.
SCRB-6 | Scrub-Up / Gowning, 6m2 | Yes | 1 | 6 | 2 scrub stations. May be shared between Angiography Rooms if design permits.
ANSS | Angiography Sterile Store / Set-Up Room | Yes | 1 | 16
ANPR | Angiography Room | Yes |  | 55 | Number of rooms adjusted to suit service plan. This is consistent with HPU 170 Cardiac Investigations Unit Angiography Room.
ANCRT | Angiography Control / Reporting Room | Yes | 1 | 14
| Computer Equipment / Machine Room |  | 1 | 6 | 1 per Angiography Room. Refer to specifications for equipment type.
BPPE | Bay – PPE | Yes |  | 0.5 | No dependent on number of Angiography / DSA rooms. Accommodates lead aprons.
SDTUR-S | Dirty Utility – Sub, 8m2 | Yes | 1 | 8 (o) | May be shared with other modalities in smaller units. Not in level 5 & 6 services where instruments require cleaning prior to sterilising.
XRRR | X-Ray Viewing & Reporting Room | Yes | 1 | 8 | Adjust as necessary. Angiography
BRES | Bay - Resuscitation | Yes | 1 | 5 (o) | May be shared with recovery areas recovery trolley.
BLIN | Bay – Linen | Yes | 1 | 2 (o) | May access allocated linen trolleys.
STGN-8 | Store - General, 8m2 |  | 1 | 8 (o) | If a stand-alone service only for films/ CDs/ videos – this can be shared space.
| Discounted Circulation |  | 35%
CT SCANNING
CTPR | CT Scanning | Yes | 1 | 45 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan.
CTCR | CT Control Room | Yes | 1 | 12
CTCR | CT Control Room - Shared | Yes
ANPR | CT Computer Room |  | 1 | 10 (o) | Dependent on equipment selected.
XRRR | X-Ray Viewing & Reporting Room | Yes | 1 | 9 | One workstation per two CT Scanning Rooms.
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 1 | 4 | 1 per CT room.
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 2 | 5 (8 seats) | Optional. Shared for MRI and CT and provides male/female separation.
WCPT | Toilet - Patient, 4m2 | Yes
PBTR-H-6 | Patient Bay - Holding, 6m2 | Yes | 1 | 6 | 1 outside each CT scanning room
| Preparation / Chair Bay |  |  |  | Space for a chair and storage with a curtain for privacy for inserting cannula if IV contrast is required.
| Patient Locker Bay |  | 1 | 1 | Shared between MRI and CT.
CLUR-8 | Clean Utility – Sub, 8m2 | Yes | 1 | 8 | May be shared by MRI and CT and other modalities.
DTUR-8 | Dirty Utility – Sub, 8m2 | Yes | 1 | 8 (o) | May be shared with other modalities.
SCRB-6 | Scrub-up/ Gowning, 6m2 | Yes | 1 | 6 | 2 stations. May be shared between rooms.
BHWS-B | Bay – Handwashing, Type B | Yes | 1 | 1 | Part of preparation area.
| Discounted Circulation |  | 35%
MRI
ANIN | Anaesthetic Induction Room | Yes | 1 | 15 | Include small staff base/clean utility and 1 patient trolley/bed bay per room. Also used for recovery
BHWS-B | Bay – Handwashing, Type B | Yes | 1 | 1 | Part of preparation area.
| MRI Scanning Room |  | 1 | 42 | Size will depend on equipment selected.
| MRI Control Room |  | 1 | 14 | Must oversee and control entry into magnet room.
| MRI Computer Room |  | 1 | 10
| MRI Reporting Room |  | 1 | 9
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4
BLIN | Bay - Linen | Yes | 1 | 2 | Optional if stand-alone unit.
| Patient Locker Bay | Yes | 1 | 1
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 1 | 5
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
| Bay – Wheelchair Park/ Trolley Park |  | 1 | 6
BRES | Bay - Resuscitation | Yes | 1 | 1.5 | Non-ferrous construction
| Store - Dewar Tank |  | Remote areas only |  | Must provide easy access into MRI room for top-up using Dewar.
| Discounted Circulation |  | 35%
PATIENT HOLDING / RECOVERY
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  | 9 | At least two patient holding / recovery bays per interventional room are required (three bed spaces for CT), generally configured as a single area to maximise staff efficiencies. MRI patient holding / recovery bays however are excluded and are instead located in the MRI suite – MRI Induction / Preparation / Recovery Room. Bays should have direct observation from Staff Station. They may be used for preparation
SSTN-10 | Staff Station, 10m2 | Yes | 1 | 10
CLUR-12 | Clean Utility / Medication Room | Yes | 1 | 10
BLIN | Bay - Linen | Yes | 1 | 2
BRES | Bay - Resuscitation | Yes | 1 | 1.5
DTUR-S | Dirty Utility - Sub, 8m2 | Yes | 1 | 8
STEQ-14 | Store - Equipment | Yes | 1 | 12
| Discounted Circulation |  | 35%
STAFF OFFICES AND REPORTING
OFF-S12 | Office - Single Person, 12m2 | Yes |  |  | Director
OFF-S9 | Office - Single Person, 9m2 | Yes |  | 9 | Chief Radiologist
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 | Chief Radiographer
| PACS Reporting Workstation |  |  | 5.5 | Number of spaces as per staff establishment.
| Office - Workstation, 5.5m2 |  |  |  | Number of spaces as per staff establishment.
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 (o) | NUM
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment for Transcription
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment for IT & Administration
| Discounted Circulation |  | 35%
STAFF AREAS
| Film Library / Study |  |  |  | 5 places
SSPS-8 | Store - Photocopy / Stationery, 8m2 | Yes
MEET-L-20 | Meeting Room, 20m2 | Yes
MEET-L-30 | Meeting Room, 30m2 | Yes |  |  | Access to larger meeting room.
SRM-15 | Staff Room, 15m2 | Yes | 1 | 15 | Depending on Staff Establishment. Shared staff room for small Units.
PROP-2 | Property Bay - Staff | Yes | 1 | 2 | Numbers dependent on staff establishment.
SHST | Shower - Staff, 3m2 | Yes | 1 | 3 (o) | Numbers dependent on staff establishment.
WCST | Toilet - Staff, 3m2 | Yes | 2 | 3 | No. dependent on peak staffing numbers. Easily accessible if main change rooms remote.
CHST-10 | Change Room - Staff (Male / Female), 10m2 | Yes |  |  | 1 Shower, 2 WCs plus lockers. Required for facilities with high interventional workload.
| Discounted Circulation % |  | 35%

## Level 5
Level 5
AusHFG Room Code | Room / Space | SC / SC-D | Qty | m2 | Remarks
ENTRY / RECEPTION / CLERICAL
WAIT-20 | Waiting, 20m2 | Yes | 1 | 35 | 1.2 m2 per seat, 1.5m2 for w/chair. Respectively 15, 25 & 35 seats/ wheelchairs
BWD-1 | Bay - Water Dispenser | Yes | 1 | 1
BVM-3 | Bay - Vending | Yes | 1 | 3 (o)
PLAP-10 | Play Area - Paediatric, 10m2 | Yes | 1 | 10 (o) | 4-5 places for children
WCAC | Toilet – Accessible, 6m2 | Yes | 1 | 5 | Unless available nearby
WCPU-3 | Toilet - Public, 3m2 | Yes | 1 | 3 | Unless available nearby
RECL-10 | Reception / Clerical, 10m2 | Yes | 1 | 12 | 1, 2 & 2 staff
OFF-3P | Office -Shared | Yes | 1 | 15 | Clerical staff. May need larger area if they also combine booking system for the facility/AHS
| Film Storage - Current |  | 1 | 70 (o) | Check need according to use of PACS and film management policy
CONS | Consult Room | Yes | 2 | 12 | Number dependent on volumes and patient requirements
BMEQ-4 | Bay - Mobile Equipment, 4m2 | Yes | 3 | 4 | For mobile units depending on policy re decentralising
| Trolley / Wheelchair Park |  | 1 | 10 | Smaller area required if porter service provided by central Transport Unit
| Transport Staff Workbase |  | 1 | 6 (o) | Not required if porter service provided by central Transport Unit.
| Discounted Circulation |  | 35%
SUPPORT AREAS
CLRM-5 | Cleaner's Room, 5m2 | Yes |  |  | A separate cleaners room is scheduled for MRI if no collocated with Medical Imaging
| Laser Printing / Digitiser Room |  |  |  | This may be included as work space in the radiographers
PACS Server Room
STGN-9 | Store - General, 9m2 | Yes |  |  | Bulk stock e.g. contrast media, CT syringes, bulk items related to patient care
STSS_12 | Store – Sterile Stock, 12m2 | Yes |  |  | Bulk clinical sterile stock
STGN-9 | Store – General, 9m2 | Yes |  |  | Dependent on service/site storage management of hard film
STGN-9 | Store – General, 9m2 | Yes |  |  | For optical discs
OFF-S9 | Office – Single Person, 9m2 | Yes |  |  | Optional. PACS Operations/Management Team
DISP-8 | Disposal Room, 8m2 | Yes
| Discounted Circulation |  | 35%
GENERAL X-RAY AND FLUOROSCOPY
GENXR | General X-Ray | Yes | 3 | 35 | Includes Control console space. Number of rooms adjusted to suit service plan
SCRN | Screening Room (Fluoroscopy) | Yes | 1 | 36 (o) | Includes Control & area for contrast media prep. Number of rooms adjusted to suit service plan
ENS-ACC | Ensuite - Accessible, 7m2 | Yes | 1 | 7 | Dual access from fluoroscopy room and corridor
CHPT | Change Cubicle - Patient, 2m2 | Yes | 4 | 2 | 1 cubicle per imaging room. Less required if centralised
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 4 | 4 | 1 cubicle per imaging room. Less required if centralised
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
WAIT-SUB | Waiting – Sub, 5m2 | Yes | 2 | 8 (10 seats) | Required if Change Cubicles are centralised. Separate male & female patients.
| Patient Locker Bay |  | 1 | 2 (o) | Operational process may be for patients to take clothes/possessions with them. For centralised changed waiting.
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  |  | Recovery bays. For Level 5 & 6 refer to combined general nursing holding/recovery unit.
BLIN | Bay - Linen | Yes | 1 | 2
BPPE | Bay - PPE | Yes | 4 | 0.5 | Storage outside of modality room for Lead Apron
| Computed Radiography (CR) Processing Workroom |  | 1 | 30 | Adjacent to control rooms.
| Digital Radiography (DR) Processing Workroom |  | 1 | 30 | For DR system – 8m2 per Imaging Room for equipment & workstations. Adjacent to control rooms
DTUR-10 | Dirty Utility, 10m2 | Yes | 1 | 10 (o) | Disposal of contrast media, dirty linen storage. May be shared with other modalities
BRES | Bay - Resuscitation | Yes | 1 | 1.5
| Discounted Circulation |  | 35%
ULTRASOUND AND MAMMOGRAPHY
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 1 | 10 | Size according to activity data. Assumed patients are waiting and are changed
| Patient Locker Bay |  | 1 | 2 (o) | Operational process may be for patients to take clothes / possessions with them. For Centralised Changed Waiting.
ULTR | Ultrasound | Yes | 2 | 14 | Number of rooms adjusted to suit service plan.
| Ultrasound - Procedures |  | 1 | 20 | For interventional procedures. Recovery from US procedures is recommended to be in the combined holding / nursing recovery area for all interventional imaging procedures.
CHPT | Change Cubicle - Patient, 2m2 | Yes | 2 | 2 | Total 1 cubicle per US room (including accessible).
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4 | Total 1 cubicle per US room.
DTUR-10 | Sonographer Work Room |  |  | 5.5 | Per person. Benching. Minimum size 8m.
XRRR | X-Ray Viewing & Reporting Room |  | 1 | 12 | 1 workstation per 2 rooms. Adjust size accordingly. Altrasound and Mammography.
MAMMO | Mammography | Yes | 1 | 16 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4 | Total 2 cubicles per mammography room.
| Ultrasound / Mammography Prep Room / Lab |  | 1 | 10 | Direct access to mammography and reporting rooms.
| Discounted Circulation |  | 35%
ANGIOGRAPHY / DSA
ANIN | Anaesthetic Induction Room | Yes |  | 15 (o) | More common alternative is anaesthetic induction in the angiography room.
SCRB-6 | Scrub-Up / Gowning, 6m2 | Yes | 1 | 6 | 2 scrub stations. May be shared between Angiography Rooms if design permits.
ANSS | Angiography Sterile Store / Set-Up Room | Yes | 1 | 18
ANPR | Angiography Room | Yes |  | 55 | Number of rooms adjusted to suit service plan. This is consistent with HPU 170 Cardiac Investigations Unit Angiography Room.
ANCRT | Angiography Control / Reporting Room | Yes | 1 | 14
| Computer Equipment / Machine Room |  | 2 | 6 | 1 per Angiography Room. Refer to specifications for equipment type.
BPPE | Bay – PPE | Yes |  | 0.5 | No dependent on number of Angiography / DSA rooms. Accommodates lead aprons.
SDTUR-S | Dirty Utility – Sub, 8m2 | Yes | 1 | 8 (o) | May be shared with other modalities in smaller units. Not in level 5 & 6 services where instruments require cleaning prior to sterilising.
XRRR | X-Ray Viewing & Reporting Room | Yes | 1 | 12 | Adjust as necessary. Angiography
BRES | Bay - Resuscitation | Yes | 1 | 5 (o) | May be shared with recovery areas recovery trolley.
BLIN | Bay – Linen | Yes | 1 | 2 (o) | May access allocated linen trolleys.
STGN-8 | Store - General, 8m2 |  | 1 | 8 (o) | If a stand-alone service only for films/ CDs/ videos – this can be shared space.
| Discounted Circulation |  | 35%
CT SCANNING
CTPR | CT Scanning | Yes | 2 | 45 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan.
CTCR | CT Control Room | Yes
CTCR | CT Control Room - Shared | Yes | 1 | 20
ANPR | CT Computer Room |  | 2 | 10 (o) | Dependent on equipment selected.
XRRR | X-Ray Viewing & Reporting Room | Yes | 1 | 9 | One workstation per two CT Scanning Rooms.
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 2 | 4 | 1 per CT room.
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 2 | 8 (10 seats) | Optional. Shared for MRI and CT and provides male/female separation.
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
PBTR-H-6 | Patient Bay - Holding, 6m2 | Yes | 2 | 6 | 1 outside each CT scanning room
| Preparation / Chair Bay |  | 1 | 9 | Space for a chair and storage with a curtain for privacy for inserting cannula if IV contrast is required.
| Patient Locker Bay |  | 1 | 1 | Shared between MRI and CT.
CLUR-8 | Clean Utility – Sub, 8m2 | Yes | 1 | 8 | May be shared by MRI and CT and other modalities.
DTUR-8 | Dirty Utility – Sub, 8m2 | Yes | 1 | 8 (o) | May be shared with other modalities.
SCRB-6 | Scrub-up/ Gowning, 6m2 | Yes | 1 | 6 | 2 stations. May be shared between rooms.
BHWS-B | Bay – Handwashing, Type B | Yes | 1 | 1 | Part of preparation area.
| Discounted Circulation |  | 35%
MRI
ANIN | Anaesthetic Induction Room | Yes | 1 | 15 | Include small staff base/clean utility and 1 patient trolley/bed bay per room. Also used for recovery
BHWS-B | Bay – Handwashing, Type B | Yes | 1 | 1 | Part of preparation area.
| MRI Scanning Room |  | 1 | 42 | Size will depend on equipment selected.
| MRI Control Room |  | 1 | 14 | Must oversee and control entry into magnet room.
| MRI Computer Room |  | 1 | 10
| MRI Reporting Room |  | 1 | 9
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4
BLIN | Bay - Linen | Yes | 1 | 2 | Optional if stand-alone unit.
| Patient Locker Bay | Yes | 1 | 1
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 1 | 5
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
| Bay – Wheelchair Park/ Trolley Park |  | 1 | 6
BRES | Bay - Resuscitation | Yes | 1 | 1.5 | Non-ferrous construction
| Store - Dewar Tank |  | Remote areas only |  | Must provide easy access into MRI room for top-up using Dewar.
| Discounted Circulation |  | 35%
PATIENT HOLDING / RECOVERY
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  | 9 | At least two patient holding / recovery bays per interventional room are required (three bed spaces for CT), generally configured as a single area to maximise staff efficiencies. MRI patient holding / recovery bays however are excluded and are instead located in the MRI suite – MRI Induction / Preparation / Recovery Room. Bays should have direct observation from Staff Station. They may be used for preparation
SSTN-10 | Staff Station, 10m2 | Yes | 1 | 10
CLUR-12 | Clean Utility / Medication Room | Yes | 1 | 10
BLIN | Bay - Linen | Yes | 1 | 2
BRES | Bay - Resuscitation | Yes | 1 | 1.5
DTUR-S | Dirty Utility - Sub, 8m2 | Yes | 1 | 8
STEQ-14 | Store - Equipment | Yes | 1 | 12
| Discounted Circulation |  | 35%
STAFF OFFICES AND REPORTING
OFF-S12 | Office - Single Person, 12m2 | Yes | 1 | 12 | Director
OFF-S9 | Office - Single Person, 9m2 | Yes |  | 9 | Chief Radiologist
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 | Chief Radiographer
| PACS Reporting Workstation |  |  | 5.5 | Number of spaces as per staff establishment.
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment.
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 | NUM
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment for Transcription
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment for IT & Administration
| Discounted Circulation |  | 35%
STAFF AREAS
| Film Library / Study |  | 1 | 20 | 5 places
SSPS-8 | Store - Photocopy / Stationery, 8m2 | Yes | 1 | 8
MEET-L-20 | Meeting Room, 20m2 | Yes | 1 | 20
MEET-L-30 | Meeting Room, 30m2 | Yes | 1 | 30 | Access to larger meeting room.
SRM-15 | Staff Room, 15m2 | Yes | 1 | 20 | Depending on Staff Establishment. Shared staff room for small Units.
PROP-2 | Property Bay - Staff | Yes |  |  | Numbers dependent on staff establishment.
SHST | Shower - Staff, 3m2 | Yes |  |  | Numbers dependent on staff establishment.
WCST | Toilet - Staff, 3m2 | Yes |  | 3 | No. dependent on peak staffing numbers. Easily accessible if main change rooms remote.
CHST-10 | Change Room - Staff (Male / Female), 10m2 | Yes | 1 | 14 | 1 Shower, 2 WCs plus lockers. Required for facilities with high interventional workload.
| Discounted Circulation % |  | 35%

## Level 6
Level 6
AusHFG Room Code | Room / Space | SC / SC-D | Qty | m2 | Remarks
ENTRY / RECEPTION / CLERICAL
WAIT-20 | Waiting, 20m2 | Yes | 1 | 45 | 1.2 m2 per seat, 1.5m2 for w/chair. Respectively 15, 25 & 35 seats/ wheelchairs
BWD-1 | Bay - Water Dispenser | Yes | 1 | 1
BVM-3 | Bay - Vending | Yes | 1 | 3 (o)
PLAP-10 | Play Area - Paediatric, 10m2 | Yes | 1 | 10 (o) | 4-5 places for children
WCAC | Toilet – Accessible, 6m2 | Yes | 1 | 5 | Unless available nearby
WCPU-3 | Toilet - Public, 3m2 | Yes | 2 | 3 | Unless available nearby
RECL-10 | Reception / Clerical, 10m2 | Yes | 1 | 12 | 1, 2 & 2 staff
OFF-3P | Office -Shared | Yes | 1 | 20 | Clerical staff. May need larger area if they also combine booking system for the facility/AHS
| Film Storage - Current |  | 1 | 100 (o) | Check need according to use of PACS and film management policy
CONS | Consult Room | Yes | 2 | 12 | Number dependent on volumes and patient requirements
BMEQ-4 | Bay - Mobile Equipment, 4m2 | Yes | 4 | 4 | For mobile units depending on policy re decentralising
| Trolley / Wheelchair Park |  | 1 | 12 | Smaller area required if porter service provided by central Transport Unit
| Transport Staff Workbase |  | 1 | 8 (o) | Not required if porter service provided by central Transport Unit.
| Discounted Circulation |  | 35%
SUPPORT AREAS
CLRM-5 | Cleaner's Room, 5m2 | Yes |  |  | A separate cleaners room is scheduled for MRI if no collocated with Medical Imaging
| Laser Printing / Digitiser Room |  |  |  | This may be included as work space in the radiographers
PACS Server Room
STGN-9 | Store - General, 9m2 | Yes |  |  | Bulk stock e.g. contrast media, CT syringes, bulk items related to patient care
STSS_12 | Store – Sterile Stock, 12m2 | Yes |  |  | Bulk clinical sterile stock
STGN-9 | Store – General, 9m2 | Yes |  |  | Dependent on service/site storage management of hard film
STGN-9 | Store – General, 9m2 | Yes |  |  | For optical discs
OFF-S9 | Office – Single Person, 9m2 | Yes |  |  | Optional. PACS Operations/Management Team
DISP-8 | Disposal Room, 8m2 | Yes
| Discounted Circulation |  | 35%
GENERAL X-RAY AND FLUOROSCOPY
GENXR | General X-Ray | Yes | 4 | 35 | Includes Control console space. Number of rooms adjusted to suit service plan
SCRN | Screening Room (Fluoroscopy) | Yes | 1 | 36 (o) | Includes Control & area for contrast media prep. Number of rooms adjusted to suit service plan
ENS-ACC | Ensuite - Accessible, 7m2 | Yes | 1 | 7 | Dual access from fluoroscopy room and corridor
CHPT | Change Cubicle - Patient, 2m2 | Yes | 5 | 2 | 1 cubicle per imaging room. Less required if centralised
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 5 | 4 | 1 cubicle per imaging room. Less required if centralised
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
WAIT-SUB | Waiting – Sub, 5m2 | Yes | 2 | 10 (14 seats) | Required if Change Cubicles are centralised. Separate male & female patients.
| Patient Locker Bay |  | 1 | 2 (o) | Operational process may be for patients to take clothes/possessions with them. For centralised changed waiting.
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  |  | Recovery bays. For Level 5 & 6 refer to combined general nursing holding/recovery unit.
BLIN | Bay - Linen | Yes | 2 | 2
BPPE | Bay - PPE | Yes | 5 | 0.5 | Storage outside of modality room for Lead Apron
| Computed Radiography (CR) Processing Workroom |  | 1 | 40 | Adjacent to control rooms.
| Digital Radiography (DR) Processing Workroom |  | 1 | 40 | For DR system – 8m2 per Imaging Room for equipment & workstations. Adjacent to control rooms
DTUR-10 | Dirty Utility, 10m2 | Yes | 1 | 12 (o) | Disposal of contrast media, dirty linen storage. May be shared with other modalities
BRES | Bay - Resuscitation | Yes | 1 | 1.5
| Discounted Circulation |  | 35%
ULTRASOUND AND MAMMOGRAPHY
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 1 | 10 | Size according to activity data. Assumed patients are waiting and are changed
| Patient Locker Bay |  | 1 | 2 (o) | Operational process may be for patients to take clothes / possessions with them. For Centralised Changed Waiting.
ULTR | Ultrasound | Yes | 2 | 14 | Number of rooms adjusted to suit service plan.
| Ultrasound - Procedures |  | 1 | 24 | For interventional procedures. Recovery from US procedures is recommended to be in the combined holding / nursing recovery area for all interventional imaging procedures.
CHPT | Change Cubicle - Patient, 2m2 | Yes | 2 | 2 | Total 1 cubicle per US room (including accessible).
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4 | Total 1 cubicle per US room.
DTUR-10 | Sonographer Work Room |  |  | 5.5 | Per person. Benching. Minimum size 8m.
XRRR | X-Ray Viewing & Reporting Room |  | 1 | 12 | 1 workstation per 2 rooms. Adjust size accordingly. Altrasound and Mammography.
MAMMO | Mammography | Yes | 1 | 16 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan. Number of rooms adjusted to suit service plan.
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4 | Total 2 cubicles per mammography room.
| Ultrasound / Mammography Prep Room / Lab |  | 1 | 10 | Direct access to mammography and reporting rooms.
| Discounted Circulation |  | 35%
ANGIOGRAPHY / DSA
ANIN | Anaesthetic Induction Room | Yes |  | 15 (o) | More common alternative is anaesthetic induction in the angiography room.
SCRB-6 | Scrub-Up / Gowning, 6m2 | Yes | 1 | 6 | 2 scrub stations. May be shared between Angiography Rooms if design permits.
ANSS | Angiography Sterile Store / Set-Up Room | Yes | 1 | 24
ANPR | Angiography Room | Yes |  | 55 | Number of rooms adjusted to suit service plan. This is consistent with HPU 170 Cardiac Investigations Unit Angiography Room.
ANCRT | Angiography Control / Reporting Room | Yes | 1 | 14
| Computer Equipment / Machine Room |  | 3 | 6 | 1 per Angiography Room. Refer to specifications for equipment type.
BPPE | Bay – PPE | Yes |  | 0.5 | No dependent on number of Angiography / DSA rooms. Accommodates lead aprons.
SDTUR-S | Dirty Utility – Sub, 8m2 | Yes | 1 | 8 (o) | May be shared with other modalities in smaller units. Not in level 5 & 6 services where instruments require cleaning prior to sterilising.
XRRR | X-Ray Viewing & Reporting Room | Yes | 1 | 16 | Adjust as necessary. Angiography
BRES | Bay - Resuscitation | Yes | 1 | 5 (o) | May be shared with recovery areas recovery trolley.
BLIN | Bay – Linen | Yes | 1 | 2 (o) | May access allocated linen trolleys.
STGN-8 | Store - General, 8m2 |  | 1 | 8 (o) | If a stand-alone service only for films/ CDs/ videos – this can be shared space.
| Discounted Circulation |  | 35%
CT SCANNING
CTPR | CT Scanning | Yes | 2 | 45 | Mammography room – Not all Units offer a mammography service. Inclusion of areas for mammography is subject to the service plan.
CTCR | CT Control Room | Yes
CTCR | CT Control Room - Shared | Yes | 1 | 24
ANPR | CT Computer Room |  | 2 | 10 (o) | Dependent on equipment selected.
XRRR | X-Ray Viewing & Reporting Room | Yes | 1 | 9 | One workstation per two CT Scanning Rooms.
CHPT-D | Change Cubicle - Accessible, 4m2 | Yes | 2 | 4 | 1 per CT room.
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 2 | 8 (10 seats) | Optional. Shared for MRI and CT and provides male/female separation.
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
PBTR-H-6 | Patient Bay - Holding, 6m2 | Yes | 2 | 6 | 1 outside each CT scanning room
| Preparation / Chair Bay |  | 1 | 9 | Space for a chair and storage with a curtain for privacy for inserting cannula if IV contrast is required.
| Patient Locker Bay |  | 1 | 1 | Shared between MRI and CT.
CLUR-8 | Clean Utility – Sub, 8m2 | Yes | 1 | 8 | May be shared by MRI and CT and other modalities.
DTUR-8 | Dirty Utility – Sub, 8m2 | Yes | 1 | 8 (o) | May be shared with other modalities.
SCRB-6 | Scrub-up/ Gowning, 6m2 | Yes | 1 | 6 | 2 stations. May be shared between rooms.
BHWS-B | Bay – Handwashing, Type B | Yes | 1 | 1 | Part of preparation area.
| Discounted Circulation |  | 35%
MRI
ANIN | Anaesthetic Induction Room | Yes | 1 | 15 | Include small staff base/clean utility and 1 patient trolley/bed bay per room. Also used for recovery
BHWS-B | Bay – Handwashing, Type B | Yes | 1 | 1 | Part of preparation area.
| MRI Scanning Room |  | 1 | 42 | Size will depend on equipment selected.
| MRI Control Room |  | 1 | 4 | Must oversee and control entry into magnet room.
| MRI Computer Room |  | 1 | 10
| MRI Reporting Room |  | 1 | 9
CHPT | Change Cubicle - Patient, 2m2 | Yes | 1 | 2
CHPT-D | Change Cubicle – Accessible, 4m2 | Yes | 1 | 4
BLIN | Bay - Linen | Yes | 1 | 2 | Optional if stand-alone unit.
| Patient Locker Bay | Yes | 1 | 1
WAIT-SUB | Waiting - Sub, 5m2 | Yes | 1 | 5
WCPT | Toilet - Patient, 4m2 | Yes | 1 | 4
| Bay – Wheelchair Park/ Trolley Park |  | 1 | 6
BRES | Bay - Resuscitation | Yes | 1 | 1.5 | Non-ferrous construction
| Store - Dewar Tank |  | Remote areas only |  | Must provide easy access into MRI room for top-up using Dewar.
| Discounted Circulation |  | 35%
PATIENT HOLDING / RECOVERY
PBTR-H-9 | Patient Bay - Holding, 9m2 | Yes |  | 9 | At least two patient holding / recovery bays per interventional room are required (three bed spaces for CT), generally configured as a single area to maximise staff efficiencies. MRI patient holding / recovery bays however are excluded and are instead located in the MRI suite – MRI Induction / Preparation / Recovery Room. Bays should have direct observation from Staff Station. They may be used for preparation
SSTN-10 | Staff Station, 10m2 | Yes | 1 | 10
CLUR-12 | Clean Utility / Medication Room | Yes | 1 | 10
BLIN | Bay - Linen | Yes | 1 | 2
BRES | Bay - Resuscitation | Yes | 1 | 1.5
DTUR-S | Dirty Utility - Sub, 8m2 | Yes | 1 | 8
STEQ-14 | Store - Equipment | Yes | 1 | 12
| Discounted Circulation |  | 35%
STAFF OFFICES AND REPORTING
OFF-S12 | Office - Single Person, 12m2 | Yes | 1 | 12 | Director
OFF-S9 | Office - Single Person, 9m2 | Yes |  | 9 | Chief Radiologist
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 | Chief Radiographer
| PACS Reporting Workstation |  |  | 5.5 | Number of spaces as per staff establishment.
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment.
OFF-S9 | Office - Single Person, 9m2 | Yes | 1 | 9 | NUM
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment for Transcription
| Office - Workstation, 5.5m2 |  |  | 5.5 | Number of spaces as per staff establishment for IT & Administration
| Discounted Circulation |  | 35%
STAFF AREAS
| Film Library / Study |  | 1 | 20 | 5 places
SSPS-8 | Store - Photocopy / Stationery, 8m2 | Yes | 1 | 8
MEET-L-20 | Meeting Room, 20m2 | Yes | 1 | 20
MEET-L-30 | Meeting Room, 30m2 | Yes | 2 | 30 | Access to larger meeting room.
SRM-15 | Staff Room, 15m2 | Yes | 1 | 25 | Depending on Staff Establishment. Shared staff room for small Units.
PROP-2 | Property Bay - Staff | Yes |  |  | Numbers dependent on staff establishment.
SHST | Shower - Staff, 3m2 | Yes |  |  | Numbers dependent on staff establishment.
WCST | Toilet - Staff, 3m2 | Yes |  | 3 | No. dependent on peak staffing numbers. Easily accessible if main change rooms remote.
CHST-10 | Change Room - Staff (Male / Female), 10m2 | Yes | 1 | 14 | 1 Shower, 2 WCs plus lockers. Required for facilities with high interventional workload.
| Discounted Circulation % |  | 35%