## SOA Rev 7
AUSHFG HPU 520 OPERATING SUITE
SCHEDULE OF ACCOMMODATION, REVISION 7
ENTRY / RECEPTION / WAITING
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
WAIT-30 | Waiting, 30m2 | Yes | 1 | 10 | 1 | 30 | 1 | 50 | Indicative area allocation. Requirements will depend on the proportion of cases anticipated to be day only and DOSA admissions and casemix / throughput. 1.2m2 recommended per seat, 1.5m2 per wheelchair / bariatric space.
WCAC | Toilet - Accessible, 6m2 | Yes | 1 | 6 (o) | 1 | 6 | 1 | 6 | Optional. Include if no shared facilities available nearby.
WCPU | Toilet - Public, 3m2 | Yes | 1 | 3 (o) | 2 | 3 | 2 | 3 | Optional. Include if no shared facilities available nearby.
RECP-15 | Reception, 15m2 | Yes | 1 | 10 | 1 | 15 | 1 | 20
OFF-WS | Office-Workstation | Yes |  | 4.5 |  | 4.5 |  | 4.5 | Administration workstations collocated with reception. Number of workstations dependent on staff profile.
OFF-CLN | Office- Clinical Workroom | Yes |  |  | 1 | 15 | 1 | 20 | Central coordination point for unit.
BMFD-3 BMFD-7 | Bay - Multifunction Device | Yes | 1 | 3 | 1 | 7 | 1 | 7
Intradepartmental (discounted) circulation |  |  |  | 30% |  | 30% |  | 30%
PREOPERATIVE HOLDING
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
INTV | Interview Room | Yes | 1 | 12 | 1 | 12 | 2 | 12 | Interviews with nursing staff to check details and undertake baseline observations. Clinical examinations are typically undertaken in the patient bays however consult rooms may be required for some services.
PT-HOLD-B | Patient Bay - Holding, Bed | Yes | 2 | 9 | 8 | 9 | 16 | 9 | 1 per OR, sized for trolleys, but some may be recliner chairs. Holding capacity for inpatient flows will depend on local operational policies. These may be combined with the day only holding area, PACU or a designated area depending on travel distances, patient flow and workforce implications. For high volume services eg opthalmology, 2 bays per theatre / procedure room may be required in line with HPU 270 (1 may be provided as an anaesthetic preparation room as noted below). 6.5m2 bay may be used for high volume day surgery services as per HPU 270. Isolation rooms are not typically provided in this zone, however project teams are advised to refer to local operational policies.
SHPT | Shower - Patient, 4m2 | Yes | 1 | 4 | 1 | 4 | 1 | 4 | Number provided will depend on shared access with Stage 2/3 recovery.
WCPT | Toilet - Patient, 4m2 | Yes |  |  | 1 | 4 | 1 | 4
WCAC | Toilet - Accessible, 6m2 | Yes | 1 | 6 | 1 | 6 | 1 | 6
CHPT | Change Cubicle - Patient, 2m2 | Yes |  |  | 2 | 2 | 3 | 2 | Number provided will depend on shared access with Stage 2/3 recovery.
CHPT-AC | Change Cubicle - Patient, Accessible | Yes | 1 | 4 | 1 | 4 | 1 | 4
| Property Bay - Patient |  | 1 | 1 (o) | 1 | 2 (o) | 1 | 3 (o) | Optional. Assumes quarter height lockers. Property may instead travel with the patient. Access from Stage 2 recovery also required.
SSTN-10 | Staff Station, 10m2 | Yes |  |  |  |  | 1 | 10 | Only allocated for larger units as reception could be the base used for smaller units. Locate to provide oversight of holding bay areas. Can be shared to support pre and post-operative areas.
BHWS-B | Bay - Handwashing, Type B | Yes | 1 | 1 | 2 | 1 | 4 | 1 | Refer to AusHFG Part D for further details, accessible from OR and patient holding areas.
BLIN | Bay - Linen | Yes |  |  | 1 | 2 | 1 | 2 | Min. 1 per 16 Patient Bay - Holding; corridor location with ready access to Bays. Assume share with recovery for small units.
BBW | Bay - Blanket/ Fluid Warmer | Yes |  |  | 1 | 1 | 1 | 1 | Blanket warmer only
CLN-10 | Clean Store | Yes |  |  |  |  | 1 | 8 | Shared with recovery area for smaller units.
DTUR-S | Dirty Utility - Sub, 8m2 | Yes |  |  |  |  | 1 | 8 | May be shared with recovery depending on travel distances.
Intradepartmental circulation |  |  |  | 40% |  | 40% |  | 40%
STAGE 2 / 3 RECOVERY
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
PT-HOLD-B | Patient Bay – Holding, Bed | Yes |  | 9 |  | 9 |  | 9 | Stage 2 Recovery. Number of bays will be guided by clinical services planning - Recommend 3 patient bays per OR used for day surgery. 6.5m2 bay may be used for high volume day surgery services as per HPU 270.
| Recovery Stage 3 / Discharge Lounge |  |  |  |  | 4 (o) |  | 4 (o) | Optional. Recommend 2 per OR used for day surgery.
WCPT | Toilet - Patient, 4m2 | Yes |  |  |  |  | 1 | 4 | Number provided will depend on number of holding bays. Shared access to toilets, change cubicles and shower in preoperative holding.
WCAC | Toilet - Accessible, 6m2 | Yes |  |  | 1 | 6 | 1 | 6 | Number provided will depend on number of holding bays. Shared access to toilets, change cubicles and shower in preoperative holding.
SSTN-10 | Staff Station | Yes |  |  | 1 | 8 | 1 | 12 | Assume smaller units will share with PACU or pre-operative holding. Size will depend on number of bays provided.
BHWS-B | Bay – Handwashing, Type B | Yes |  | 1 |  | 1 |  | 1 | 1 per 4 patient bays.
BBEV | Bay- Beverage | Yes | 1 | 4 | 1 | 4 | 1 | 4
BLIN | Bay – Linen | Yes |  |  |  | 2 | 1 | 2 | 1 per 16 spaces. Assume smaller units will share with PACU or pre-operative holding.
BBW | Bay - Blanket / Fluid Warmer | Yes |  |  |  | 1 | 1 | 1 | 1 per 16 spaces. Assume smaller units will share with PACU or pre-operative holding.
BMEQ | Bay - Mobile Equipment | Yes |  |  | 1 | 3 | 2 | 2 | Assume smaller units will share with PACU or pre-operative holding.
BRES | Bay - Resuscitation Trolley | Yes |  |  | 1 | 1.5 | 1 | 1.5 | Assume smaller units will share with PACU or pre-operative holding.
CLN-10 | Clean Store | Yes |  |  |  |  | 1 | 10 | Assume smaller units will share with PACU or pre-operative holding. Size will depend on volume of day only bays.
MED-14 | Medication Room | Yes |  |  |  |  | 1 | 12 | Assume smaller units will share with PACU. Size will depend on volume of day only bays.
STEQ-14 | Store - Equipment | Yes |  |  |  |  | 1 | 10 | Size will depend on volume and range of equipment. Include power for re-charging.
| Dirty Utility / Disposal Room | [5 empty] | 1 | 14 | Assume smaller units will share with PACU or pre-operative holding. Area is indicative and will need to be confirmed depending on the waste streams / types of bins and bin sizes to be accommodated.
CLRM | Cleaner's Room | Yes |  |  |  |  | 1 | 5 | Assume smaller units will share with PACU or pre-operative holding. Provide at least 1 room per 1000m2.
SRM-35 | Staff Room | Yes | 1 | 12 | 1 | 18 | 1 | 24 | Indicative only, requirements will depend on staff profile / capacity of unit that will be informed by the percentage of ORs used for day surgery. External window desirable.
WCST | Toilet - Staff | Yes | 1 | 3 | 1 | 3 | 2 | 3 | Also consider access to disabled accessible staff toilets.
Intradepartmental circulation |  |  |  | 40% |  | 40% |  | 40%
OPERATING ROOM AREA
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
ANAE-16 | Anaesthetic Preparation Room, 16m2 | Yes | 2 | 16 (o) | 8 | 16 (o) | 16 | 16 (o) | Optional. Anaesthetic preparation rooms are typically provided to facilitate patient flow and support patient preparation. Provision to be determined during planning with consideration of local operational policies and patient throughput. A larger size may be required in some cases although the use of sliding doors at the entry to the room will maximise available space.
ORGN | Operating Room - General | Yes | 2 | 60 | 8 | 60 | 16 | 60 | This is the minimum recommended area requirement and is sufficient for the majority of routine surgery. Larger sized ORs may be required for some surgical specialties where additional equipment and staff need to be accommodated, eg cardiac surgery, neurosurgery, trauma and obstetrics. The number of larger sized theatres will be informed by an analysis of the projected casemix and additional equipment and staff to be accommodated.
| Operating Room, Hybrid / Large |  |  |  |  | 75 (o) |  | 75 (o) | Optional depending on confirmed scope. Minumum area noted. Requirements will depend on equipment and staff to be accommodated. 75m2 for hybrid OR assumes a single plane angiography system. Also refer to local jurisdictional guidelines where available. Hybrid ORs incorporating biplane systems and where flexibility is required to support flexible use for non-hybrid procedures will require a larger area, subject to service and equipment requirements.
| Control Room, Hybrid OR |  |  |  |  | 15 (o) |  | 15 (o) | Optional depending on confirmed scope. 1 per Hybrid OR.
| Computer Room, Hybrid OR |  |  |  |  | 10 (o) |  | 10 (o) | Optional depending on confirmed scope. 1 per Hybrid OR.
SCRB-4 | Scrub Up, 4m2 | Yes | 2 | 4 | 8 | 4 | 16 | 4 | For surgical hand antisepsis.
| Exit Bay |  | 2 | 11 | 8 | 11 | 16 | 11 | 1 per Operating Room; if shared between 2 rooms, increase to 16m2.
| AV Integration Server Cupboard |  | 2 | 1 | 8 | 1 | 16 | 1 | Audiovisual integration equipment. May be consolidated between a number of ORs. Temperature control is required. Cabling distances must be considered.
CLUP | Clean-Up Room - Shared, 12m2 | Yes | 1 | 12 | 4 | 12 | 8 | 12 | 1 per 2 ORs maximum. 1 per 3 or 4 ORs may be achievable depending on travel distances and operational practices. Area allocation includes fluid management system dock which will not be required to every Clean-Up Room. Area requirements will vary depending on local operational arrangements relating to processing of RMDs, waste streaming and provision of fluid management systems. This area allocation may be reduced where trolleys / carts with used / soiled RMDs are transferred direct from theatre to SSU, and the 'clean up' room function is focussed on cleaning and waste management only.
Intradepartmental (discounted) circulation |  |  |  | 40% |  | 40% |  | 40%
ENDOSCOPY ROOM AREA (OPTIONAL)
CODE | ROOM/SPACE | SC / SC-D | 1 - 2 Procedure Rooms | [5 empty] | REMARKS
|  |  | Qty | m2
ENPR | Procedure Room - Endoscopy | Yes |  | 45 (o) |  |  |  |  | Optional. To be provided, depending on the service profile, where dedicated endoscopy rooms are required. Procedural services may also be provided in general ORs for flexible use of rooms and to support future changes to the case mix. Specialised units requiring angiography eg for ERCPs will also require a control room.
| Endoscope Reprocessing - Dirty |  | 1 | 13 |  |  |  |  | One double basin sink per one-two procedure rooms is assumed.
| Automated Flexible Endoscope Reprocessors (AFERs) |  | 1 | 4 |  |  |  |  | The number of AFERs should provide capacity to process two endoscopes for each procedure room. Some AFERs are able to process two endoscopes at a time or asynchronously. Rural and remote health services will require a minimum of 2 AFERs. Check dimensions of preferred supplier's AFERs to confirm allowances. Note that the Quantity "1" refers to the number of spaces, not the number of AFERs or other equipment.
| Endoscope Reprocessing - Clean |  | 1 | 8
| Endoscope Reprocessing - Storage (CESC or equivalent system) |  | 1 | 4 |  |  |  |  | Fleet size and the number of storage positions required must be confirmed with the user. Note that the Quantity "1" refers to the number of spaces, not the number of storage cabinets or storage positions. Where a storage bag system is used instead of CESCs, space will be required for the storage of the bagged endoscopes.
| Store - Chemicals |  | 1 | 2 |  |  |  |  | Area allocation will depend on the volume of chemicals required to service the projected workload. Some reprocessing agents used for endoscopy reprocessing are flammable and/or toxic. Storage of these agents will comply with jurisdictional workplace health and safety regulations. This space may need to be ventilated to exhaust air to a safe location outside the building.
CLINICAL SUPPORT AREAS
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
BBW | Bay - Blanket/ Fluid Warmer | Yes |  |  | 2 | 1 | 4 | 1 | Final number required will depend on the unit layout and associated travel distances.
BLIN | Bay - Linen | Yes | 1 | 2 | 4 | 2 | 8 | 2 | 1 per 2 Operating Rooms; corridor recess with ready access to Operating Rooms.
BMEQ | Bay - Mobile Equipment, 4m2 | Yes | 1 | 4 | 4 | 4 | 8 | 4 | 1 bay per 2 Operating Rooms; provide power outlets for recharging and data.
STSS-20 | Store - Sterile Stock | Yes | 2 | 20 | 8 | 20 | 16 | 20 | 20m2 per Operating Room; direct relationship to SSU; may be provided as a single central sterile core or a number of smaller rooms to support a pair or pod of Operating Rooms. Requirements will vary depending on the surgical casemix and location of the unit, eg for paediatrics an increased range of equipment sizes is required, a high volume of orthopaedics will require greater storage than opthalmology and remote locations may need to hold a larger volume of stock.
| Store - Prosthesis | [5 empty] | 1 | 20 (o) | Optional. For services undertaking high volume orthopaedics, locate close to orthopaedic ORs.
STGN | Store - General | Yes | 1 | 10 | 1 | 12 | 1 | 20 | IV and other fluid storage.
STGN | Store - General | Yes | 1 | 20 | 1 | 30 | 1 | 40 | For Non-Sterile/ Deboxing storage. Storage of medical gas cylinders to be considered eg if reticulated nitrous oxide is not provided.
STEQ-20 | Store - Equipment (Major) | Yes | 1 | 12 | 1 | 48 | 1 | 96 | Plan at 6m2 per Operating Room, for major equipment.
STEQ-20 | Store - Equipment (Minor) | Yes | 1 | 10 | 1 | 40 | 1 | 80 | Plan at 5m2 per Operating Room, for minor equipment.
| Anaesthetic Workroom & Biomedical Equipment |  | 1 | 10 | 1 | 15 | 1 | 20
STGN-9 | Store - General | Yes | 1 | 15 | 1 | 25 | 1 | 50 | Anaesthetic store for consumables.
| Perfusion Room - Set-up | [5 empty] | 1 | 40 | Assumes area shared between 2 cardiothoracic operating rooms
| Store - Perfusion | [5 empty] | 1 | 20 | Assumes area shared between 2 cardiothoracic operating rooms.
| Clean-Up - Perfusion | [5 empty] | 1 | 9 (o) | Optional. Cleaning and disinfection of heater-cooler units is recommended to be undertaken in SSU, however some units may provide this service within the unit. Appropriate exhaust of room is essential. Refer to HPU Section 2.2.9 for further information.
| Dispatch - Dirty (offsite sterilising) |  | 1 | 18 (o) |  |  |  |  | Optional, for off-site sterilisation models. Includes packing and storage in preparation of dispatch. May include washer disinfector for gross decontamination of instruments prior to transport (this may be located with the OR cleanup area). Refer to HPU 190 for further information.
| Receiving - Clean (offsite sterilising) |  | 1 | 10 (o) |  |  |  |  | Optional, for off-site sterilisation models. For receipt of sterilised instruments. Refer to HPU 190 for further information.
CLRM | Cleaner's Room, 5m2 | Yes |  | 5 |  | 5 |  | 5 | Provide at least 1 room per 1000m2; ready access to all areas of the unit, preferred on perimeter; one room may be sized to accommodate a scrubber
DISP-10 | Disposal Room | Yes | 1 | 10 | 1 | 15 | 1 | 20 | Area is indicative and will need to be confirmed depending on the waste streams / types of bins and bin sizes to be accommodated
BBLD | Bay - Blood | Yes | 1 | 2 | 1 | 2 | 1 | 2 | May be for whole health care facility.
BPATH | Bay - Pathology | Yes | 1 | 6 | 1 | 9 | 1 | 14 | To support POCT, frozen sections etc. May also be used to store fridges/ freezers for human tissue including freezer to support orthopaedic work (bone fridge). Refer to HPU Section 2.2.11 for further informtion regarding management of formalin.
MED-14 | Medication Room | Yes | 1 | 6 | 1 | 10 | 1 | 14
OFF-WI-5 | Office - Write-up, 3m2 | Yes | 1 | 3 (o) | 4 | 3 (o) | 8 | 3 (o) | Optional depending on staff profile; 1 per 2 Operating Rooms
OFF-S9 | Office - Single Person, 9m2 | Yes |  |  |  |  | 1 | 9 | Requirements will depend on staff profile.
| Radiographer Workroom |  |  |  | 1 | 15 (o) | 1 | 25 (o) | Optional, depending on radiography staff profile based in the operating suite. Work area for imaging processing and workspace.
SRM-15 | Staff Room | Yes | [5 empty] | tbc (o) | Optional, satellite staff room for those units where staff amenities are located on a separate floor.
WCST | Toilet - Staff, 3m2 | Yes |  | 3 |  | 3 |  | 3 | Number and location so staff have access close to where they work
Intradepartmental (discounted) circulation |  |  |  | 40% |  | 40% |  | 40%
PACU / STAGE 1 RECOVERY
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
PT-RS1 | Patient Bay - Recovery, Stage 1 | Yes | 5 | 9 | 11 | 9 | 22 | 9 | 1.5 - 3 bays per Operating Room depending on casemix (refer to HPU Section 2.4.6)
PT-RS1 | Patient Bay - Recovery, Stage 1 | Yes | 1 | 12 (o) | 1 | 12 (o) | 2 | 12 (o) | Optional, for infectious patients, feeding mothers, paediatrics etc. Alternative management of infectious patients includes placing patients last on the list and recovering in the OR. Ratio of enclosed rooms to open bays will depend on the anticipated casemix acknowledging that open bays should be prioritised to support optimal patient observation. Obstetric patients require the larger size space to accomodate a cot and additional staff. Refer to HPU Sections 2.4.6 and 3.4.1.
WCAC | Toilet - Accessible, 6m2 | Yes |  |  |  |  | 1 | 6 | Not frequently used in PACU. Smaller units will share with Stage 2 recovery. Close access to toilets will be required for units where advanced recovery models are an endorsed model of care.
SSTN-10 | Staff Station, 10m2 | Yes | 1 | 10 | 1 | 14 | 2 | 10
BRES | Bay - Resuscitation Trolley | Yes |  | 3 |  | 3 |  | 3 | Access from Operating Rooms and Preoperative Holding Area; number to be determined on local requirements. Number to be determined based on local requirements / casemix. Includes difficult intubation trolley and paediatric trolley where applicable.
BHWS-B | Bay - Handwashing, Type B | Yes | 1 | 1 | 3 | 1 | 6 | 1
BLIN | Bay - Linen | Yes | 1 | 2 | 1 | 2 | 2 | 2 | 1 per 16 spaces
BMEQ | Bay - Mobile Equipment | Yes | 1 | 2 | 1 | 3 | 2 | 2 | Number depends on equipment stored and frequency of use. Include power for recharging equipment.
BBW | Bay - Blanket/ Fluid Warmer | Yes | 1 | 1 | 1 | 1 | 2 | 1 | 1 per 16 spaces
CLN-10 | Clean Store | Yes | 1 | 8 | 1 | 10 | 1 | 10 | Assumed to be shared with pre-operative holding and recovery for smaller scenarios. May be provided as a combined space including medication store depending on local jurisdictional policies.
MED-14 | Medication Room | Yes | 1 | 10 | 1 | 12 | 1 | 14 | Assumed to be shared with pre-operative holding and recovery. May be provided as a combined space including medication store depending on local jurisdictional policies.
DTUR-12 | Dirty Utility, 12m2 | Yes | 1 | 10 | 1 | 12 | 1 | 14 | Direct access from Recovery Area, may be shared with Preoperative Holding Area
STGN-8 | Store - Equipment | Yes |  | 6 |  | 8 |  | 10 | Large enough for cots, IV poles, blood warmers, etc. Wide and shallow room size preferred.
MEET-9 | Meeting Room, 9m2 | Yes |  |  |  |  | 1 | 9 | Optional; may be used for interview and other purposes
WCST | Toilet - Staff | Yes |  |  |  | 3(o) |  | 3 (o) | Number and location dependent on travel distances to staff change rooms.
Intradepartmental circulation |  |  |  | 40% |  | 40% |  | 40%
STAFF AREAS - AMENITIES
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
CHST-35 | Change - Staff (Male/Female/All Gender) | Yes | 1 | 30 | 1 | 100 | 1 | 180 | Indicative only; peak access periods need to be assessed; separate male and female and support for all gender / gender neutral facilities needed in line with local policies.
SRM-35 | Staff Room | Yes | 1 | 20 | 1 | 50 | 1 | 90 | Smaller units may share as appropriate; external window desirable. Satellite staff room may be required where staff amenities are located on a different level to the main OR floor.
WCAC | Toilet - Accessible, 6m2 | Yes |  |  | 1 | 6 | 1 | 6 | Unless readily available elsewhere
Intradepartmental (discounted) circulation |  |  |  | 25% |  | 25% |  | 25%
STAFF AREAS - OFFICE AND SUPPORT
CODE | ROOM/SPACE | SC / SC-D | 2 ORs |  | 8 ORs |  | 16 ORs |  | REMARKS
|  |  | Qty | m2 | Qty | m2 | Qty | m2
OFF-S12 | Office - Single Person, 12m2 | Yes |  |  |  | 12 |  | 12 | Number and area allocation will depend on staff profile and local jurisdictional policies.
OFF-S9 | Office - Single Person, 9m2 | Yes |  | 9 |  | 9 |  | 9 | Number and area allocation will depend on staff profile and local jurisdictional policies.
OFF-WS | Office - Workstation | Yes |  | 4.5 |  | 4.5 |  | 4.5 | Number and area allocation will depend on staff profile and local jurisdictional policies.
BMFD-3 | Bay - Multifunction Device | Yes |  |  | 1 | 3 | 1 | 3 | Assume desktop printers in smaller facilities.
MEET-15 | Meeting Room | Yes |  | 15 |  | 20 |  | 30 | Quantity to be determined by service demand, may be used for educational purposes including simulation training. This will require consideration of the types of training, e.g. advanced life support and manual handling; volume of staff participating; and any networked training arrangements.
MEET-L-55 | Meeting Room | Yes |  |  |  | 30 |  | 50 | Quantity to be determined by service demand, may be used for educational purposes including simulation training. This will require consideration of the types of training, e.g. advanced life support and manual handling; volume of staff participating; and any networked training arrangements.
Intradepartmental (discounted) circulation |  |  |  | 25% |  | 25% |  | 25%